Grand Bay (GND) National Estuarine Research Reserve Water Quality Metadata

January – December 2005

Latest Update:  October 1, 2019

I. Data Set & Research Descriptors

1) Principal investigators & contact persons:

Address:		Grand Bay National Estuarine Research Reserve

6005 Bayou Heron Road

Moss Point, MS  39562

Phone:  (228) 475-7047

Fax:      (228) 475-8097

Contact Persons:	Dr. Mark Woodrey, Research Coordinator

E-mail:  mark.woodrey@dmr.state.ms.us

(228) 475-7047, ext. 104

Christine Walters, Research Technician

E-mail:  christine.walters@dmr.state.ms.us

(228) 475-7047, ext. 103

2) Entry verification:

Deployment data (.DAT files) are uploaded from the YSI 6600 data logger to a Personal Computer (IBM compatible).  Data plots are produced and examined and erroneous data are detected using the PC6000 or EcoWatch software from YSI.  Notes are made of any unusual data during that deployment.  Files are exported from EcoWatch in a comma-delimited format (CSV) and opened in Microsoft Excel for pre-processing with the EQWin format macro that was developed by the Centralized Data Management Office (CDMO).  Data files are revised using the EQWinformat.xls macros that: 1) ensure parameters are in the correct order; 2) convert depth in feet to depth in meters; 3) format the data to the correct number of decimal points; 4) insert identifying station codes; 5) insert a corrected time column; and 6) allows the user to remove any pre- and post-deployment data from the file.  Where deployment overlaps occur between files, the data produced by the newly calibrated sonde is accepted as being the most accurate.

The pre-processed data are then copied into EQWin where: 1) the station and parameters are verified against the database; 2) the file configurations are checked; 3) the date and time formats are verified; 4) data are checked for duplication; 5) data are checked for outliers (data outside the acceptable range for the sensors) and outlier reports produced; and 6) graphs are generated.  The reports and graphs are printed and the data reviewed.  Data at the beginning and end of each data record when the instrument was out of the water are deleted.  These data are identified by field notes that document the times of deployment and retrieval and by the sudden change in Specific Conductivity and Depth records.  Data in the remainder of the record are not deleted unless supporting reasons can be documented.  EQWin is also used to generate statistics, view graphs, create customized queries and reports of the data, cross query the water, weather and nutrient data and finally export the data to the CDMO.

The anomalous data are evaluated to determine whether to flag or delete the suspect data.  Data are flagged if the values are: 1) outside the range expected for the site (see Site Location and Character, Section 5), 2) outside the range of measurements and accuracy established for the sensors (see Section 9), or 3) outside the range established for good water quality conditions (e.g., Turbidity >150 NTU, negative Turbidity, variable DO, or sudden data spikes).  Data outside the acceptable range of water quality for a particular site are investigated for validity based on weather data, field observations, QC checks, data printouts, and instrument diagnostics.  Data are deleted if the anomalies are attributed to: 1) sensor malfunction - if the voltage reading of the sensor is outside the range established for the sensor or the sensor will not calibrate, 2) exposure of sensors during low tide - identified by unusual Depth, Temperature, and Specific Conductivity data, or 3) fouling of sensors by aquatic organisms, debris or sediment - detected by comparing in situ sensor with measurements (see Section 4).  Turbidity readings > 1000 NTU are deleted.  In addition, sensor readings that differ significantly (>10%) from standard solutions (e.g., Conductivity) suggest that the sensor was fouled during deployment.  After corrections are made to the pre-processed data in Excel, it is imported into the EQWin database and text files are submitted to the CDMO.  All raw .csv files are also sent to the CDMO via FTP for archival.  Christine Walters is responsible for data management.

3) Research objectives:

The National Estuarine Research Reserve (NERR) System-wide Monitoring Program (SWMP) was designed to fulfill two major overall goals:  1) to support state-specific non-point pollution control programs by establishing local networks of continuous water quality monitoring stations in representative protected estuarine ecosystems and 2) to develop a nation-wide database on baseline environmental conditions in the NERR system of estuaries.  The specific goal of SWMP is to identify and track short-term variability and long-term changes in the integrity and biodiversity of representative estuarine ecosystems and coastal watersheds for the purpose of contributing to effective national, regional, and site specific coastal zone management.  This comprehensive program consists of three phased components:  1) abiotic factors such as water quality and meteorological monitoring; 2) biodiversity monitoring; and 3) land-use planning analysis.  With the initial focus of phase 1, the NERR SWMP will provide data necessary for intra- and inter- site baseline studies, trend analyses, and impact assessment.

Four long-term monitoring stations have been established across the Grand Bay NERR in order to collect essential baseline water quality information to improve our understanding of the tidal dynamics and freshwater inputs into this system.  Specifically, the Grand Bay NERR system wide monitoring program stations collect continuous data to address the following objectives: (1) track short-term variability and long-term changes in estuarine water parameters within three different regions within the reserve and (2) provide bayou-specific water quality data to be applied towards the development of a hydrologic model for the Grand Bay NERR.

The Grand Bay NERR is part of the Grand Bay coastal streams watershed which consists of three primary sub-watersheds that provide much of the freshwater inputs into the system. The Bayou Heron site, located in the upper reaches of the Bayou Heron sub-watershed, monitors water quality for a semi-pristine area with no development and serves as a reference site for the reserve. The Bayou Cumbest site monitors water quality for the Bayou Cumbest sub-watershed which is a moderately impacted area with some residential housing development and non-point source pollution issues related to failing septic tanks (i.e., elevated levels of fecal coliforms).  Located towards the southern end of the Bangs Lake sub-watershed, the Bangs Lake site monitors water quality for a second impacted area with minimal residential development.  Although this sub-watershed has less development than the Bayou Cumbest sub-watershed, a recent study found extremely elevated levels of fecal coliforms.  Subsequent investigations of the Bangs Lake area showed a small drainage ditch running directly into the northern end of the sub-watershed from several residential housing units with failing septic systems.  The fourth water quality monitoring site was located at Crooked Bayou, near the south-central boundary of the reserve.  During 2005, this site was relocated to Point Aux Chenes Bay, near the southern boundary of the Reserve.  Both of these sites provide baseline data of the relative influence of marine inputs and tidal influence from the East Mississippi Sound.

Current and future research projects will be developed around this monitoring program.  The four sites identified here monitor areas with varying degrees of human disturbance and impacts, providing an excellent framework for developing reserve-wide research projects.

4) Research methods:

Sonde cleaning and calibrations of the DO, Conductivity, Depth, pH, and Turbidity probes are performed as outlined in the YSI manual.  For Conductivity and Salinity, YSI calibrator solution 3169 (50,000 µS/cm) is used without dilution.  For pH calibrations, YSI calibrator solution 3822 and 3823 are used for pH 7 and 10, respectively.  A two-point calibration is used for Turbidity, distilled water serves as the calibrator solution for 0 NTU and YSI 6073G is used for 123 NTU.  Depth is calibrated in air and is barometrically corrected.  The DO membrane is replaced and calibrated every deployment.  It is allowed to stretch for 16-20 hours, and re-calibrated prior to actual deployment.  The sonde is programmed to brush the probes 1 minute prior to the actual measurement.  All data are recorded in Central Standard Time.

One data logger is deployed at each permanent monitoring station in the Grand Bay estuary at all times.  Two permanently assigned data loggers are interchanged among each site.  Sites are accessed using a small skiff equipped with an outboard motor.  During transport, each sonde is wrapped in a white towel soaked in tap water and placed horizontally in a cooler for insulation against jarring.  For deployment, the data loggers are affixed to a PVC rack using a stainless steel hose clamp, as well as, hung by the pail from a PVC hook built into the rack for additional support.  This rack is lowered into a five inch diameter stainless steel (SS) pipe that has been bolted to a log piling driven into the mud at each site.  The SS pipes have cut outs to ensure adequate tidal flushing and exposure of the probes to ambient water conditions.  A grate across the bottom of the pipe prevents the rack from descending beyond the bottom of the pipe and ensures that the logger monitors at the same depth on every deployment.  The pipes, along with the probes, are also coated with an anti-fouling paint to minimize biofouling.

As a quality assurance measure, a discreet reading is taken with a freshly calibrated sonde and recorded on the calibration sheet during sonde deployment and retrieval.  During retrieval, the data loggers are again wrapped in a saturated white towel and placed in a cooler for transport to the lab.  At least two data points are recorded while the sonde is wrapped in the towel to record post-deployment dissolved oxygen in 100% water-saturated air.  The other post-deployment calibrations are performed in the laboratory prior to cleaning to determine if instrument drift has occurred and to evaluate the validity of the data.  After post-deployment calibrations, the sondes are cleaned and stored in calibration cups until the next deployment.

5) Site location and character:

Thousands of years ago, the Escatawpa River built a sizable delta in southeast Mississippi.  The deterioration of this delta began when the Escatawpa River changed its course hundreds of years ago.  In 1999, the retrograding delta became home to the 24th National Esturarine Research Reserve, the Grand Bay NERR.

The four SWMP water quality monitoring sites (all 5 locations used in 2005) within the Grand Bay NERR have a tidal range of approximately 2 feet.  Additional site specific characteristics are as follows:

a) Bayou Heron (BH):  [30° 25.068’ N, 88° 24.324’W].   The Bayou Heron site provides water quality data for the Bayou Heron sub-watershed.  The water depth range is from .16 m to 1.69 m and the salinity range is from 0 ppt to 25.3 ppt.  Salinity tends to decrease rapidly when rain events occur at the site or within the watershed.  Because the sensor is located .25m from the bottom and there are no major sources of freshwater into this bayou, it has been concluded that some source of freshwater exist close to the bottom of the bayou.  Hydrology studies are being designed to further understand this occurrence.  This site tends to become hypoxic during warmer months (March – October).  The bottom habitat is soft sediments.

b) Crooked Bayou (CR):  [30° 21.551’N, 88° 25.202’W] January 1, 2006 – August 9, 2006.  The Crooked Bayou site is not associated with a freshwater sub-watershed but rather is located to provide data on the marine inputs and tidal influence of the East Mississippi Sound into the reserve.  The water depth range is from 0 m to 1.27 m and the salinity range is from 9.1 ppt to 29.1 ppt.  This site also contains soft sediments with unconsolidated oyster shell reefs.

c) Point Aux Chenes (PC):  [30° 20.916’N, 88° 25.112’W] August 11, 2006 – December 31, 2006.  The Point Aux Chenes site is the most southern water quality site within the boundaries of Grand Bay NERR.  It is located in Point Aux Chenes Bay and is highly influenced by the Mississippi Sound.   The water depth range is from .3 m to 1.6 m and the salinity range is from 17.2 ppt to 32.1 ppt.  The bottom habitat is soft sediments.

d) Bayou Cumbest (BC):  [30° 23.016’N, 88° 26.184’W].   The Bayou Cumbest site monitors water quality for the Bayou Cumbest sub-watershed.  The water depth range is from 0 m to 1.10 m and the salinity range is from 0.5 ppt to 27.8 ppt.  The site is located in soft sediments with unconsolidated oyster shell reefs.

e) Bangs Lake (BL):  [30° 21.426’N, 88° 27.774’W].  The Bangs Lake site monitors the water quality of the Bangs Lake sub-watershed.  The water depth range is from 0 m to 1.37 m and the salinity range is from 9.6 ppt to 28.0 ppt. The site is located on a soft sediment bottom.

6) Data Collection Periods:

Data loggers were first deployed at Bayou Heron (BH) and Crooked Bayou (CR) on January 22, 2004.  Additional data loggers were deployed at Bayou Cumbest (BC) and Bangs Lake (BL) on March 25, 2004.   Bayou Heron, Bayou Cumbest, and Bangs Lake have been in service continuously since inception.  Crooked Bayou (CR) was relocated to the weather station in August 2004 [30° 21.551’N, 88° 25.202’W] due to the loss of the permanent log piling at the original site [30° 21.597’N, 88° 25.143’W].  During August 2005, the Crooked Bayou site was discontinued due to the magnitude of data lost during low tide events.  The Point Aux Chenes site (PC) [30° 20.916’N, 88° 25.112’W] was designated to replace the Crooked Bayou site.  Also during August 2005, all sites were lowered from .5m to .25m to increase data collection during low tide events.  In 2005, the collection dates and times were as follows:

Site	Deploy Date	Time	Retrieve Date	Time

Bayou Heron

12/07/04	1330	01/14/05	0830

01/14/05	0900	02/16/05	1430

02/16/05	1500	03/17/05	0930

03/17/05	1000	04/14/05	1230

04/14/05	1300	05/02/05	1100

05/02/05	1130	05/17/05	0900

05/17/05	0930	06/23/05	1300

06/23/05	1330	07/15/05	0900

07/15/05	0930	08/11/05	0930

08/11/05	1000	08/25/05	1230

08/25/05	1300	09/15/05	0830

09/15/05	0900	10/07/05	1200

10/07/05	1230	10/28/05	0930

10/28/05	1000	11/22/05	0800

11/22/05	0830	12/21/05	1300

12/21/05	1330	01/10/06	0930

Site	Deploy Date	Time	Retrieve Date	Time

Crooked Bayou

12/07/04	1400	01/06/05	1200

01/06/05	1230	02/16/05	1500

02/16/05	1530	03/17/05	1000

03/17/05	1030	04/14/05	1300

04/14/05	1330	05/02/05	1230

05/02/05	1330	05/17/05	0930

05/17/05	1000	06/23/05	1400

06/23/05	1500	07/15/05	0930

07/15/05	1000	08/09/05	0700

Site	Deploy Date	Time	Retrieve Date	Time

Point Aux Chenes

08/11/05	1030	08/25/05	1300

08/25/05	1330	09/21/05	0900

09/21/05	1000	10/07/05	1230

10/07/05	1300	10/28/05	1000

10/28/05	1030	11/28/05	0830

11/28/05	0900	12/21/05	1330

12/21/05	1430	01/10/06	1000

Site	Deploy Date	Time	Retrieve Date	Time

Bayou Cumbest

12/07/04	1530	01/06/05	1100

01/06/05	1130	02/16/05	1530

02/16/05	1600	03/17/05	1030

03/17/05	1100	04/14/05	1330

04/14/05	1400	05/02/05	1200

05/02/05	1230	05/17/05	0930

05/17/05	1000	06/23/05	1330

06/23/05	1400	07/15/05	1000

07/15/05	1030	08/08/05	0930

08/11/05	1100	08/25/05	1330

08/25/05	1400	09/15/05	1000

09/15/05	1030	10/07/05	1300

10/07/05	1330	10/28/05	1200

10/28/05	1230	11/22/05	0900

11/22/05	0930	12/21/05	1400

12/21/05	1430	01/10/06	1030

Site	Deploy Date	Time	Retrieve Date	Time

Bangs Lake

12/07/04	1500	01/06/05	1130

01/06/05	1200	02/16/05	1530

02/16/05	1600	03/17/05	1100

03/17/05	1130	04/14/05	1400

04/14/05	1430	05/02/05	1130

05/02/05	1200	05/17/05	1000

05/17/05	1100	06/23/05	1330

06/24/05	1430	07/15/05	1030

07/15/05	1100	08/09/05	0830

08/11/05	1130	08/25/05	1330

08/25/05	1430	09/15/05	1100

09/15/05	1130	10/07/05	1300

10/07/05	1330	10/28/05	1100

10/28/05	1200	11/22/05	0900

11/22/05	0930	12/21/05	1430

12/21/05	1500	01/10/06	1030

7) Distribution

According to the Ocean and Coastal Resource Management Data Dissemination Policy for the NERRS System-wide Monitoring Program, NOAA/ERD retains the right to analyze, synthesize and publish summaries of the NERRS System-wide Monitoring Program data.  The Principal Investigator (PI) retains the right to be fully credited for having collected and processed the data.  Following academic courtesy standards, the PI and NERR site where the data were collected will be contacted and fully acknowledged in any subsequent publications in which any part of the data are used.  Manuscripts resulting from the NOAA/OCRM supported research that are produced for publication in open literature, including refereed scientific journals, will acknowledge that the research was conducted under an award from the Estuarine Reserves Division, Office of Ocean and Coastal Resource Management, National Ocean Service, National Oceanic and Atmospheric Administration.  The data set enclosed within this package/transmission is only as good as the quality assurance and quality control procedures outlined by the enclosed metadata reporting statement.  The user bears all responsibility for its subsequent use/misuse in any further analyses or comparisons.  The Federal government does not assume liability to the Recipient of third persons, nor will the Federal government reimburse or indemnify the Recipient for its liability due to any losses resulting in any way from the use of this data.

NERR water quality data and metadata can be obtained from the Research Coordinator at the individual NERR site (please see section 1. Principal investigators and contact persons), from the Data Manager at the Centralized Data Management Office (please see personnel directory under general information link on CDMO home page) and online at the CDMO home page http://cdmo.baruch.sc.edu.  Data are available in text tab-delimited format, Microsoft Excel spreadsheet format and comma-delimited format.

8) Associated researchers and projects:

Several research and monitoring projects are currently using the water quality data from the Grand Bay NERR.  These projects are listed below:

- A Regional Synthesis of Hurricane Impacts in the Gulf of Mexico:  Water quality changes at several National Estuarine Research Reserves Sites – Collaborative project between the Rookery Bay, Apalachicola, Weeks Bay, and Grand Bay NERRs;

- Development of a Hydrology Model for the Grand Bay NERR – collaborative project between Grand Bay NERR staff and Mississippi State University;

- Distribution and Abundance of Submerged Aquatic Vegetation at the Grand Bay NERR – a project being conducted by Grand Bay NERR staff;

- Influence of Climatic Factors on Fecal Coliform Dynamics in the Grand Bay NERR  - collaborators: Dr. Paulinus Chigbu (Jackson State University - JSU), Scott Gordon (MS Department of Marine Resources - DMR), Tom Strange (Stennis Space Center), Dr. Paul Tchounwou (JSU), Dr. Ibrahim Farah (JSU);

- Assessment of the Microbial Quality of Water in the Pearl River and the Grand Bay NERR - collaborators: Dr. Ibrahim Farah (JSU),  Dr. Paul Tchounwou (JSU), Dr. Paulinus Chigbu (JSU), Dr. Mark Woodrey (Mississippi State University/Grand Bay NERR);

- Assessment of the Spatial and Temporal Distribution of Fish and Macroinvertebrates in Relation to Water Quality in the Grand Bay NERR - collaborators: Dr. Paulinus Chigbu (JSU), Dr. Kenneth Goldman (JSU), Dr. Paul Tchounwou (JSU), Dr. Ibrahim Farah (JSU), Dr. Mark Woodrey (MSU/NERR);

- Ecotoxicology Studies of the Pearl River and the Grand Bay NERR - collaborators: Dr. Paul Tchounwou (JSU), Dr. Paulinus Chigbu (JSU), Dr. Ibrahim Farah (JSU), Dr. Mark Woodrey (MSU/NERR);

- Breeding Ecology and Ecotoxiocology of Clapper Rails in Coastal Mississippi Marshes – collaborators: Dr. Bob Cooper (University of Georgia), Scott Rush (Ph.D. Candidate, UGA), Dr. Courtney Conway (University of Arizona), Dr. Mark Woodrey (MSU/NERR); and

- Ecotoxicology of Oysters of the Grand Bay NERR – Jim Weston (Ph.D. Candidate, Univeristy of Missisippi).

II. Physical Structure Descriptors

9) Sensor Specifications:

YSI 6600/YSI 6600EDS data logger

Parameter: Temperature

Units: Celsius (C)

Sensor Type: Thermistor

Model #: 6560

Range: -5 to 45 °C

Accuracy: +/-0.15 °C

Resolution: 0.01 °C

Parameter: Conductivity

Units: milli-Siemens per cm (mS/cm)

Sensor Type: 4-electrode cell with autoranging

Model #: 6560

Range: 0 to 100 mS/cm

Accuracy: +/-0.5% of reading + 0.001 mS/cm

Resolution: 0.001 mS/cm to 0.1 mS/cm (range dependent)

Parameter: Salinity

Units: parts per thousand (ppt)

Sensor Type: Calculated from conductivity and temperature

Range: 0 to 70 ppt

Accuracy: +/- 1.0% of reading or 0.1 ppt, whichever is greater

Resolution: 0.01 ppt

Parameter: Dissolved Oxygen % saturation (EDS sonde)

Units: percent air saturation (%)

Sensor Type: Rapid Pulse – Clark type, polarographic

Model #: 6562

Range: 0 to 500 % air saturation

Accuracy: 0-200 % air saturation, +/- 2 % of the reading or 2 % air saturation, whichever is greater; 200-500 % air saturation, +/- 6 % of the reading

Resolution: 0.1 % air saturation

Parameter: Dissolved Oxygen mg/L (Calculated from % air saturation, temperature and salinity) (EDS sonde)

Units: milligrams per Liter (mg/L)

Sensor Type: Rapid Pulse – Clark type, polarographic

Model #: 6562

Range: 0 to 50 mg/L

Accuracy: 0 to 20 mg/L, +/- 2 % of the reading or 0.2 mg/L, whichever is greater; 20 to 50 mg/L, +/- 6 % of the reading

Resolution: 0.01 mg/L

Parameter: Non-Vented Level – Shallow (Depth)

Units: feet or meters (ft or m)

Sensor Type: Stainless steel strain gauge

Range: 0 to 30 ft (9.1 m)

Accuracy: +/- 0.06 ft (0.018 m)

Resolution: 0.001 ft (0.001 m)

Parameter: pH (EDS sonde)

Units: units

Sensor Type: Glass combination electrode

Model #: 6561

Range: 0 to 14 units

Accuracy: +/- 0.2 units

Resolution: 0.01 units

Parameter: Turbidity

Units: nephelometric turbidity units (NTU)

Sensor Type: Optical, 90 ° scatter, with mechanical cleaning

Model #: 6136

Range: 0 to 1000 NTU

Accuracy: +/- 5 % reading or 2 NTU (whichever is greater)

Resolution: 0.1 NTU

Dissolved Oxygen qualifier:  The reliability of the dissolved oxygen (DO) data after 96 hours post-deployment for non-EDS (Extended Deployment System) data sondes may be problematic due to fouling which forms on the DO probe membrane during some deployments (Wenner et al. 2001).  Many reserves have upgraded to YSI 6600 EDS data sondes, which increase DO accuracy and longevity by reducing the environmental effects of fouling.  The user is therefore advised to consult the metadata and to exercise caution when utilizing the DO data beyond the initial 96-hour time period.  However, this potential drift is not always problematic for some uses of the data, ie. periodicity analysis. It should be noted that the amount of fouling is site specific and that not all data are affected.  The Research Coordinator at the specific NERR site should be contacted concerning the reliability of the DO data because of the site and seasonal variation in the fouling of the DO sensor. All data sondes used at GND NERR sites in 2005 were EDS models.

Depth qualifier:  The NERRS System-Wide Monitoring Program utilizes YSI data sondes that can be equipped with either depth or water level sensors.  Both sensors measure water depth, but by convention, level sensors refer to atmospherically vented measurements and depth refers to non-vented measurements.  Standard calibration protocols for the non-vented sensor use the atmosphere pressure at the time of calibration.  Therefore, changes in atmospheric pressure between calibrations appear as changes in water depth.  The error is equal to approximately 1.03 cm for every 1 millibar change in atmospheric pressure.  This error is eliminated for level sensors because they are vented to the atmosphere throughout the deployment time interval.  If proper atmospheric pressure data are available, non-vented sensor depth measurements can be corrected for deployments between calibrations. Readings for both vented and non-vented sensors are automatically compensated for water density changes due to variations in temperature and salinity.  The Research Coordinator at the specific NERR site should be contacted in order to obtain information regarding atmospheric pressure data availability. All data sondes used at the GND NERR sites in 2005 were non-vented models.

10) Coded variable definitions:

Sampling station:		Sampling site code:		Station Code:

Bayou Heron				BH			gndbhwq

Crooked Bayou			CR			gndcrwq

Bayou Cumbest			BC			gndbcwq

Bangs Lake				BL			gndblwq

Point Aux Chenes			PC			gndpcwq

11) Anomalous/Suspect Data:

On 01/14/2019 this dataset was updated to include embedded QAQC flags for anomalous/suspect data.  System-wide monitoring data beginning in 2007 were processed to allow for QAQC flags and codes to be embedded in the data files rather than detailed in the metadata alone (as in the anomalous/suspect, deleted, and missing data sections above).  Prior to 2006, rejected data were deleted from the dataset so they are unavailable to be used at all, but suspect data were only noted in the metadata document.  Suspect data flags <1> were embedded retroactively in order to allow suspect data to be easily identified and filtered from the dataset if desired for analysis and reporting purposes.  No other flags or codes were embedded in the dataset and users should still refer to the detailed explanations above for more information.

There were instances at this NERR site where negative depth readings were not deleted.  These data were considered to be accurate based upon the changes in barometric pressure during storm events experienced at Grand Bay.  For more information on non-vented level – shallow depth sensors and the effects of barometric pressure, see Sensor Specification (9.).

There were a few instances at this NERR site where turbidity recorded small negative values (-0001 and -0002). Because turbidity has a range of accuracy of +/-2 NTU, the technician did not edit or deleted these values in any way.

Bayou Heron

Turbidity readings from 07/06/05 0500 – 07/06/05 0600 are unusually high.

Crooked Bayou

There were a large number of negative depth readings recorded at this site, these values should be considered anomalous and used with caution, however, negative depths do not necessarily mean that the sonde was out of the water (see above).  All data recorded while the sonde was out of the water, as evidenced by Specific Conductivity and other parameters, were deleted.

Bayou Cumbest

None

Bangs Lake

a) Salinity readings from 04/18/05 0130 – 04/18/05 0530 are lower than usual.

b) pH readings are unusually low from 04/15/05 0330 – 04/19/05 1130 due to highly acidic runoff escaping a breached containment system owned and operated by a nearby fertilizer manufacturer.

c) Dissolved oxygen readings are unusually high after Hurricane Katrina from 09/02/05 0500 – 09/05/05 1930.

Point Aux Chenes

None

12) Deleted Data:

Note: The average overall depth of 70-80% of the waterways of the Grand Bay NERR is between .6-.9 meters.  With an average tidal range of .6 meters, instruments located .5 meters (SWMP protocol) from the bottom were often out of the water.  Amendments to SWMP protocol were made at the Grand Bay NERR during 2005 to lessen the effects of the tide on data collection.  The instruments were lowered to .25 meters during August to increase the number of data points collected during low tide events.  All data were deleted when sudden changes in Specific Conductivity and Depth records indicated that the sondes were out of the water.

Bayou Heron

January

1/06/05 0100 – 1/14/05 0830 dissolved oxygen percent and mg/L were deleted due to bad post-calibration of sensor

1/11/05 1130 – 1/14/05 0830 turbidity deleted due to fouling and bad post-calibration of sensor

May

5/17/05 1500 – 6/23/05 1300 dissolved oxygen percent and mg/L were deleted due to a hole in the sensor’s membrane

5/28/05 2030 – 6/11/05 1230 turbidity deleted due to malfunctioning of sensor

June

6/20/05 1200 – 6/22/05 2000 turbidity deleted due to malfunctioning of sensor

August

8/18/05 0830 – 8/18/05 0900 all readings were deleted due to depth values indicating sonde was out of water for deployment sleeve replacement

September

9/14/05 0900 – 9/15/05 0830 turbidity deleted due to fouling

October

10/08/05 1400 – 10/28/05 0930 turbidity deleted due to fouling and bad post-calibration of sensor

November

11/22/05	0830 turbidity deleted due to sensor out of range

Crooked Bayou

January

All data were deleted due to depth values indicating sonde was out of water:

1/5/2005	130	-	1/05/2005	400

1/6/2005	100	-	1/6/2005	730

1/7/2005	100	-	1/7/2005	900

1/8/2005	200	-	1/8/2005	1100

1/9/2005	130	-	1/9/2005	1400

1/10/2005	300	-	1/10/2005	1400

1/11/2005	400	-	1/11/2005	1400

1/12/2005	600	-	1/12/2005	1330

1/13/2005	1500	-	1/13/2005	1530

1/14/2005	600	-	1/14/2005	1600

1/14/2005	1730	-	1/14/2005	2100

1/15/2005	630	-	1/15/2005	1130

1/15/2005	1900	-	1/17/2005	1330

1/17/2005	1600	-	1/18/2005	1100

1/18/2005	2100	-	1/19/2005	1030

1/19/2005	2330	-	1/20/2005	1100

1/21/2005	130	-	1/21/2005	1030

1/22/2005	230	-	1/22/2005	1130

1/22/2005	2300	-	1/23/2005	1600

1/24/2005	0000	-	1/24/2005	1430

1/25/2005	230	-	1/25/2005	1430

1/26/2005	500	-	1/26/2005	1200

1/27/2005	400	-	1/27/2005	1430

February

Turbidity deleted due to sensor out of range:

2/02/05 	1100

All data were deleted due to depth values indicating sonde was out of water:

2/3/2005	0000	-	2/3/2005	830

2/3/2005	2300	-	2/4/2005	1100

2/5/2005	0000	-	2/5/2005	830

2/6/2005	100	-	2/6/2005	930

2/7/2005	400	-	2/7/2005	1030

2/8/2005	630	-	2/8/2005	830

2/9/2005	700	-	2/9/2005	900

2/10/2005	430	-	2/10/2005	1400

2/10/2005	1530

2/11/2005	530	-	2/11/2005	900

2/11/2005	1100	-	2/11/2005	1330

2/11/2005	1800	-	2/11/2005	1930

2/12/2005	530	-	2/12/2005	830

2/16/2005	0000	-	2/16/2005	500

2/17/2005	100	-	2/17/2005	900

2/17/2005	2330	-	2/18/2005	930

2/19/2005	030	-	2/19/2005	930

2/20/2005	200	-	2/20/2005	930

2/21/2005	330	-	2/21/2005	930

2/22/2005	330	-	2/22/2005	1030

2/23/2005	330	-	2/23/2005	1000

2/24/2005	600	-	2/24/2005	930

2/25/2005	400	-	2/25/2005	1100

2/26/2005	500

2/28/2005	1730	-	3/1/2005	1030

March

Turbidity deleted due to sensor out of range:

3/15/05 	1330

3/26/05 	2200

3/30/05	0900

All data were deleted due to depth values indicating sonde was out of water:

3/1/2005	1630	-	3/2/2005	830

3/2/2005	1900	-	3/3/2005	730

3/3/2005	2130	-	3/4/2005	800

3/4/2005	2300	-	3/5/2005	1030

3/5/2005	2230	-	3/6/2005	1000

3/7/2005	030	-	3/7/2005	1000

3/8/2005	300	-	3/8/2005	1500

3/9/2005	130	-	3/9/2005	1130

3/10/2005	130	-	3/10/2005	1800

3/11/2005	430	-	3/12/2005	700

3/12/2005	1630	-	3/13/2005	630

3/13/2005	1830	-	3/14/2005	330

3/14/2005	1630	-	3/15/2005	530

3/15/2005	2130	-	3/16/2005	030

3/16/2005	2030	-	3/17/2005	900

3/17/2005	2030	-	3/18/2005	830

3/18/2005	2200	-	3/19/2005	830

3/19/2005	2330	-	3/20/2005	900

3/21/2005	200	-	3/21/2005	900

3/23/2005	430	-	3/23/2005	730

3/24/2005	300	-	3/24/2005	800

3/27/2005	1930	-	3/28/2005	730

3/28/2005	1630	-	3/29/2005	630

3/29/2005	1700	-	3/30/2005	400

3/30/2005	1930	-	3/31/2005	430

3/31/2005	2100	-	4/1/2005	000

April

Turbidity deleted due to sensor out of range:

4/3/05 		2130

4/27/05	0800

All data were deleted due to depth values indicating sonde was out of water:

4/1/2005	2130	-	4/2/2005	800

4/2/2005	2030	-	4/3/2005	930

4/3/2005	2200	-	4/4/2005	1030

4/5/2005	130	-	4/5/2005	900

4/8/2005	1630	-	4/8/2005	2030

4/9/2005	1630	-	4/9/2005	2030

4/13/2005	2000	-	4/14/2005	430

4/14/2005	1930	-	4/15/2005	630

4/15/2005	2100	-	4/16/2005	600

4/16/2005	2200	-	4/17/2005	700

4/17/2005	2330	-	4/18/2005	700

4/19/2005	100	-	4/19/2005	400

4/23/2005	1500	-	4/24/2005	130

4/24/2005	1530	-	4/25/2005	100

4/25/2005	1700	-	4/25/2005	2300

4/26/2005	1830	-	4/27/2005	200

4/27/2005	1900	-	4/28/2005	400

4/28/2005	2030	-	4/29/2005	200

4/30/2005	2230	-	5/1/2005	630

May

Turbidity deleted due to sensor out of range:

5/9/05 		1600

5/13/05 	1200

All data were deleted due to depth values indicating sonde was out of water:

5/1/2005	2330	-	5/2/2005	600

5/3/2005	030	-	5/3/2005	500

5/4/2005	030	-	5/4/2005	130

5/6/2005	1600	-	5/6/2005	1830

5/7/2005	1500	-	5/7/2005	2230

5/8/2005	1600	-	5/8/2005	2230

5/9/2005	1700	-	5/9/2005	2330

5/10/2005	1900	-	5/11/2005	100

5/11/2005	1830	-	5/12/2005	230

5/12/2005	2000	-	5/13/2005	200

5/16/2005	000	-	5/16/2005	330

5/21/2005	1530	-	5/21/2005	2130

5/22/2005	1530	-	5/22/2005	2130

5/23/2005	1600	-	5/23/2005	2330

5/24/2005	1630	-	5/25/2005	130

5/25/2005	1800	-	5/26/2005	230

5/26/2005	1900	-	5/27/2005	230

5/27/2005	2000	-	5/28/2005	330

5/28/2005	2100	-	5/29/2005	330

5/29/2005	2200	-	5/29/2005	2230

June

Turbidity deleted due to malfunctioning of sensor:

6/5/05 		2330 	– 	6/23/05 	1400

6/30/05 	0900 	– 	7/1/2005 	1300

6/10/05 1830 – 6/23/05 1400 dissolved oxygen percent and mg/L were deleted due to bad post-calibration of sensor

All data were deleted due to depth values indicating sonde was out of water:

6/6/2005	1730	-	6/6/2005	2030

6/6/2005	2230

6/6/2005	2330

6/7/2005	1730	-	6/7/2005	2330

6/8/2005	1900	-	6/9/2005	000

6/9/2005	2030	-	6/9/2005	2330

6/17/2005	1630	-	6/17/2005	2000

6/18/2005	1600	-	6/18/2005	1830

6/19/2005	1400	-	6/19/2005	2200

6/20/2005	1500	-	6/20/2005	2200

6/21/2005	1630	-	6/21/2005	2300

6/22/2005	1730	-	6/23/2005	030

6/23/2005	1830	-	6/24/2005	200

6/24/2005	1930	-	6/25/2005	200

6/25/2005	2030	-	6/26/2005	200

6/30/2005	1530	-	6/30/2005	1800

July

Turbidity deleted due to malfunctioning of sensor:

7/2/05 		1030 	– 	7/3/05 		1230

7/10/05 	1700 	–	7/15/05 	0930

7/22/05 	0900 	– 	7/22/05 	1830

7/26/05 	0200 	– 	7/27/05 	1100

7/29/05 	0400 	– 	8/9/05 		0700

7/15/05 1000 – 8/9/05 0700 pH deleted, no sensor installed

All data were deleted due to depth values indicating sonde was out of water:

7/1/2005	1330	-	7/1/2005	2100

7/2/2005	1330	-	7/2/2005	2330

7/3/2005	1400	-	7/3/2005	2300

7/4/2005	1430	-	7/4/2005	2300

7/5/2005	1630	-	7/5/2005	2030

7/6/2005	1800	-	7/7/2005	030

7/7/2005	1800	-	7/8/2005	030

7/8/2005	1830	-	7/9/2005	100

7/9/2005	2000	-	7/9/2005	2100

7/17/2005	1500	-	7/17/2005	1800

7/18/2005	1500	-	7/18/2005	2000

7/19/2005	1600	-	7/19/2005	2100

7/20/2005	1700	-	7/20/2005	2200

7/21/2005	1730	-	7/22/2005	000

7/22/2005	1900	-	7/23/2005	030

7/23/2005	1830	-	7/24/2005	200

7/24/2005	2130	-	7/25/2005	100

7/28/2005	1330	-	7/28/2005	1900

7/29/2005	1230	-	7/29/2005	1930

7/30/2005	1230	-	7/30/2005	2030

7/31/2005	1400	-	7/31/2005	2030

August

All data were deleted due to depth values indicating sonde was out of water:

8/1/2005	1430	-	8/1/2005	2130

8/2/2005	1430	-	8/2/2005	2230

8/3/2005	1630	-	8/3/2005	2230

8/4/2005	1630	-	8/4/2005	2200

8/5/2005	1900	-	8/5/2005	2130

8/6/2005	1830	-	8/6/2005	2000

8/6/2005	2200	-	8/6/2005	2300

Point Aux Chenes

August

Turbidity deleted due to sensor out of range:

8/19/05	0100

8/20/05 0400 – 8/25/05 1300 dissolved oxygen percent and mg/L were deleted due to bad post-calibration of sensor

September

9/1/05 1330 – 9/21/05 0900 dissolved oxygen percent and mg/L were deleted due to bad post-calibration of sensor

9/4/05 0230 – 9/21/05 0900 salinity, specific conductivity, dissolved oxygen percent and dissolved oxygen mg/L deleted due to bad post-calibration of sensor

Turbidity deleted due to sensor out of range:

9/6/05		1430

9/29/05 	0800	-	9/29/05	0900

9/29/05	1030

9/29/05	0130

Turbidity deleted due to malfunctioning of sensor:

9/5/05		1530	-	9/7/05		0600

9/21/05 	1730 	– 	9/25/05 	2330

October

Turbidity deleted due to malfunctioning of sensor:

10/1/05	0530	-	10/1/05	0600

10/1/05 	1730 	– 	10/7/05 	1230

Dissolved oxygen percent and mg/L were deleted due to bad post-calibration of sensor:

10/5/05	1030	-	10/7/05	1230

10/21/05	0230	-	10/28/05	1000

November

Turbidity deleted due to malfunctioning of sensor:

11/7/05	2330	-	11/28/05	0830

Bayou Cumbest

January

All data were deleted due to depth values indicating sonde was out of water:

1/8/2005	530	-	1/8/2005	830

1/9/2005	400	-	1/9/2005	1130

1/10/2005	530	-	1/10/2005	1200

1/11/2005	700	-	1/11/2005	1130

1/14/2005	900	-	1/14/2005	1100

1/16/2005	830

1/16/2005	2000	-	1/17/2005	1000

1/17/2005	2130	-	1/18/2005	0730

1/19/2005	030	-	1/19/2005	730

1/20/2005	200	-	1/20/2005	900

1/21/2005	500	-	1/21/2005	800

1/22/2005	600	-	1/22/2005	830

1/23/2005	100	-	1/23/2005	1300

1/24/2005	330	-	1/24/2005	1130

1/25/2005	500	-	1/25/2005	1130

1/27/2005	630	-	1/27/2005	1100

February

All data were deleted due to depth values indicating sonde was out of water:

2/4/2005	230	-	2/4/2005	700

2/5/2005	300	-	2/5/2005	600

2/6/2005	400	-	2/6/2005	630

March

All data were deleted due to depth values indicating sonde was out of water:

3/18/2005	100	-	3/18/2005	500

3/28/2005	2030	-	3/29/2005	0000

April

Turbidity deleted due to malfunctioning of sensor:

4/10/05	1200	-	4/13/05	300

All data were deleted due to depth values indicating sonde was out of water:

4/3/2005	100	-	4/3/2005	630

4/4/2005	230	-	4/4/2005	730

4/13/2005	2330	-	4/14/2005	200

4/14/2005	2200	-	4/15/2005	300

4/16/2005	0000	-	4/16/2005	400

4/17/2005	0000	-	4/17/2005	500

4/23/2005	1730

4/23/2005	2130	-	4/23/2005	2200

4/24/2005	1930	-	4/24/2005	2000

4/24/2005	2100	-	4/24/2005	2300

May

Turbidity deleted due to sensor out of range:

5/29/05	700

All data were deleted due to depth values indicating sonde was out of water:

5/1/2005	130	 -	5/1/2005	500

5/24/2005	1930	 -	5/24/2005	2330

5/25/2005	2130	 -	5/26/2005	0000

June

Turbidity deleted due to sensor out of range:

6/8/05		1200

Turbidity deleted due to malfunctioning of sensor:

6/11/05	1430	-	6/12/05	2130

6/15/05	0000	-	6/23/05	1330

Dissolved oxygen percent and mg/L were deleted due to fouling:

6/17/05	1730	-	6/23/05	1330

July

Turbidity deleted due to malfunctioning of sensor:

7/2/05		200	-	7/9/05		1000

Turbidity deleted, erroneous reading (extremely high)

7/26/05	1030

August

Turbidity deleted due to malfunctioning of sensor:

8/2/05		1600	-	8/2/05		2030

Turbidity deleted due to sensor out of range:

8/3/05		1530

8/4/05		1230

8/5/06		1430

8/6/05		1300

8/7/05		700

8/7/05		1600

8/8/05		930

September

Turbidity deleted due to malfunctioning of sensor:

9/4/05		2130	-	9/5/05		600

9/30/05	2330	-	10/7/05	1300

October

All data were deleted due to depth values indicating sonde was out of water:

10/24/2005	1030	 -	10/24/2005	1700

10/25/2005	1200	-	10/25/2005	1700

November

Turbidity deleted due to malfunctioning of sensor:

11/9/05	2330	-	11/10/05	200

11/11/05	1530	-	11/12/05	130

11/12/05	900	-	11/12/05	1030

All data were deleted due to depth values indicating sonde was out of water:

11/16/2005	730	-	11/16/2005	1000

11/17/2005	500	-	11/17/2005	1230

11/18/2005	630	-	11/18/2005	1230

11/21/2005	1130	-	11/21/2005	1800

11/22/2005	900	-	11/22/2005	1430

11/30/2005	430	-	11/30/2005	630

December

All data were deleted due to depth values indicating sonde was out of water:

12/1/2005 	530	-	12/1/2005 	830

12/2/2005 	530 	- 	12/2/2005	1100

12/5/2005 	1200 	- 	12/5/2005	1300

12/13/2005 	400 	- 	12/13/2005	800

12/16/2005 	730 	- 	12/16/2005	1030

12/19/2005 	800 	- 	12/19/2005	1230

12/20/2005 	830 	- 	12/20/2005	1030

12/27/2005 	200 	-	12/27/2005	730

12/29/2005 	430 	- 	12/29/2005 	930

12/30/2005 	500 	- 	12/30/2005	930

12/31/2005 	730 	- 	12/31/2005	930

Bangs Lake

January

All data were deleted due to depth values indicating sonde was out of water:

1/8/2005	500	-	1/8/2005	830

1/9/2005	400	-	1/9/2005	1100

1/10/2005	500	-	1/10/2005	1200

1/11/2005	630	-	1/11/2005	1130

1/14/2005	900	-	1/14/2005	1030

1/16/2005	2000	-	1/17/2005	930

1/17/2005	2130	-	1/18/2005	730

1/19/2005	030	-	1/19/2005	730

1/20/2005	200	-	1/20/2005	900

1/21/2005	400	-	1/21/2005	800

1/22/2005	530	-	1/22/2005	900

1/23/2005	100	-	1/23/2005	1230

1/24/2005	330	-	1/24/2005	1200

1/25/2005	500	-	1/25/2005	1130

1/27/2005	600	-	1/27/2005	1100

February

Turbidity deleted due to sensor out of range:

2/21/05	1730

2/27/05	930

All data were deleted due to depth values indicating sonde was out of water:

2/4/2005	200	-	2/4/2005	700

2/5/2005	230	-	2/5/2005	630

2/6/2005	330	-	2/6/2005	700

2/10/2005	930	-	2/10/2005	1100

2/18/2005	230	-	2/18/2005	600

March

Turbidity deleted due to sensor out of range:

3/2/05		700

All data were deleted due to depth values indicating sonde was out of water:

3/1/2005	400	-	3/1/2005	700

3/1/2005	2030	-	3/2/2005	500

3/2/2005	2330	-	3/3/2005	300

3/4/2005	030	-	3/4/2005	330

3/5/2005	130	-	3/5/2005	730

3/6/2005	100	-	3/6/2005	830

3/7/2005	300	-	3/7/2005	800

3/8/2005	430	-	3/8/2005	1130

3/9/2005	530	-	3/9/2005	630

3/10/2005	330	-	3/10/2005	1000

3/11/2005	1800	-	3/11/2005	1830

3/14/2005	2000	-	3/15/2005	200

3/17/2005	130	-	3/17/2005	500

3/18/2005	0000	-	3/18/2005	530

3/19/2005	130	-	3/19/2005	600

3/20/2005	300	-	3/20/2005	600

3/27/2005	2230	-	3/28/2005	100

3/28/2005	1900	-	3/29/2005	300

3/29/2005	2100	-	3/29/2005	2200

April

Turbidity deleted due to sensor out of range:

4/3/05		1630

4/20/05	1300

All data were deleted due to depth values indicating sonde was out of water:

4/1/2005	2330	-	4/2/2005	600

4/2/2005	2300	-	4/3/2005	730

4/4/2005	100	-	4/4/2005	800

4/13/2005	2200	-	4/14/2005	230

4/14/2005	2130	-	4/15/2005	300

4/15/2005	2330	-	4/16/2005	400

4/16/2005	2330	-	4/17/2005	530

4/23/2005	1700	-	4/23/2005	1900

4/23/2005	2100	-	4/23/2005	2230

4/24/2005	1700	-	4/24/2005	2300

4/25/2005	1830	-	4/25/2005	2030

4/26/2005	2000	-	4/27/2005	0000

4/27/2005	2030	-	4/28/2005	300

May

Turbidity deleted due to sensor out of range:

5/2/05		830

5/24/05	1400

Turbidity deleted due to malfunctioning of sensor:

5/31/05	130	-	6/23/05	1330

All data were deleted due to depth values indicating sonde was out of water:

5/1/2005 	030	- 	5/1/2005 	500

5/7/2005 	1630	- 	5/7/2005 	1900

5/8/2005 	1830	- 	5/8/2005 	1930

5/9/2005 	1930	- 	5/9/2005 	2030

5/11/2005 	2030	- 	5/11/2005 	2200

5/12/2005 	2200	- 	5/12/2005 	2330

5/23/2005 	1730 	- 	5/23/2005	2200

5/24/2005	1800	-	5/25/2005	030

5/25/2005 	1930	- 	5/26/2005 	100

5/26/2005 	2030 	- 	5/27/2005	100

5/27/2005	2130	-	5/28/2005	230

June

Dissolved oxygen percent and mg/L were deleted due to bad post-calibration of sensor:

6/8/05		2130	-	6/23/05	1330

All data were deleted due to depth values indicating sonde was out of water:

6/6/2005	1900

6/19/2005	1600	-	6/19/2005	2030

6/20/2005	1700	-	6/20/2005	2000

6/21/2005	1800	-	6/21/2005	2130

6/22/2005	1900	-	6/22/2005	2230

6/24/2005	2100	-	6/25/2005	030

6/25/2005	2230	-	6/26/2005	030

July

Dissolved oxygen percent and mg/L were deleted due to bad post-calibration of sensor:

7/11/05	1600	-	7/15/05	1030

7/31/05	730	-	8/9/05		830

pH data were deleted due to anomalous readings and bad post-calibration of sensor:

7/31/05	0730	-	8/9/05		830

Turbidity deleted due to malfunctioning of sensor:

7/12/05	2030	-	7/15/05	0000

7/23/05	930	-	7/27/05	1300

All data were deleted due to depth values indicating sonde was out of water:

7/1/2005	1630	-	7/1/2005	2000

7/2/2005	1500	-	7/2/2005	2200

7/3/2005	1500	-	7/3/2005	2200

7/4/2005	1530	-	7/4/2005	2200

7/6/2005	2030	-	7/6/2005	2300

7/21/2005	2130	-	7/21/2005	2200

August

Turbidity deleted due to sensor out of range:

8/2/05		2330

8/3/05		600

8/28/05	1430

Turbidity deleted due to malfunctioning of sensor:

8/4/05		0000	-	8/4/05		630

8/21/05	2230	-	8/25/05	1330

Dissolved oxygen percent and mg/L were deleted due to bad post-calibration of sensor:

8/23/05	300	-	8/25/05	1330

September

Turbidity deleted due to malfunctioning of sensor:

9/3/05		200 	- 	9/3/05 		500

9/5/05		930	-	9/15/05	1100

9/27/05	2030	-	10/7/05	1300

October

Turbidity deleted due to malfunctioning of sensor:

10/18/05	800	-	10/18/05	2300

10/26/05	1130	-	10/26/05	1230

All data were deleted due to depth values indicating sonde was out of water:

10/24/2005 	1130 	- 	10/24/2005	1700

10/25/2005 	1130 	- 	10/25/2005	1700

November

Turbidity deleted due to sensor out of range:

11/13/05	1530

11/15/05	1430

Turbidity deleted due to malfunctioning of sensor:

11/21/05	0930	-	12/21/05	1430

All data were deleted due to depth values indicating sonde was out of water:

11/16/2005 	700 	-	11/16/2005 	1030

11/17/2005 	430 	-	11/17/2005 	1230

11/18/2005 	600 	-	11/18/2005 	1230

11/21/2005 	1100 	-	11/21/2005 	1800

11/22/2005 	800 	-	11/22/2005 	1600

11/30/2005 	330 	-	11/30/2005 	730

December

All data were deleted due to depth values indicating sonde was out of water:

12/1/2005	500	 -	12/1/2005	900

12/2/2005	500	 -	12/2/2005	1100

12/3/2005	830

12/5/2005	1100	 -	12/5/2005	1300

12/6/2005	1330

12/13/2005	330	 -	12/13/2005	830

12/15/2005	900	 -	12/15/2005	1030

12/16/2005	700	 -	12/16/2005	1030

12/17/2005	800

12/19/2005	730	 -	12/19/2005	1200

12/20/2005	800	 -	12/20/2005	1230

12/27/2005	200	 -	12/27/2005	730

12/28/2005	530	 -	12/28/2005	630

12/29/2005	430	 -	12/29/2005	930

12/30/2005	500	 -	12/30/2005	1000

12/31/2005	700	 -	12/31/2005	930

13) Missing data:

Data are missing due to equipment failure, specific probes not being deployed, and time of maintenance or calibration of equipment.  For more details on deleted data, see the Deleted Data Section (12.).  If additional information on missing data is needed, contact the Research Coordinator at the reserve submitting the data.

Bayou Heron

All data are missing due to a programming error pre-deployment.

2/16/05	1500	-	3/17/05	930

All data are missing due to alterations to the deployment system.

8/18/05	830	-	8/18/05	900

Bayou Cumbest

All data are missing due to insufficient voltage during deployment.

2/17/05 	2330	-	3/17/05	1030

All data are missing due to alterations to the deployment system.

8/8/05		930	-	8/11/05	1030

Bangs Lake

All data are missing due to maintenance and calibration of sonde.

6/23/05	1400	-	6/24/05	1400

All data are missing due to alterations to the deployment system.

8/9/05		900	-	8/11/05	1100

Point Aux Chenes

None???

14) Post deployment information:

End of deployment Post-calibration readings in standard solutions prior to probe cleaning.

The following variable codes are used to explain missing post-calibration readings:

PA	probe absent/not functioning

NA	reading not available

Bayou Heron

| Date | DO% Air Sat | pH (7.00) | pH (10.00) | Turbidity 0 | Turbidity 123 | Sp Cond 50 |
| --- | --- | --- | --- | --- | --- | --- |
| 1/19/05 | 68.2 | 7.02 | 10.00 | 0.0 | 129.8 | 49.38 |
| 2/17/01 | 99.4 | 7.16 | 10.13 | 0.7 | 123.7 | 49.48 |
| 3/17/05 | NA | 7.19 | 10.01 | 12.2 | 46.9 | 48.97 |
| 4/15/05 | 98.2 | 7.20 | 10.24 | 0.0 | 121.8 | 49.68 |
| 5/03/05 | 88.0 | 7.07 | 9.94 | 0.3 | 122.7 | 50.29 |
| 5/18/05 | 98.8 | 6.99 | 10.00 | -0.8 | 108.9 | 50.80 |
| 6/24/05 | 47.4 | 7.07 | 9.95 | 0.7 | 108.8 | 51.24 |
| 7/18/05 | 101.0 | 7.00 | 9.98 | -0.9 | 124.0 | 49.75 |
| 8/12/05 | 91.7 | 6.96 | 9.67 | 1.0 | 115.8 | 50.04 |
| 8/29/05 | 99.4 | NA | NA | NA | NA | NA |
| 9/20/05 | 98.6 | 7.06 | 10.02 | 0.5 | 122.4 | 49.53 |
| 10/10/05 | 96.2 | 6.57 | 9.45 | -1.6 | 122.0 | 50.75 |
| 10/31/05 | 100.8 | 6.71 | 9.68 | 2.3 | 45.5 | 49.92 |
| 11/23/05 | 100.5 | 7.01 | 10.01 | -0.3 | 129.5 | 50.27 |
| 12/22/05 | 101.7 | 7.15 | 10.24 | -0.2 | 117.6 | 49.97 |
| 1/11/06 | 99.7 | 7.03 | 9.82 | 0.2 | 123.5 | 51.25 |

Crooked Bayou

| Date | DO% Air Sat | pH (7.00) | pH (10.00) | Turbidity 0 | Turbidity 123 | Sp Cond 50 |
| --- | --- | --- | --- | --- | --- | --- |
| 1/07/05 | 101.6 | 7.14 | 10.12 | 0.8 | 126.9 | 49.55 |
| 2/17/01 | 96.2 | 7.02 | 9.91 | 0.5 | 125.9 | 48.75 |
| 3/17/05 | 101.9 | 7.16 | 10.10 | 0.3 | 125.1 | 49.32 |
| 4/15/05 | 101.5 | 6.94 | 9.89 | 0.0 | 125.8 | 49.83 |
| 5/03/05 | 109.1 | 7.00 | 10.17 | -0.9 | 121.9 | 50.53 |
| 5/18/05 | 109.7 | 7.11 | 10.06 | -2.0 | 121.3 | 49.97 |
| 6/24/05 | 29.1 | 7.14 | 9.69 | -2.0 | 115.0 | 47.12 |
| 7/18/05 | 73.4 | 7.01 | 10.06 | 0.6 | 124.7 | 49.70 |
| 8/12/05 | 109.6 | PA | PA | 1.4 | 119.8 | 50.25 |

Point Aux Chenes

| Date | DO% Air Sat | pH (7.00) | pH (10.00) | Turbidity 0 | Turbidity 123 | Sp Cond 50 |
| --- | --- | --- | --- | --- | --- | --- |
| 8/29/05 | NA | NA | NA | NA | NA | NA |
| 10/02/05 | 87.9 | 7.09 | 9.40 | -0.8 | 114.6 | 42.88 |
| 10/10/05 | 73.7 | 6.31 | 9.21 | -6.4 | 118.1 | 49.27 |
| 10/31/05 | 65.7 | 6.85 | 9.99 | 0.2 | 115.4 | 49.37 |
| 11/29/05 | 94.0 | 6.90 | 9.81 | 0.5 | 123.7 | 49.94 |
| 12/22/05 | 103.5 | 6.98 | 10.03 | -0.4 | 124.7 | 51.38 |
| 1/11/06 | 106.2 | 7.14 | 10.11 | -0.9 | 125.7 | 50.24 |

Bayou Cumbest

| Date | DO% Air Sat | pH (7.00) | pH (10.00) | Turbidity 0 | Turbidity 123 | Sp Cond 50 |
| --- | --- | --- | --- | --- | --- | --- |
| 1/7/05 | 107.4 | 7.19 | 10.20 | 0.6 | 127.0 | 49.57 |
| 2/17/05 | 112.0 | 7.04 | 9.95 | -0.4 | 123.9 | 49.06 |
| 3/17/05 | NA | 7.13 | 10.11 | -0.6 | 123.5 | 50.48 |
| 4/15/05 | 122.5 | 7.11 | 9.99 | -0.7 | 122.5 | 50.18 |
| 5/03/05 | 105.5 | 7.16 | 10.04 | -1.5 | 121.7 | 49.88 |
| 5/18/05 | 90.1 | 7.10 | 10.00 | -0.8 | 121.1 | 50.32 |
| 6/24/05 | 30.3 | 7.17 | 9.81 | 1.2 | 114.9 | 45.29 |
| 7/18/05 | 97.0 | 7.00 | 9.87 | 1.5 | 125.5 | 48.93 |
| 8/12/05 | 149.0 | 7.25 | 10.33 | 0.0 | 119.9 | 50.08 |
| 8/29/05 | NA | NA | NA | NA | NA | NA |
| 9/20/05 | 101.6 | 6.98 | 10.00 | -4.3 | 120.4 | 50.30 |
| 10/10/05 | 109.8 | 6.33 | 9.18 | -6.5 | 118.5 | 50.15 |
| 10/31/05 | 115.2 | 6.75 | 9.82 | 0.4 | 119.7 | 50.50 |
| 11/23/05 | 106.4 | 6.83 | 9.83 | -0.1 | 122.4 | 50.22 |
| 12/22/05 | 117.1 | 7.21 | 10.22 | 0.0 | 120.6 | 50.94 |
| 1/11/06 | 101.9 | 7.12 | 10.11 | -0.5 | 126.4 | 50.51 |

Bangs Lake

| Date | DO% Air Sat | pH (7.00) | pH (10.00) | Turbidity 0 | Turbidity 123 | Sp Cond 50 |
| --- | --- | --- | --- | --- | --- | --- |
| 1/7/05 | 107.3 | 7.06 | 10.31 | -2.0 | 128.1 | 49.97 |
| 2/17/05 | 109.2 | 6.98 | 9.86 | -0.9 | 124.3 | 49.11 |
| 3/17/05 | 115.2 | 7.02 | 9.96 | -1.5 | 126.1 | 49.30 |
| 4/15/05 | 109.3 | 7.04 | 9.99 | -3.8 | 124.8 | 49.61 |
| 5/03/05 | 113.4 | 7.08 | 9.87 | 0.4 | 122.5 | 48.94 |
| 5/18/05 | 76.1 | 6.97 | 9.89 | -1.6 | 120.3 | 48.93 |
| 6/23/05 | 61.5 | 7.76 | 8.95 | -0.3 | 123.2 | 41.88 |
| 7/18/05 | 85.4 | 7.11 | 9.95 | 0.7 | 122.1 | 40.84 |
| 8/12/05 | 35.8 | 7.50 | 7.63 | 1.0 | 125.4 | 48.06 |
| 8/29/05 | NA | NA | NA | NA | NA | NA |
| 9/20/05 | 96.7 | 6.85 | 9.96 | -1.0 | 122.9 | 50.73 |
| 10/10/05 | 110.5 | 5.33 | 8.67 | -5.8 | 115.5 | 50.38 |
| 10/31/05 | 112.5 | 6.85 | 9.80 | 0.4 | 123.9 | 50.46 |
| 11/23/05 | 109.6 | 7.22 | 9.96 | 0.0 | 115.0 | 50.36 |
| 12/22/05 | 109.7 | 7.03 | 10.10 | -2.6 | 121.4 | 50.25 |
| 1/11/06 | 103.3 | 7.39 | 10.29 | -0.4 | 117.4 | 51.01 |

15)  Other remarks/notes:

On 10/01/2019 this dataset was updated to include embedded QAQC flags for anomalous/suspect data.  System-wide monitoring data beginning in 2007 were processed to allow for QAQC flags and codes to be embedded in the data files rather than detailed in the metadata alone (as in the anomalous/suspect, deleted, and missing data sections above).  Prior to 2006, rejected data were deleted from the dataset so they are unavailable to be used at all, but suspect data were only noted in the metadata document.  Suspect data flags <1> were embedded retroactively in order to allow suspect data to be easily identified and filtered from the dataset if desired for analysis and reporting purposes.  No other flags or codes were embedded in the dataset and users should still refer to the detailed explanations above for more information.

All data collected on or around the following dates were subject to tropical storm and/or hurricane conditions at the Grand Bay NERR.

7/05/05 – 7/06/05 	Tropical Storm Cindy		Sustained winds in excess of 55mph

7/10/05 – 7/11/05	Hurricane Dennis 		Category 3, light wind and rain

8/28/05 – 8/29/05	Hurricane Katrina 		Category 5, storm surge 15-18ft and

100 mph winds

9/22/05 – 9/24/05	Hurricane Rita			Category 5, storm surge 3-5ft

During August, the deployment sleeves located at Bayou Heron, Bayou Cumbest, and Bangs Lake were lowered .25m in order to collect as much data as possible during low tide events, see Missing Data (13.) for the exact dates and times that these changes occurred.  Also at this time, the Crooked Bayou site was abandoned (08/9/05 0700) and the Point Aux Chenes site designated (08/11/05 1030) as part of this effort to capture more low tide data.

All readings collected from Bangs Lake on or around 04/15/05 – 04/19/05 may be out of the normal range for the site due to highly acidic runoff escaping a breached containment system owned and operated by a nearby fertilizer manufacturer.