South Slough (SOS) NERR Meteorological Metadata

January – December 2008

Latest Update: 8/29/2012

I. Data Set and Research Descriptors

- Principal investigators and contact persons:

### Address: 	South Slough National Estuarine Research Reserve

### PO Box 5417

### Charleston, OR 97420

### Phone:  (541) 888-5558

### Fax:      (541) 888-2733

Contact Persons:

###### Dr. Steven Rumrill, Research Program Coordinator

### E-mail:  steve.rumrill@state.or.us; (541) 888-2581 ext. 302

Alicia Helms, Estuarine Monitoring Coordinator

E-mail: alicia.r.helms@state.or.us; (541) 888-2581 ext 304

Adam DeMarzo, Estuarine Monitoring Assistant

E-mail: adam.demarzo@state.or.us (541) 888-2581 ext 305

- Entry verification

Data are uploaded from the CR1000 data logger to a Personal Computer (IBM compatible).  Files are exported from LoggerNet v. 3.2.2 in a comma-delimited format and uploaded to the CDMO where they undergo automated primary QAQC and become part of the CDMO’s online provisional database.  During primary QAQC, data are flagged if they are missing or out of sensor range.  The edited file is then returned to the Reserve where it is opened in Microsoft Excel and processed using the CDMO’s NERRQAQC Excel macro.  The macro inserts station codes, creates metadata worksheets for flagged data, and graphs the data for review.  It allows the user to apply QAQC flags and codes to the data, append files, and export the resulting data file to the CDMO for tertiary QAQC and assimilation into the CDMO’s authoritative online database.  For more information on QAQC flags and QAQC codes, see Sections 11 and 12.

The meteorological information from each sensor is sampled every 5 seconds and stored on the Campbell Scientific, Inc. (CSI) CR1000 datalogger (installed Feb 5, 2007).  The station was hardwired into the Estuarine and Coastal Science laboratory IBM-compatible PC originally using CSI’s Real-Time Data Management Program December 13, 2002; however, on Oct 6, 2006, the hardwire connection was removed due to conflict with server usage for a new employee.  On February 5, 2007, a CR1000 datalogger was installed to replace the CR10x.  The CR1000 interfaces with the LoggerNet 3.2.2 software supplied by Campbell Scientific.  In June 2007, the NL100 Lan Link network connection was re-established using LoggerNet 3.2; this allowed the weather data to be automatically downloaded to a real-time file on a personal computer; however, this link was not maintained due to possible firewall security issues with the Department of State Lands and/or potential equipment problems.  Beginning Jan 2008, a laptop computer is connected directly to the CR1000 datalogger via an RS232 connection to download data monthly.    A Sutron Sat-Link2 transmitter was installed at this station on 11/17/05 and transmits data to the NOAA GOES satellite, NESDIS ID #3B0032E0.  The transmissions are scheduled hourly and contain four (4) datasets reflecting fifteen minute data sampling intervals.

Steve Rumrill has been managing all aspects of the weather component of SWMP since installation in August 2001.  Sue Powell was responsible for the weather station from August 2001 through December 16, 2005.  Ali Helms has been responsible for the weather station since July 2006, including data management with Adam DeMarzo assisting with management of the weather station since October 2006.

- Research objectives:

The South Slough NERR operates the SWMP Meteorological Station to provide continuous measurements of local weather conditions affecting the South Slough estuary.  The nearest National Weather Service certified station is located at the North Bend Airport, 11.5 km NE from the mouth of South Slough.  At that location, the weather observers’ primary function is to provide current meteorological information to airplane pilots.  Only 24-hour weather data is recorded, including maximum and minimum temperature, precipitation, and prevailing wind direction.

The SWMP weather station is located on the campus of the University of Oregon’s Institute of Marine Biology campus near the mouth of South Slough; recorded data provides information more pertinent to the microclimates affecting the estuary.  The principal objective is to record long-term meteorological data that is particularly relevant to the South Slough estuary in order to gain a better understanding of climatic forcing and to observe any environmental changes or trends over time.  In addition, the station provides real-time data to assess the short-term effects of local weather on measurements of water parameters within the estuary.  The station is part of a framework to evaluate the relative contributions of oceanic forcing and watershed inputs to hydrodynamics, nutrient dynamics, and other ecological processes within the South Slough estuary.  Data can be used as 1) a vital reference of atmospheric data for various research projects at the Reserve – an integral part of the NERR mission to provide a platform for estuarine research, 2) for meteorological context or atmospheric forcing for 15 minute water quality data, and other long-term environmental monitoring paradigms at the Reserve

- Research methods:

Sutron data telemetry equipment was installed at the Charleston Meterological Station on Nov. 17, 2005 and transmits data to the NOAA GOES satellite, NESDIS ID #3B0032E0.  (Where #3B0032E0 is the GOES ID for that particular station).  The transmissions are scheduled hourly and contain four (4) data sets reflecting fifteen minute data sampling intervals.  Upon receipt by the CDMO, the data undergoes the same automated primary QAQC process detailed in Section 2 above.  The “real-time” telemetry data become part of the provisional dataset until undergoing secondary and tertiary QAQC and assimilation in the CDMO’s authoritative online database.  Provisional and authoritative data are available at http://cdmo.baruch.sc.edu.

The 15 minute Data are collected in the following formats for the CR1000:

Averages from 5-second data:

Air temperature (ºC), Relative Humidity (%), Barometric pressure (mb), Wind Speed (m/s), Wind Direction (degrees),  Wind Direction Standard Deviation (degrees) (began collecting as a standard parameter on 9/30/2008), Battery Voltage (volts)

Maximum, Minimum, and their times for 5-second data:

Air Temperature (ºC) (these data are not included in the dataset but are available from the SOS Reserve)

Max and time for 5-second data:

Wind Speed (m/s), from 5-second data (began collecting as a standard parameter on 6/12/2008)

Totals:

Precipitation (mm), PAR (millimoles/m2), and Cumulative Precipitation (mm) (began collecting as a standard parameter on 4/3/2008)

On 17 Nov 2005, new satellite telemetry equipment was installed on the South Slough NERR SWMP meteorological station.  The new equipment will use the Geostationary Operational Environmental Satellite (GOES) transmission system to send data in near real-time to the NERRS Centralized Data Management Office.  The new equipment includes a GOES transmission antenna, a Global Position System (GPS) receiver, a 20 watt solar panel and charge controller, a 12 V DC power supply, and a new weatherproof datalogger enclosure.

The meteorological sensors are inspected monthly for damage or fouling and the solar panels, PAR sensor, and the rain gauge are cleaned if necessary.  A Kestrel 4000 handheld and a sling psychrometer are used as a general check of the sensors. In the laboratory, data recorded as close as possible to the field check are downloaded from the nearby North Bend Weather Station and/or the NOAA Charleston NWLON station and compared to South Slough station data. Annually, all sensors were field tested and calibrated on site by Robb Hibbs of MeasureTek, Inc. for drift from acceptable sensor ranges through 2005 (1/14/02, 12/13/02, 10/20/03, and 9/15/05.  Beginning in 2006, spare sensors from CDMO are loaned to the reserve and swapped out with the South Slough sensors while they are sent to individual manufactures for re-calibration. Data recorded during station maintenance are rejected.

The TE525WS-L Texas Electronics 8' Tipping Bucket Rain Gage (0.01" tip) from Campbell Scientific, Inc. was installed on 4/30/2007.

The recommended sensor calibration intervals are:

- Temperature/Humidity- annual recalibration

- Rain Gauge- annual recalibration

- Wind Speed/Direction- bi-annual recalibration

- Barometric Pressure- bi-annual recalibration

- PAR- bi-annual recalibration

- Site location and character:

South Slough is characterized as a drowned river-mouth estuary.  Located 8.5 km southwest of Coos Bay, Oregon, South Slough is one of seven tidal inlets that collectively form the Coos estuary.  The South Slough National Estuarine Research Reserve is representative of the mid-Pacific portion of the Lower Columbian Bioregion, extending from Cape Mendocino, CA north to the Columbia River, OR.  The Reserve includes 4,700 acres with 600 acres of tidal marshes, mudflats and open water channels.

Within the South Slough, there are three primary sub-systems that provide much of the freshwater input to the area.  Only two sub-systems, the Sengstacken and Winchester arms, are within Reserve boundaries.  The Sengstacken arm drains a semi-pristine area with no development and serves as a reference site for the riverine portion of the estuary.  The Winchester arm is similarly located in an area with development limited to one small farm (<10 cattle) high in the watershed.  The area functions as a management-treatment site for salt marsh restoration activities in the Winchester Tidelands Restoration Project.  Valino Island marks the northern boundary of the Reserve and is a marine-dominated area.

Charleston Meteorological Station:  [43 20’ 42.18 N, 124 19’ 43.44 W]

The station is located 1.2 km northwest of the Charleston water quality monitoring site.  A 10-meter tall navigation tower that formerly belonged to the Corps of Engineers is situated on the northeast corner of the Oregon Institute of Marine Biology campus.  The tower is about 20 meters from the inner boat marina at Charleston harbor.  Three 12-13 meter tall shore pines lie directly south of the tower and may slightly shelter the station from southerly winds, but the sensor array is never shaded.  The weather station is above the nearest building, which is about 8 meters away.  A 40-meter-tall forested ridge that has a northeast to southwest alignment is located about 60 meters from the weather tower and may interfere with winds from the southwest.  On the tower platform, a 3-meter tall mast holds the Wind Monitor, Temperature and Relative Humidity, Barometric Pressure, and PAR sensors array and the CR-1000.  Sensor heights from the ground are Rain 8.73 m, Barometric Pressure 9.18 m, Relative Humidity/Temp 9.68 m, Wind 11.08 m, and PAR 11.08 m. Because there is no open field available, the Rain Gauge is set on top of a platform on the southwest corner of the tower.  Other sensor placement is in accordance with the CDMO Manual (Version 4.0, Part D, Section 2.0).

The CR1000 is powered by an internal battery.  Two solar power panels (40W and 20W) generate power during the daylight hours to maintain battery charge. The panels also generate power to the Optimus 12V deep-cycle sealed gel battery (installed on 11/17/05) during the day. The Optimus provides power during the night.

Safety apparatus for ascending the tower is in accordance with OSHA regulations.  A sliding trolley runs in a track during ascension of the ladder.  The trolley is connected by a carabiner to a chest harness and belt.  While working on the tower platform, a 4-meter-long tether secures the person to a tower strut.

- Data collection period:

The Charleston station was installed on 08/21/01.  The station was hard-wired to the SWMP laboratory in 2002 and the data were downloaded automatically.  Instruments were installed on 11/17/05 to telemeter the meteorological data to NOAA GOES satellite. Data were collected for the entire year in 2008 from 01/01/2008 00:00 to 12/31/2008 23:45.

- Distribution

According to the Ocean and Coastal Resource Management Data Dissemination Policy for the NERRS System-wide Monitoring Program:  NOAA/ERD retains the right to analyze, synthesize and publish summaries of the NERRS System-wide Monitoring Program data.  The PI retains the right to be fully credited for having collected and processed the data.  Following academic courtesy standards, the PI and NERR site where the data were collected will be contacted and fully acknowledged in any subsequent publications in which any part of the data are used.  Manuscripts resulting from the NOAA/OCRM-supported research that are produced for publication in open literature, including refereed scientific journals, will acknowledge that the research was conducted under an award from the Estuarine Reserves Division, Office of Ocean and Coastal Resource Management, National Ocean Service, National Oceanic and Atmospheric Administration.  The data set enclosed within this package/transmission is only as good as the quality assurance/quality control procedures outlined by the enclosed metadata reporting statement.  The user bears all responsibility for its subsequent use/misuse in any further analyses or comparisons.  The Federal government does not assume liability to the Recipient or third persons, nor will the Federal government reimburse or indemnify the Recipient for its liability due to any losses resulting in any way from the use of this data.

NERR weather data and metadata can be obtained from the Research Coordinator at the individual NERR site (please see Section 1 - Principal investigators and contact persons), from the Data Manager at the Centralized Data Management Office (please see personnel directory under the general information link on the CDMO home page) and online at the CDMO home page http://cdmo.baruch.sc.edu/.  Data are available in comma separated format.

- Associated researchers and projects:

The South Slough NERR maintains four water quality monitoring stations along the estuarine gradient within the South Slough estuary.  These stations collect continuous information to address three primary topics:  1) short-term variability and long-term changes in estuarine water parameters within different regions of the estuary (i.e., localized impacts of seasonal storm events, variability in water quality parameters in association with re-establishment of tidal circulation, interannual differences in rainfall, magnitude and influence of El Niño – La Niña events, spatial extent of oceanic and tidal forcing, and long-term changes associated with the Pacific Decadal Oscillation, etc.); 2) verification, calibration, and future refinement of an existing two-dimensional water quality model (CE-QUAL-W2) developed to predict levels of coliform bacteria and dispersion of environmental toxins; and 3) collection of fundamental baseline data for future development of a three-dimensional tidal hydrodynamic model for the greater South Slough/Coos Bay estuaries.

A water chemistry component is also associated with the core monitoring program.  Variables measured include: nitrogen, phosphorus, chlorophyll, dissolved organic carbon, total suspended solids, and silica.  Grab samples are collected for both high and low tides at the four water quality sampling sites, and automated samplers collect diel samples to compare marine and riverine sites during wet and dry seasons.

The SSNERR water quality monitoring project has been incorporated into the State of Oregon Coastal Non-Point Pollution Control Program (S. 6217) developed by the Oregon Department of Land Conservation and Development and the Oregon Department of Environmental Quality in July 1995.

#### II. Physical Structure Descriptors

- Sensor specifications, operating range, accuracy, date of last calibration

Parameter:			Barometric Pressure

Units:				millibars (mb)

Sensor Type:			Vaisala Analog Barocap silicon capacitive pressure sensor

Model #:			CS 015

Operating Range:		Pressure: 600-1060 mbar; Temperature: -40-+60°C

Humidity:  			non-condensing

Accuracy: 	±0.5 mb at 20°C; ± 2 mb 0°C-40°C; ± 4 mb -20°C to 45°C; ± 6 mb -40 – 60°C

Stability: 			± 0.1 mb per year

Date of Last calibration:	12/12/95 (s/n P5020003), 5/11/01 (P5020003), 8/30/02 (X3540012), 10/14/02 (X4150040), 12/13/02, 10/20/03 (X3540012), 09/15/2004, 09/15/05 (X3540012), 09/04/2007 (X3540012)

Calibration Drift:		0.0%

Parameter:			Temperature

Units:				Celsius

Sensor Type:			Platinum resistance temperature detector (PRT)

Model:	HMP45C, changed to HMP35C on 11/17/05 and returned to manufacturer for repair

Operating Temperature:	-40 - +60°C

Range:				-40 - +60°C

Accuracy: 			± 0.2 °C @ 20°C

Date of Last calibration:	4/11/01 (s/n R2720011), 12/13/02 (HMP35C), 10/20/03 (s/n 4150040), 09/15/04, 09/15/05 (s/n 4150040), 11/22/05 (s/n 467113), 8/28/07 (s/n 467113), 4/26/07 (s/n R2720011),

Parameter:			Relative Humidity

Units:				Percent

Sensor Type:			Vaisala HUMICAP 180 capacitive relative humidity sensor

Model:	HMP45C, changed to HMP35C on 11/17/05 and returned to manufacturer for repair

Range:				0 - 100% non-condensing

Accuracy:  			± 2% RH (0-90%) and ± 3% (90-100%) at 20oC

Temperature dependence

of RH measurement:		± 0.05% HR/°C

Date of Last calibration:	4/11/01 (s/n R2720011), 12/13/02 (HMP35C), 10/20/03 (s/n 4150040), 09/15/04, 09/15/05 (s/n 4150040), 11/22/05 (s/n 467113), 8/28/07 (s/n 467113), 4/26/07 (s/n R2720011)

Parameter:			LiCor Quantum Sensor

Units:				mmoles m-2 (total flux)

Sensor Type:			High stability silicon photovoltaic detector (blue enhanced)

Model:				LI190SB

Light spectrum wavelength:	400 to 700 nm

Temperature Dependence:	0.15% per °C maximum

Stability:			<±2% change over 1 yr

Operating Temperature:	-40 to 65°C; humidity: 0 to 100%

Sensitivity:			typically 5 μA per 1000μmoles s-1 m-2

Date of Last calibration:	5/03/01 (Q22455), 08/07/02 (Q31375), 12/13/02, 10/20/03, 09/15/04, 09/15/05 (Q31375), 10/01/2007 (Q31375)

Multiplier:			1.26

Parameter:			Wind Speed

Units:				meters per second (m/s)

Sensor Type:	18 cm diameter 4-bladed helicoids propeller molded of polypropylene

Model:				RM Young 05106 Wind Monitor

Range:				0-60 m/s (130 mph); gust survival 100 m/s (220 mph)

Accuracy:			± 2%

Date of last calibration: 	11/22/02 (WM54625) installed new 12/13/02, 10/20/03 (5106), 09/15/04, 09/15/05, 9/20/2007

Parameter:			Wind Direction

Units:				degrees

Sensor Type:	balanced vane, 38 cm turning radius

Model:				RM Young 05106 Wind Monitor

Range:				360° mechanical, 355° electrical (5° open)

Accuracy:			± 5%

Date of last calibration: 	11/22/02 (WM54625) installed new 12/13/02, 10/20/03 (5106), 09/15/04, 09/15/05, 09/20/2007

Parameter:			Precipitation

Units:				milliliters

Sensor Type:			Tipping Bucket Rain Gauge

Model:				TE525WS (8" tipping bucket)

Range:				0.01 in per bucket tip

Operating Range:		0° to ± 50°C; Humidity: 0 to 100%

Accuracy:			±1.0% at 14”/hr

Date of Last calibration:	12/13/02, 10/20/03 (952183), 9/15/05 (952183), 5/01/2007, 3/5/2008

Campbell Scientific CR1000 Wiring Panel has two MB Flash EEPROM that is used to store the operating system and another 128 K of flash is used to store configuration settings.  A minimum of 2 MB SRAM (4 MB optional) is available for program storage (16K), operating system use, and data storage.

CR1000 installed:  Feb. 5, 2007

- Coded variable definitions:

Sampling Station		Sampling Site Code		Station Code

Charleston Meteorological		CM			soscmmet

11) QAQC flag definitions – This section details the automated primary and secondary QAQC flag definitions.

QAQC flags provide documentation of the data and are applied to individual data points by insertion into the parameter’s associated flag column (header preceded by an F_).   During primary automated QAQC (performed by the CDMO), -5, -4, -2flags are applied automatically to indicate data that is above or below sensor range or missing.  All remaining data are then flagged 0, as passing initial QAQC checks.   During secondary and tertiary QAQC 1, -3, and 5 flags may be used to note data as suspect, rejected due to QAQC, or corrected.

-5	Outside High Sensor Range

-4	Outside Low Sensor Range

-3	Data Rejected due to QAQC

-2	Missing Data

-1	Optional SWMP supported parameter

0		Passed initial QAQC checks

1		Suspect Data

2		Open - reserved for later flag mean

3		Open - reserved for later flag

4		Historical Data:  Pre-Auto QAQC

5		Corrected Data

12)  QAQC code definitions – This section details the secondary QAQC Code definitions used in combination with the QAQC flags above.

QAQC codes are used in conjunction with QAQC flags to provide further documentation of the data and are also applied by insertion into the associated flag column.  There are three (3) different code categories, general, sensor, and comment.  General errors document general problems with the CR1000 datalogger, sensor errors are sensor specific, and comment codes are used to further document conditions or a problem with the data.  Only one general or sensor error and one comment code can be applied to a particular data point.

General Errors

GIM	Instrument Malfunction

GIT	Instrument Recording Error, Recovered Telemetry Data

GMC	No Instrument Deployed due to Maintenance/Calibration

GMT 	Instrument Maintenance

GPD	Power Down

GPF	Power Failure / Low Battery

GPR	Program Reload

GQR	Data Rejected Due to QA/QC Checks

GSM	See Metadata

Sensor Errors

SIC		Incorrect Calibration Constant, Multiplier or Offset

SSN		Not a Number / Unknown Value

SNV		Negative Value

SOC		Out of Calibration

SSM		Sensor Malfunction

SSR		Sensor Removed

Comments

CAF	Acceptable Calibration/Accuracy Error of Sensor

CDF	Data Appear to Fit Conditions

CRE	Significant Rain Event

CSM       See Metadata

CVT	Possible Vandalism/Tampering

13) Other remarks/notes

Data are missing due to equipment or associated specific sensors not being deployed, equipment failure, time of maintenance or calibration of equipment, or repair/replacement of a sampling station platform.  Any NANs in the dataset stand for “not a number” and are the result of low power, disconnected wires, or out of range readings.  If additional information on missing data is needed, contact the Research Coordinator at the reserve submitting the data.

Small negative PAR values are within range of the sensor and are due to normal errors in the sensor and the CR1000 Datalogger. The Maximum signal noise error for the Licor sensor is +/- 2.214 mmoles/m2 over a 15 minute interval.

Relative Humidity data greater than 100 are within range of the sensor accuracy of +/-3%.

Cumulative precipitation data are recorded from 00:00 to 23:59 with the daily total recorded at the midnight mark (00:00). The midnight CumPrcp value is actually the total from the previous day.

All Reserves were required to align their wind sensors to True North by April 1, 2008.  SOS NERR’s sensor was aligned on April 3, 2008 at 15:45.

Elevated nighttime PAR values for 2008 greater than 2.214 mmoles/m2 were rejected and flagged <-3> (CSM); these elevated readings/offsets occurring at night are most likely due to slow moisture intrusion into the body of the sensor over time.  As a result, elevated daytime PAR events occur as well; therefore, some daytime PAR values are marked as suspect <1> CSM (3/6/2008, 3/7/2008, and 3/12/2008) .  Elevated nighttime data that were less than 2.214mmoles.m2 were flagged as suspect.  The sensor (Q31375) was calibrated 11/23/09 and the photodiode had water damage inside and the cable was corroded/water damage.  A new PAR sensor (Q42152) was installed 2/18/2010. Par data from 3/17 22:00 – 12/31 23:45 were rejected due to this damaged photodiode on the sensor.

Data from 1/2/08 6:30 – 1/7/08 10:30 & 1/7/08 14:45 – 1/15/08 12:45 are missing due to battery failure.  There was not enough sunlight to keep the battery charged.  Because of the missing data, 1/7/08 10:45 and 1/15/08 13:00 data were rejected since they are probably not a full 15 minutes of 5-second data.  A 40 W solar panel was reinstalled to prevent loss of data due to battery failure.  On 1/2/08, the power cord was replaced Part # 18133 TX312 Power Cable, 2 ft. but, the station was still not transmitting.  A temporary marine deep cycle battery Orbital exide select (ORB34DC-36) 14.4V was placed temporarily at the weather station to bring power back while the original battery was charged.

Data from 1/15/08 15:15 – 1/22/08 15:15 were retrieved from the real-time data file (downloaded from CDMO) and coded with [GIT]; however, data records are missing from 1/21/08 9:45 – 10:30 (was not transmitted by satellite).    Data was not available from CR1000 because we were relying on NL100 lan link to lab to retrieve data and data was not being transmitted all the time; also, the LoggerNet program was supposed to stay running and open and computer user was not aware of this.

The rain gauge was calibrated in house using the static method on 3/5/08 beginning around 14:22.  The calibration affected precipitation data from 3/5/08 14:30-15:45.

We reoriented the wind sensor to True North on 4/3/08 15:45; therefore, there are missing  and rejected data 4/3/08 15:45-16:15.  Wind direction data were flagged as suspect from 4/1/2008 00:00 through 4/3/2008 15:30 since the wind sensor was not aligned to True North by the required date, April 1, 2008.  Uploaded new met program soscmmet_v5.0_030708.CR1 on 4/3/08 16:15 (added cumulative precipitation).  Data at 16:15 were rejected because it was not a full 15 minutes of 5-second data.  Also, uploaded new met program soscmmet_V5.5_040408.CR1 on 5/13/08 16:00 because version 5.0 upload missed coordination with HADS and CDMO.

On 4/3/08 16:15, we began collecting cumulative precipitation parameter.

On 6/12/08 10:30, we began submitting Max Wind Speed as a standard parameter.

On 9/30/08 10:00, we began submitting Wind Direction Standard Deviation as a standard parameter.

Relative humidity values on 11/13/08 04:30-09:30 are just above (101%) the high measurement specifications for the sensor but were coded as <1> [GSM] (CAF).  There were significant rain events just prior to this period (Nov 8-13, 2008).

Relative humidity data were rejected <-3> [SSM] (CSM) from 12/10/08 12:15-12/31/08 23:45 and values were low (range 50-78%; avg 72 %) when compared to relative humidity values from the North Bend Airport (values range 56-93 %, average 80%).  The sensor was calibrated (biannually) on 11/13/09 and the RH was non-functional and the hybrid circuit card was damaged (sensor was non-repairable).   The sensor was replaced on 1/11/2010 13:30 with a previously calibrated sensor (s/n R2720011).  The sensor (s/n R2720011) was calibrated on 4/26/07 and remained in the lab after calibration and was never installed in the field.  Campbell Scientific confirmed that the instrument was recalibrated with < 3 % change for maximum drift per year and they don’t differentiate max drift per year based on field vs lab; therefore, we determined that the calibration was still good.

Precipitation data recorded at the North Bend Airport compared to soscmmet weather station:

2008 Monthly Precipitation in mm

| Site | Jan | Feb | Mar | Apr | May | Jun | Jul | Aug | Sep | Oct | Nov | Dec |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| soscmmet | 150.46* | 108.97 | 147.32 | 122.17 | 10.67 | 29.46 | 2.54 | 27.43 | 3.05 | 61.90 | 155.60 | 212.50 |
| NBWS | 361.19 | 113.54 | 151.64 | 144.27 | 12.19 | 32.77 | 1.78 | 27.69 | 3.30 | 58.93 | 174.24 | 233.68 |
| Histor | 258.32 | 202.69 | 194.56 | 114.30 | 76.71 | 41.40 | 11.43 | 18.03 | 42.93 | 131.32 | 241.05 | 270.26 |

soscmmet		2008 data from SWMP Charleston Meteorological Station – located 11.5 km

SW of NBWS and 8 km SW of local weather spotter

*Jan reflects missing data from 1/2/08 6:30 – 1/7/08 10:30 & 1/7/08 14:45 – 1/15/08 12:45

NBWS		North Bend Weather Station at FAA Airport 2008 data

Histor		Mean data recorded at NBWS from 1931-2001