# ACE Basin National Estuarine Research Reserve

# Meteorological Metadata Report

January – July 2017

Latest Update: 07/31/2017 by S. Upchurch

# I.  Data Set and Research Descriptors

- Principal investigator(s) and contact persons

SCDNR/Marine Resources Division	ACE Basin NERR field station

217 Fort Johnson Road			15717 Bennett's Point Road

Charleston, SC  29422			Green Pond, SC 29441

(843) 953-9300				(843) 844-8822

Contact Persons:

Dr. Denise Sanger, Research Coordinator

# Contact: Sangerd@dnr.sc.gov  (843) 953-9074

# Saundra Upchurch, Reserve Biologist

# Contact: UpchurchS@dnr.sc.gov  (843) 953-9223

- Entry verification

# a)  Data Input Procedures:

Meteorological data is recorded in 15-minute intervals and stored in a Campbell Scientific CR1000 data logger.  ACE staff uses a data storage card to serve as a back-up for the CR1000 data logger.  The program that controls the sampling, schedule, and storage of data by the CR1000 is provided by CDMO.

The CR1000 is interfaced with the Campbell Scientific’s LoggerNet software.  ACE staff downloads and troubleshoots LoggerNet programs with a direct connection to the weather station, using a laptop computer and a 9-pin serial cable.  The GOES satellite system also uploads all 15-minute meteorological data summaries to a NOAA server every hour.  These multiple methods of data retrieval ensure that all measured parameters are reported.

# b)  QA/QC Procedures:

Data are uploaded from the CR1000 data logger to a Personal Computer (IBM compatible).  Files are exported from LoggerNet in comma-delimited format and uploaded to the CDMO where they undergo automated primary QAQC and become part of the CDMO’s online provisional database.  During primary QAQC, data are flagged if they are missing or out of sensor range.  The edited file is then returned to the Reserve where it is processed in Microsoft Excel using the CDMO’s NERRQAQC Excel macro.  The macro inserts station codes, creates metadata worksheets for flagged data and summary statistics, and graphs the data for review.  It allows the user to apply QAQC flags and codes to the data, append files, and export the resulting data file to the CDMO for tertiary QAQC and assimilation into the CDMO’s authoritative online database.  Outliers, suspect data, erroneous data, and other data flagged by the NERRQAQC macro are evaluated based on monthly field logs and NOAA historical data for the area.  For more information on QAQC flags and QAQC codes, see Sections 11 and 12.

# Amanda Fornal, Denise Sanger, and Saundra Upchurch were responsible for the above tasks during 2017 Q1. Denise Sanger and Saundra Upchurch were responsible for the above tasks during 2017 Q2.

3)  Research objectives – (Campbell Weather Station):

The principal objective of the Weather Monitoring Program is to record long-term meteorological data for the ACE Basin NERR in order to observe any environmental changes or trends over time.

4)  Research methods –

Campbell Scientific data telemetry equipment was installed at the Bennett’s Point station on 06/30/2006 and transmits data to the NOAA GOES satellite, NESDIS ID #3B01E672.  The transmissions are scheduled hourly and contain four (4) data sets reflecting fifteen minute data sampling intervals.  Upon receipt by the CDMO, the data undergoes the same automated primary QAQC process detailed in Section 2 above.  The “real-time” telemetry data become part of the provisional dataset until undergoing secondary and tertiary QAQC and assimilation in the CDMO’s authoritative online database.  Provisional and authoritative data are available at http://cdmo.baruch.sc.edu.

The 15 minute Data are collected in the following formats for the CR1000:

Averages from 5-second data:

Air Temperature (°C), Relative Humidity (%), Barometric Pressure (mb), Wind Speed (m/s), Wind Direction (degrees) and Battery Voltage (volts)

Maximum and Minimum Air Temperature (°C) and their times from 5-second data (these data are available from the Reserve)

Maximum Wind Speed (m/s) and time from 5-second data

Wind Direction Standard Deviation (degrees)

Totals:

Precipitation (mm), PAR (μmol/m2), and Cumulative Precipitation (mm)

In-situ weather conditions are measured to verify the accuracy of the readings by the sensors, using a Kestrel 500 hand-held.  After downloading data, sensors on the weather station are inspected for damage or debris.  If any problems are found, it is repaired and/or cleaned. Tree limbs and other shrubs are cut back to prevent obstruction of the sensors.  Sensors are removed and sent back to Campbell Scientific for calibration according to the following guidelines:

- Temperature/Humidity- yearly recalibration

- Rain Gauge- yearly recalibration (in house)

- Wind Speed/Direction- every 2 years recalibration

- Barometric Pressure- every 2 years recalibration

- PAR- every 2 years recalibration

- CR1000-every 5 years

# Data collected underwent the QA/QC process described in 2b.

5)  Site location and character –

The ACE Basin National Estuarine Research Reserve (NERR) is located on the Southeastern Atlantic coast of the United States, including portions of Charleston, Colleton and Beaufort Counties in South Carolina.  The study area encompasses the Ashepoo, Combahee, and South Edisto River basins, which empty into St. Helena Sound.  Diverse estuarine wetlands provide extensive and complex habitat types for fish and wildlife.  The NERR consists of approximately 92,000 acres of tidal marshes.  Of this, 65,600 acres are salt marshes, 13,600 acres are brackish marshes and 12,100 acres are freshwater marshes.  Interspersed within these three tidal marsh zones are approximately 26,000 acres of managed wetlands, marsh impoundments, and hammock islands.  St. Helena Sound comprises approximately 23,870 acres of open coastal marine and estuarine waters.

The weather station is located at the Bennett's Point field station on Mosquito Creek, a navigable tributary off of the Ashepoo River.  The weather station’s latitudinal and longitudinal coordinates are: 32.55934 N and -80.45456 W and is located at sea level elevation.  The station is approximately 90 m from the creek (800 m from the Ashepoo River) in a grassy field, 80 m to the southwest of the field station.  The closest wind obstructions are oak trees, 25 m to the south and southwest of the weather station and 70 m from a public paved road.  The closest SWMP water quality station, Mosquito Creek, is approximately 1655 m from the weather station.

The Campbell Scientific data logger and the barometric sensor (sensor body at 1.5 m and sensor tube at 1.0 m) are enclosed in a Campbell Scientific enclosure box with moisture/humidity indicators and 4 desiccant packets.  A 6.1 m (15 foot) galvanized steel tower elevates the sensors above potential barriers and enhances the performance of each sensor.  The PAR sensor and anemometer are attached to a cross-bar at the top of the tower at a height of 6.1 m and the temperature/relative humidity sensor is attached approximately half way up the tower at a height of 3.35 m.  The solar panel is attached to a 2.5 m long arm, and is oriented to the east at approximately 47-degrees.  The tipping bucket rain gauge (1.3 m height) is located 2.5 m to the southeast of the box on a concrete level platform.  The sensors are wired to the CR1000 data logger following protocol in the CDMO Manual, with minor changes made in order to facilitate the function of upgraded sensors.  A new larger enclosure box and solar panel were necessary with the installation of the GOES telemetry system.

6)  Data collection period –

Weather data have been collected at the ACE Basin NERR Bennett's Point weather station since the station became operational in March 2001.  Data has been collected for the first quarter (~January – March) and second quarter (~April – June) of 2017.

| File Start Date and Time | File End Date and Time |
| --- | --- |
| 12/12/2016 11:45 | 01/09/2017 12:15 |
| 01/09/2017 12:30 | 02/06/2017 12:00 |
| 02/06/2017 12:15 | 03/03/2017 10:15 |
| 03/03/2017 14:45 | 03/04/2017 01:10 |
| 03/04/2017 01:15 | 04/10/2017 10:15* |
| 04/10/2017 11:15 | 05/03/2017 10:15 |
| 05/03/2017 12:15 | 05/31/2017 21:30 |
| 05/31/2017 21:45 | 06/12/2017 17:00 |
| 06/12/2017 17:15 | 07/13/2017 10:45 |

7)  Distribution –

According to the Ocean and Coastal Resource Management Data Dissemination Policy for the NERRS System-wide Monitoring Program, is as follows

NOAA retains the right to analyze, synthesize and publish summaries of the NERRS System-wide Monitoring Program data.  The NERRS retains the right to be fully credited for having collected and process the data.  Following academic courtesy standards, the NERR site where the data were collected should be contacted and fully acknowledged in any subsequent publications in which any part of the data are used.  The data set enclosed within this package/transmission is only as good as the quality assurance and quality control procedures outlined by the enclosed metadata reporting statement.  The user bears all responsibility for its subsequent use/misuse in any further analyses or comparisons.  The Federal government does not assume liability to the Recipient or third persons, nor will the Federal government reimburse or indemnify the Recipient for its liability due to any losses resulting in any way from the use of this data.

Requested citation format:

NOAA National Estuarine Research Reserve System (NERRS). System-wide Monitoring Program. Data accessed from the NOAA NERRS Centralized Data Management Office website: http://www.nerrsdata.org/ ; accessed 12 October 2016.

NERR meteorological data and metadata can be obtained from the Research Coordinator at the individual NERR site (please see Principal Investigators and Contact Persons), from the Data Manager at the Centralized Data Management Office (please see personnel directory under the general information link on the CDMO home page) and online at the CDMO home page www.nerrsdata.org.   Data are available in comma delimited format.

8)  Associated researchers and projects

The NERR Water Quality Monitoring Project is a study which records long-term water quality data for the ACE Basin in order to observe any changes or tends in water quality over time.  The six sampling sites are in separate tributaries of the South Edisto and Ashepoo Rivers.  Two sites represent an urban or "treatment" site, and two sites represent pristine areas within the Reserve.  A salinity gradient is also observed between these six sites:

Edisto Island – approximately 14 kilometers from the weather station, GPS coordinates: 32.50399 N and -80.32470 W.  In January 2015, Edisto Island station replaced Big Bay station as a primary station, and it is one km upstream of the Big Bay station (See the 2014 metadata report for a description of the Big Bay station).  It is located on the dock of the Edisto Interpretive Center at Edisto Beach State Park.  The Edisto Island station is bordered by Spartina alterniflora marsh, and American oyster (Crassostrea virginica) forms a reef along the creek bank.

St. Pierre Creek - approximately 9 kilometers from the weather station, GPS coordinates: 32.52800 N and -80.36144 W.  The station is surrounded by a wide expanse of Spartina alterniflora marsh.  Extensive mud flats and oyster reefs fringe the banks. Development in the immediate area is sparse, and this creek is subject to light boat traffic.

Fishing Creek – approximately 8 kilometers from the weather station, GPS coordinates: 32.63593 N and -80.36556 W.  The station is near Jehossee Island, a protected USFWS property, and is surrounded by extensive Spartina cynosuroides marsh and vast mud flats.  The upland area is characterized by slash pine, live oak, and cabbage palmetto.

Mosquito Creek – approximately 2 kilometers from the weather station, GPS coordinates: 32.5558 N and -80.4380 W.  The station is surrounded by agriculture fields and impounded wetlands.  A public boat landing and a commercial seafood business with three commercial shrimp boats and a fueling station are located about 0.8 km (0.5 mi) downstream of the monitoring station.

Grove Plantation – approximately 12 kilometers from the weather station, GPS coordinates: 32.6637 N and -80.4130 W.  The station is in the South Edisto River located at the Grove Plantation unit of the ACE Basin National Wildlife Refuge, which is owned and managed by the US Fish and Wildlife Service (USFWS).  The station is surrounded by tidal freshwater fringe marsh and managed wetlands (aka “waterfowl impoundments”).  The upland area is dominated by pine-mixed hardwood forest that is characterized by several species of oaks and pines.

Jehossee Island – approximately 9 kilometers from the weather station, GPS coordinates: 32.6209 N and -80.3965 W.  The station is in the South Edisto River proper.  It is located at the Jehossee Island unit of the ACE Basin National Wildlife Refuge, which is owned and managed by the US Fish and Wildlife Service (USFWS).  The station is surrounded by Spartina cynosuroides fringe marsh and managed wetlands (aka “waterfowl impoundments”).  The upland area is dominated by inland maritime forest that is characterized by slash pine, live oak, and cabbage palmetto.

Measurements for all sites are taken every fifteen minutes over an approximate two-week or four week collection period (depending upon fouling previously experienced and future fouling rate expected).

In July 1997, the Reserve staff initiated nutrient monitoring study.  The objective of the study is to ascertain the annual and tidal fluctuations in nutrient levels near our two data logger sites.  Nutrient levels are measured during a complete tidal cycle each month, and the samples are analyzed for ammonia, nitrite & nitrate, orthophosphate, and chlorophyll a concentrations.  In January of 2002, the nutrient monitoring protocol (NUT) was added to the NERR System Wide Monitoring Program (SWMP).

The results from these studies and additional studies conducted in the ACE Basin can be obtained by contacting the Reserve.

# II. Physical Structure Descriptors

9)  Sensor specifications –

Parameter: Temperature and Relative Humidity

Units: Celsius, %

Model #: HMP45C

Operating Temperature: -40 to +60 degree C

Temperature Measurement Range: -40 to +60 degree C

Temperature Accuracy: +/- 2% degree C @ 20 degree C

Relative Humidity Measurement Range: 0 - 100% non-condensing

RH Accuracy: +/- @% RH (0 - 90%) and +/- 3% (90 - 100%)

Uncertainty of calibration: +/- 1.2% RH

Serial Number: E3010019

Date of Last Calibration: 02/02/2016

Previous Calibrations: 05/03/2013, 01/28/2010

Dates of Sensor Use: 04/18/2016-03/03/2017

Serial Number: D3520073

Date of last calibration: 03/13/2014

Previous Calibrations:

Date of sensor Use: 07/23/2014-04/18/2016

Serial Number: A1710045

Date of last calibration: 01/28/2017

Previous Calibrations:

Date of sensor Use: 03/03/2017-current

# Parameter: Barometric Pressure

# Units: millibars (mb)

Model #: PTB101B

Calibrated Range: 26" - 32" (Standard)

Supply Voltage: 12 VDC at 12 mA

Accuracy: +/- 0.7 of span

Operating Temperature Range: -22 to +55 degree C

Serial Number: A195003

Date of last calibration: 02/29/2016

Previous Calibrations: 02/27/2012, 10/16/2008

Date of sensor Use: 04/18/2016-current

Serial Number: Y0930012

Date of last calibration: 05/13/2013

Previous Calibrations: 02/10/2010, 06/21/2007

Date of sensor Use: 02/27/2014-04/18/2016

Parameter: Wind speed

Units: meter per second (m/s)

Sensor type: 18 cm diameter 4-blade helicoids propeller molded of polypropylene

Parameter: Wind direction

Units: degrees

Sensor type: balanced vane, 38 cm turning radius

Model #: R.M. Young 05103 Wind Monitor (2) and 3001 MET One Wind Set (1)

Range:  0-60 m/s (134 mph); gust survival 100 m/s (220 mph)

Accuracy: +/- 0.3 m/s

Serial Number: 2365M3 (2)

Date of last calibration: 02/07/2016

Previous Calibrations: 03/02/2013, 10/31/2008

Date of sensor Use: 04/18/2016-03/03/2017

Serial Number: N/A (1)

Date of last calibration: 01/31/2017

Previous Calibrations: 02/04/2010, 06/21/2007, 05/06/2013

Date of sensor Use: 03/03/2017-current

Parameter: Photosynthetic Active Radiation (PAR)

Units: mmoles m-2 (total flux)

LI-COR Quantum Sensor

Model #: LI-190

Serial #: Q31559

Stability: < +/- 2% change over a 1 year period

Operating Temperature: -20 to +65 degree C

Sensitivity: Typically 30 µA per 100 klux

Light Spectrum Wavelength: 400 to 700 nanometers

Calibration Factor: 264.70 µmol m-2 s-1 per mV

Multiplier History and Date Changed:

Date of last calibration: 04/28/2016

Parameter: Photosynthetically Active Radiation (PAR)

Units:  mmoles m-2 (total flux)

Model #SQ110 Apogee Quantum Sensor

Serial # 22346

Light spectrum waveband:  410 to 655 nm

Temperature dependence: 0.06+/-0.06% per °C

Stability:  <±2% change over 1 yr

Operating Temperature:  -40°C to 70°C; Humidity 0 to 100%

Cosine Response: 45° zenith angle: +/- 2%; 75° zenith angle: +/- 5%

Sensitivity:  0.2mV per µmol s-1 m-2

Date of last calibration: 02/01/2017 (New)

- LI-COR PAR sensor 1 Calibrated 09/2002 – Installed May 3, 2003 – multiplier 1.27

- LI-COR PAR sensor 2 Calibrated 07/2005 – Installed Aug. 22, 2005 – multiplier 1.36

- LI-COR PAR sensor 1 Calibrated 06/2007 – Installed Apr. 22, 2008 – multiplier 1.34

- LI-COR PAR sensor 2 Calibrated 10/2008 – Installed Sept. 30, 2009 – multiplier 2.18

- LI-COR PAR sensor 1 Calibrated 02/2010 – Installed Dec. 9, 2010 – multiplier 1.38

- LI-COR PAR sensor 2 Calibrated 02/2012 – Installed Mar. 16, 2012 – multiplier 3.12

-SN Q35240

- LI-COR PAR sensor 1 Calibrated 05/2013 – Installed Feb. 27, 2014 – multiplier 1.54

-SN Q31559

- LI-COR PAR sensor 2 Retired 03/03/2016

-SN Q35240

- LI-COR PAR sensor 1 Calibrated 04/2016 – Installed May 06, 2016- multiplier 1.54

-SN Q31559

- Apogee SQ110 Installed Mar. 03, 2017-current – multiplier 0.025

-SN 22346

# Parameter: Precipitation

# Units: millimeters (mm)

# Tipping Bucket Rain Gauge

Texas Electronics

Model #: TR-525I

SN: 23864-399

Calibration: 0.01 inch per tip

Accuracy: +/- 3% (Rates of 1 to 6 inches per hour)

Date of last calibration: 04/10/2017

Previous Calibrations: 04/28/16, 02/27/2014, 12/09/2012, 06/07/2012, 05/24/2012, 02/20/2012, 02/07/2011, 03/08/2010, 07/18/2006

# Storage Module

Model # SM4M

Storage capacity: 2 million low-resolution data values

Program storage: stores up to 8 program with a total capacity of 128 KB

Processor: Hitachi H8S

Operating system: 64 KB, flash memory based, user downloadable

Operating range: Temp: -35 to +65 degree C

Baud rates: 9600, 76800

Power requirements: 5 +/- 0.3 VDC @ 100 mA

The CR1000 has 2 MB of Flash EEPROM that is used to store the Operating System.  Another 128 K Flash is used to store configuration settings.  A minimum of 2 MB SRAM is available for program storage (16K), operating system use, and data storage.  Additional storage is available by using a compact flash card in the optional CFM100 Compact Flash Module

Date CR1000 Installed: 06/29/2006

Date CR1000 Changed: 03/16/2012 SN 46152

Date Current CR1000 Installed: 03/03/2017 SN 005202 (internal battery replaced 05/03/2017)

CR1000 Firmware Version: CR1000_OS_31.02

Last update: 04/10/2017

CR1000 Program Version: ACEBPMET_V5.5.1_050317 (current)

ACEBPMET_V5.5_041017

ACEBPMET_V5.5_030217

ACEBPMET_V5.5_050616

ACEBPMET_V5.5_022714

10)  Coded variable definitions –

Sampling station:		Sampling site code:	Station code:

Bennett's Point		BP			acebpmet

11)  QAQC flag definitions –

QAQC flags provide documentation of the data and are applied to individual data points by insertion into the parameter’s associated flag column (header preceded by an F_).   During primary automated QAQC (performed by the CDMO), -5, -4, and -2 flags are applied automatically to indicate data that is above or below sensor range, or missing.  All remaining data are then flagged 0, as passing initial QAQC checks.   During secondary and tertiary QAQC 1, -3, and 5 flags may be used to note data as suspect, rejected due to QAQC, or corrected.

-5	Outside High Sensor Range

-4	Outside Low Sensor Range

-3	Data Rejected due to QAQC

-2	Missing Data

-1	Optional SWMP supported parameter

0	Passed Initial QAQC Checks

1	Suspect Data

2	Open - reserved for later flag

3	Open - reserved for later flag

4	Historical Data:  Pre-Auto QAQC

5	Corrected Data

12)  QAQC code definitions –.

QAQC codes are used in conjunction with QAQC flags to provide further documentation of the data and are also applied by insertion into the associated flag column.  There are three (3) different code categories, general, sensor, and comment.  General errors document general problems with the CR1000, sensor errors are sensor specific, and comment codes are used to further document conditions or a problem with the data.  Only one general or sensor error and one comment code can be applied to a particular data point, but some comment codes (marked with an * below)  can be applied to the entire record in the F_Record column.

General Errors

GIM	Instrument malfunction

GIT	Instrument recording error, recovered telemetry data

GMC	No instrument deployed due to maintenance/calibration

GMT	Instrument maintenance

GPD	Power down

GPF	Power failure / Low battery

GPR	Program reload

GQR	Data rejected due to QA/QC checks

GSM	See metadata

Sensor Errors

SDG	Suspect due to sensor diagnostics

SIC	Incorrect calibration constant, multiplier or offset

SIW	Incorrect wiring

SMT	Sensor maintenance

SNV	Negative value

SOC	Out of calibration

SQR	Data rejected due to QAQC checks

SSD	Sensor drift

SSN	Not a number / unknown value

SSM	Sensor malfunction

SSR	Sensor removed

Comments

CAF	Acceptable calibration/accuracy error of sensor

CCU	Cause unknown

CDF	Data appear to fit conditions

CML	Snow melt from previous snowfall event

CRE*	Significant rain event

CSM*	See metadata

CVT*	Possible vandalism/tampering

CWE*	Significant weather event

13)  Other remarks/notes –

Data are missing due to equipment or associated specific sensors not being deployed, equipment failure, time of maintenance or calibration of equipment, or repair/replacement of a sampling station platform.  Any NANs in the dataset stand for “not a number” and are the result of low power, disconnected wires, or out of range readings.  If additional information on missing data is needed, contact the Research Coordinator at the reserve submitting the data.

Small negative PAR values are within range of the LI-COR sensor and are due to normal errors in the sensor and the CR1000 data logger. The Maximum signal noise error for the LI-COR sensor is +/- 2.214 μmoles/m2 over a 15 minute interval.

After 04/10/2017 11:15, Relative Humidity data greater than 100 are within range of the sensor accuracy of +/-3% and are flagged and coded as suspect, <1> (CAF).  Values greater than 103 are rejected <-3> (SQR).

Please note that the 3001 MET One Wind Set has an offset of 0.2 and does not record values of 0.

Data recorded for all parameters (with the exception of cumulative precipitation) at the midnight timestamp (00:00) are the 15 minute averages and totals for the 23:45-23:59 time period of the previous day.  Cumulative precipitation data at the midnight timestamp (00:00) are the sum of raw (unrounded) precipitation data from 00:00 to 23:59 of the previous day.  Summing each individual 15-minute total precipitation value from the same period will result in small differences from cumulative precipitation due to rounding.  It is especially important to note how data at the midnight timestamp are recorded when using January 1st and December 31st data.  Note:  Cumulative precipitation is no longer available via export from the CDMO.  Please contact the Reserve or the CDMO for more information or to obtain these data.

All Parameters Blanket Statement for Bennett’s Point Weather Station

# Rejected Data (Flag <-3>)

# Data were missing for all parameters from 03/03/2017 at 10:30 to 03/04/2017 at 01:00 from CR1000 download.  A replacement CR1000 was installed with what has been determined as a low internal battery.  The external battery charged the computer and the CR1000 was able to begin to store data at 01:15 on 03/04/2017.  Telemetered data was available starting at 03/03/2017 at 14:45 through 03/04/2017 at 1:00.  The first record for the telemetered data (03/03/2017 at 14:45) and the first record of the CR1000 retaining data (03/04/2017 at 1:15) were flagged as <-3> [GIT] (CSM) and <-3> (CSM), respectively, due to uncertainty related to capturing all three of the 5 minute readings.

# Data are missing for all parameters from 05/03/2017 at 10:30 to 05/03/2017 at 12:45.  The CR1000 internal battery was replaced and a new program was uploaded.  The first record (05/03/2017 at 13:00) after power was reapplied was flagged as <-3> (CSM), due to uncertainty related to capturing all three of the 5 minute readings.

# Suspect Data (Flag <1>)

# Data are missing for all parameters from 03/03/2017 at 10:30 to 03/04/2017 at 01:00.  The CR1000 was replaced with what has been determined as a low internal battery.  The external battery charged the computer and the CR1000 was able to begin to store data at 01:15 on 03/04/2017.  Telemetered data was available starting at 03/03/2017 at 15:00 through 03/04/2017 at 1:00.  Telemetered data are flagged as <1> [GIT] (CSM).

# Missing Data (Flag <-2>)

# Data are missing for all parameters from 03/03/2017 at 10:30 to 03/03/2017 at 14:30.  The CR1000 was replaced with what has been determined as a low internal battery.  The external battery charged the computer and the CR1000 was able to begin to store data at 01:15 on 03/04/2017.  Telemetered data was available starting on 03/03/2017 at 14:45.  Missing data are flagged as <-2> [GMC] (CSM).

# Data are missing for all parameters from 04/10/2017 at 10:30 to 04/10/2017 at 11:00 from CR1000 download. The CR1000 was powered down for maintenance.  The missing data are flagged as <-2> [GMC] (CSM).

# Data are missing for all parameters from 05/03/2017 at 10:30 to 05/03/2017 at 12:45.  The CR1000 internal battery was replaced and a new program was uploaded. Missing data are flagged as <-2> [GMC] (CSM).

# Corrected Data (Flag <5>

# The CR1000 internal battery was replaced and a new program was uploaded which results in missing data from 05/03/2017 at 10:30 to 05/03/2017 at 12:45.  When power was reapplied, the time stamp was incorrect which results in six records being collected at the wrong time.  These values were moved down by 45 minutes to correct their time.  The first record was coded as <-3> (CSM) as detailed above.  Five records from 05/03/2017 at 13:15 to 05/03/2017 at 14:15 were coded as <5> [GPR] (CSM).  In addition, the time was corrected for MaxTempT, MinTempT, and MaxWSpdT during this time as well as for MaxTempT and MaxWSpdT for the 05/03/207 at 14:30 record by adding 45 minutes, which resulted with flagging <5> [GPR] (CSM) for temperature Max WSpd values.

# Temperature/Relative Humidity

# Passed Initial QAQC Checks (Flag <0>)

# During 2017 the CDMO discovered an incorrect line in the CR1000 programming.  If RHumidity>100 And RHumidity<108 Then RHumidity=100.  A decision was made by the DMC during 2006 to discontinue correcting >100 RH values to 100. This change was never made in our program and has remained in each updated version until it was removed during 2017.  By correcting all values >100 during data collection we may have missed erroneous values that could have indicated a problem with the RH sensor.  CSM coding was added to all RH data from 2007 until the programming change in 2017 on 04/10/2017 at 10:15.

# Rejected Data (Flag <-3>)

# None

# Suspect Data (Flag <1>)

# None

# Missing Data (Flag <-2>)

# None

# Barometric Pressure

# Rejected Data (Flag <-3>)

# None

# Suspect Data (Flag <1>)

None

# Missing Data (Flag <-2>)

# None

# Wind Speed/Direction

# Rejected Data (Flag <-3>)

# None

# Suspect Data (Flag <1>)

# None

# Missing Data (Flag <-2>)

# None

# Total Precipitation

# Rejected Data (Flag <-3>)

# None

# Suspect Data (Flag <1>)

# None

# Missing Data (Flag <-2>)

# None

# Corrected Data (Flag <5>

# The rain bucket accidentally tipped during site maintenance.  No rain fell while the technicians were onsite.  The data was corrected to a 0 reading for 03/03/2017 at 10:00.  The data was flagged <5> [SMT] (CSM).

# The rain bucket accidentally tipped during site maintenance.  A small amount of rain was in the bucket that was lost.  The data was corrected to a 0 reading for 06/20/2017 at 13:15.  The data was flagged <5> [SMT] (CSM).

# Total PAR

# Rejected Data (Flag <-3>)

PAR data for 2017 dataset is commented due to elevated nighttime PAR readings when the LI-COR sensor was deployed (switched to an Apogee on 03/03/2017).  Moisture seeping into the LI-COR sensor and/or sensor cable is thought to be the cause for the elevated nighttime PAR, however it is still under investigation.  Since the range of error for this sensor is ± 2.214 μmoles/m2 time points that fall between one half hour past sunset and one half hour before sunrise that have values greater than 2.214 µmoles/m2 are rejected.  This was done only when the nighttime readings were adjacent to the daytime readings.  Data are flagged <-3> (CSM).  Note that http://www.sunrisesunset.com was used to determine sunrise and sunset times using Green Pond, South Carolina as a city of reference.

On 03/03/2017 at 14:45, the LI-COR sensor was replaced with an Apogee sensor.  The transformer that was in the wiring for the LICOR was left in place for the Apogee.  Therefore, the data for Total PAR from 03/03/2017 at 14:45 to 03/08/2017 at 13:30 was rejected and flagged as <-3> [SIC] (CSM).

Negative PAR readings were recorded intermittently starting on 05/23/2017.  The wiring of the PAR sensor was assessed on 05/25/2017.  Water was evident within the spliced wires.  The wires were dried and re-taped.  Rainfall was recorded in the early morning of 05/22/2017 which may have initiated this issue.  The data from 05/22/2017 at 0:00 to 05/25/2017 at 15:30 were rejected and coded as <-3> [SSM] (CSM).

Suspect Data (Flag <1>)

PAR data for 2017 dataset is commented due to elevated nighttime PAR readings.  Moisture seeping into the LI-COR sensor and/or sensor cable is thought to be the cause for the elevated nighttime PAR, however it is still under investigation.  Since the range of error for this sensor is ± 2.214 µmoles/m2 time points that fall between one half hour before sunrise and one half hour past sunset that have values greater than 0 µmoles/m2 but less than 2.214 μmoles/m2, are marked as suspect.  This was done only when the nighttime readings were adjacent to the daytime readings.  Data are flagged <1> (CSM).  Those data values appear satisfactory; however, data users should be aware of possible data inconsistencies.  Negative PAR readings from either sensor type are commented and flagged as <1> (CAF).  Note that http://www.sunrisesunset.com  was used to determine sunrise and sunset times using Green Pond, South Carolina as a city of reference.

# Missing Data (Flag <-2>)

None

# Cumulative Precipitation

# Rejected Data (Flag <-3>

# None

# Suspect Data (Flag <1>)

# None

# Missing Data (Flag <-2>)

# None

# Corrected Data (Flag <5>

# The rain bucket accidentally tipped during site maintenance.  No rain fell while the technicians were onsite.  The data was corrected to a 0 reading for 03/03/2017 at 10:00 to 03/03/2017 at 10:15.  The data was flagged <5> [SMT] (CSM).

# The rain bucket accidentally tipped during site maintenance.  A small amount of rain was in the bucket that was lost.  The data was corrected to a 0 reading for 06/20/2017 at 13:15 and 0.254 was subtracted for the rest of the day (06/20/2017 at 13:15 to 06/20/2017 at 00:00).  The data was flagged <5> [SMT] (CSM).

14)  Station History

# Bennett’s Point:

2001 - Station became active with CR10X data logger. No data collected until March due to battery issues.  Data collection stopped in June.  PC208W software used to download data.  Parameters measured are: temperature, humidity, barometric pressure, rainfall, wind speed and direction and PAR.  Probes are located on 3-meter galvanized poles on top of control box.

2002 - Data collection resumed in January

2002 - NADP – National Atmospheric Deposition Program – began in January

2003 - All new probes installed on May 3 and old ones sent back to be calibrated.

2003 - A 15-foot galvanized tower is installed on which the probes are positioned

2005 - All newly calibrated probes installed on August 22

2006 - CR10X data logger replaced with the CR1000 data logger and PC208W software replaced with LoggerNet Software in June 29

2006 - NL115 Ethernet cable installed to allow for downloading of data and troubleshooting problems at Charleston Office.

2006 - GOES satellite telemetry equipment installed to view real-time weather data.

2006 - Rain tipping bucket calibrated on July 18

2006 - NL115 Ethernet cable cut by workers in December

2007 - NL115 Ethernet cable replaced and squirrels chewed through wire.  Lost connection in September.

2007 - NADP – National Atmospheric Deposition Program – ended in December

2008 - New program upload with Cumulative Precipitation parameter added in March

2008 - All newly calibrated probes installed on April 22

2009 - All newly calibrated probes installed on September 30

2010 - Rain tipping bucket calibrated on January 1

2010 - Rain tipping bucket calibrated in March 18

2010 - Newly calibrated probes installed on December 9 with the exception of the temperature probe.

2012 - Newly calibrated probes installed on March 16

2012 - New CR1000 installed on March 16

2012 - Rain bucket calibrated on June 7

2012 - Rain bucket calibrated on December 19

2013 - Barometric pressure calibrated May 16

2013 - Newly calibrated Temp/RH sensor installed July 31

2014 - Newly calibrated probes installed February 27

2014 - Rain bucket calibrated on February 27

2014 - Rain bucket calibrated on March 31

2014 - Newly calibrated Temp/RH sensor installed July 23

2015 - New battery installed November 25

2016 - Weather sensors sent to be calibrated January 15

2016 - Newly calibrated sensors (all except PAR) installed April 18

2016 - Newly calibrated PAR sensor installed May 06

2017 - Newly calibrated CR1000, Temp/RH, Wind and Apogee sensors installed March 03

2017 - New CR1000 program (ACEBPMET_V5.5_030217) uploaded March 03

2017 - Rain bucket calibrated on April 10

2017 - New Campbell OS V31.02 uploaded on April 10

2017 - New CR1000 program (ACEBPMET_V5.5_041017) uploaded on April 10

2017 - New CR1000 program (ACEBPMET_V5.5.1_050317) uploaded on May 3