# ACE Basin National Estuarine Research Reserve

# Meteorological Metadata Report

January - December 2016

Latest Update: 04/11/2017

# I.  Data Set and Research Descriptors

- Principal investigator(s) and contact persons

SCDNR/Marine Resources Division	ACE Basin NERR field station

217 Fort Johnson Road			15717 Bennett's Point Road

Charleston, SC  29422			Green Pond, SC 29441

(843) 953-9300				(843) 844-8822

Contact Persons:

Dr. Denise Sanger, Research Coordinator

# Contact: Sangerd@dnr.sc.gov  (843) 953-9074

# Saundra Upchurch, Reserve Biologist

# Contact: UpchurchS@dnr.sc.gov  (843) 953-9223

# Amanda Fornal, Reserve Biologist

# Meghan Miller, Reserve Technician

- Entry verification

# a)  Data Input Procedures:

Meteorological data is recorded in 15-minute intervals and stored in a Campbell Scientific CR1000 data logger.  ACE staff uses a data storage card to serve as a back-up for the CR1000 data logger.  The program that controls the sampling, schedule, and storage of data by the CR1000 is provided by CDMO.

The CR1000 is interfaced with the Campbell Scientific’s LoggerNet software.  ACE staff downloads and troubleshoots LoggerNet programs with a direct connection to the weather station, using a laptop computer and a 9-pin serial cable.  The GOES satellite system also uploads all 15-minute meteorological data summaries to a NOAA server every hour.  These multiple methods of data retrieval ensure that all measured parameters are reported.

# b)  QA/QC Procedures:

Data are uploaded from the CR1000 data logger to a Personal Computer (IBM compatible).  Files are exported from LoggerNet in comma-delimited format and uploaded to the CDMO where they undergo automated primary QAQC and become part of the CDMO’s online provisional database.  During primary QAQC, data are flagged if they are missing or out of sensor range.  The edited file is then returned to the Reserve where it is processed in Microsoft Excel using the CDMO’s NERRQAQC Excel macro.  The macro inserts station codes, creates metadata worksheets for flagged data and summary statistics, and graphs the data for review.  It allows the user to apply QAQC flags and codes to the data, append files, and export the resulting data file to the CDMO for tertiary QAQC and assimilation into the CDMO’s authoritative online database.  Outliers, suspect data, erroneous data, and other data flagged by the NERRQAQC macro are evaluated based on monthly field logs and NOAA historical data for the area.  For more information on QAQC flags and QAQC codes, see Sections 11 and 12.

# Denise Sanger, Saundra Upchurch, Amanda Fornal and Meghan Miller were responsible for the above tasks.

3)  Research objectives – (Campbell Weather Station):

The principal objective of the Weather Monitoring Program is to record long-term meteorological data for the ACE Basin NERR in order to observe any environmental changes or trends over time.

4)  Research methods –

Campbell Scientific data telemetry equipment was installed at the Bennett’s Point station on 06/30/2006 and transmits data to the NOAA GOES satellite, NESDIS ID #3B01E672.  The transmissions are scheduled hourly and contain four (4) data sets reflecting fifteen minute data sampling intervals.  Upon receipt by the CDMO, the data undergoes the same automated primary QAQC process detailed in Section 2 above.  The “real-time” telemetry data become part of the provisional dataset until undergoing secondary and tertiary QAQC and assimilation in the CDMO’s authoritative online database.  Provisional and authoritative data are available at http://cdmo.baruch.sc.edu.

The 15 minute Data are collected in the following formats for the CR1000:

Averages from 5-second data:

Air Temperature (°C), Relative Humidity (%), Barometric Pressure (mb), Wind Speed (m/s), Wind Direction (degrees) and Battery Voltage (volts)

Maximum and Minimum Air Temperature (°C) and their times from 5-second data (these data are available from the Reserve)

Maximum Wind Speed (m/s) and time from 5-second data

Wind Direction Standard Deviation (degrees)

Totals:

Precipitation (mm), PAR (µmol m-2 s-1) and Cumulative Precipitation (mm)

In-situ weather conditions are measured to verify the accuracy of the readings by the sensors, using a Kestrel 500 hand-held.  After downloading data, sensors on the weather station are inspected for damage or debris.  If any problems are found, it is repaired and/or cleaned. Tree limbs and other shrubs are cut back to prevent obstruction of the sensors.  Sensors are removed and sent back to Campbell Scientific for calibration according to the following guidelines:

- Temperature/Humidity- yearly recalibration

- Rain Gauge- yearly recalibration (in house)

- Wind Speed/Direction- every 2 years recalibration

- Barometric Pressure- every 2 years recalibration

- PAR- every 2 years recalibration

- CR1000-every 5 years

# Data collected underwent the QA/QC process described in 2b.

5)  Site location and character –

The ACE Basin National Estuarine Research Reserve (NERR) is located on the Southeastern Atlantic coast of the United States, including portions of Charleston, Colleton and Beaufort Counties in South Carolina.  The study area encompasses the Ashepoo, Combahee, and South Edisto River basins, which empty into St. Helena Sound.  Diverse estuarine wetlands provide extensive and complex habitat types for fish and wildlife.  The NERR consists of approximately 92,000 acres of tidal marshes.  Of this, 65,600 acres are salt marshes, 13,600 acres are brackish marshes and 12,100 acres are freshwater marshes.  Interspersed within these three tidal marsh zones are approximately 26,000 acres of managed wetlands, marsh impoundments, and hammock islands.  St. Helena Sound comprises approximately 23,870 acres of open coastal marine and estuarine waters.

The weather station is located at the Bennett's Point field station on Mosquito Creek, a navigable tributary off of the Ashepoo River.  The weather station’s latitudinal and longitudinal coordinates are: 32.55934 N and -80.45456 W and is located at sea level elevation.  The station is approximately 90 m from the creek (800 m from the Ashepoo River) in a grassy field, 80 m to the southwest of the field station.  The closest wind obstructions are oak trees, 25 m to the south and southwest of the weather station and 70 m from a public paved road.  The closest SWMP water quality station, Mosquito Creek, is approximately 1655m from the weather station.

The Campbell Scientific data logger and the barometric sensor (sensor body at 1.5 m and sensor tube at 1.0 m) are enclosed in a Campbell Scientific enclosure box with moisture/humidity indicators and 2 desiccant packets.  A 6.1 m (15 foot) galvanized steel tower elevates the sensors above potential barriers and enhances the performance of each sensor.  The PAR sensor and anemometer are attached to a cross-bar at the top of the tower at a height of 6.1 m and the temperature/relative humidity sensor is attached approximately half way up the tower at a height of 3.35m.  The solar panel is attached to a 2.5 m long arm, and is oriented to the east at approximately 47-degrees.  The tipping bucket rain gauge (1.3 m height) is located 2.5 m to the southeast of the box on a concrete level platform.  The sensors are wired to the CR1000 data logger following protocol in the CDMO Manual, with minor changes made in order to facilitate the function of upgraded sensors.   A new larger enclosure box and solar panel were necessary with the installation of the GOES telemetry system.

6)  Data collection period –

Weather data have been collected at the ACE Basin NERR Bennett's Point weather station since the station became operational in March 2001.  Data has been collected for all four quarters of the year (January 1, 2016 at 00:00 through December 31, 2016 at 23:45).

| File Start Date and Time | File End Date and Time |
| --- | --- |
| 12/09/2015 12:15 | 01/11/2016 14:45 |
| 01/11/2016 15:00 | 02/08/2016 13:15 |
| 02/08/2016 13:30 | 03/07/2016 11:15 |
| 03/07/2016 11:30 | 04/04/2016 15:00 |
| 04/04/2016 15:15 | 04/18/2016 11:00 |
| 04/18/2016 11:15 | 05/06/2016 09:45 |
| 05/09/2016 10:45 | 05/16/2016 11:30 |
| 05/16/2016 11:45 | 06/01/2016 02:00 |
| 06/01/2016 02:15 | 06/22/2016 10:00 |
| 06/22/2016 10:15 | 07/11/2016 12:30 |
| 07/11/2016 12:45 | 08/15/2016 09:45 |
| 08/15/2016 10:00 | 09/01/2016 12:45 |
| 09/01/2016 13:00 | 09/12/2016 11:45 |
| 09/12/2016 12:00 | 10/12/2016 13:30 |
| 10/12/2016 13:45 | 11/14/2016 14:00 |
| 11/14/2016 14:15 | 12/12/2016 11:30 |
| 12/12/2016 11:45 | 01/09/2017 12:15 |

7)  Distribution –

According to the Ocean and Coastal Resource Management Data Dissemination Policy for the NERRS System-wide Monitoring Program, is as follows

NOAA retains the right to analyze, synthesize and publish summaries of the NERRS System-wide Monitoring Program data.  The NERRS retains the right to be fully credited for having collected and process the data.  Following academic courtesy standards, the NERR site where the data were collected should be contacted and fully acknowledged in any subsequent publications in which any part of the data are used.  The data set enclosed within this package/transmission is only as good as the quality assurance and quality control procedures outlined by the enclosed metadata reporting statement.  The user bears all responsibility for its subsequent use/misuse in any further analyses or comparisons.  The Federal government does not assume liability to the Recipient or third persons, nor will the Federal government reimburse or indemnify the Recipient for its liability due to any losses resulting in any way from the use of this data.

Requested citation format:

NOAA National Estuarine Research Reserve System (NERRS). System-wide Monitoring Program. Data accessed from the NOAA NERRS Centralized Data Management Office website: http://www.nerrsdata.org/ ; accessed 12 October 2016.

NERR meteorological data and metadata can be obtained from the Research Coordinator at the individual NERR site (please see Principal Investigators and Contact Persons), from the Data Manager at the Centralized Data Management Office (please see personnel directory under the general information link on the CDMO home page) and online at the CDMO home page www.nerrsdata.org.   Data are available in comma delimited format.

8)  Associated researchers and projects

The NERR Water Quality Monitoring Project is a study which records long-term water quality data for the ACE Basin in order to observe any changes or tends in water quality over time. The six sampling sites are in separate tributaries of the South Edisto and Ashepoo Rivers. Two sites represent an urban or "treatment" site, and two sites represent pristine areas within the Reserve.  A salinity gradient is also observed between these six sites:

Edisto Island – approximately 14 kilometers from the weather station, GPS coordinates: 32.50399 N and -80.32470 W In January 2015, Edisto Island station replaced Big Bay station as a primary station, and it is one km upstream of the Big Bay station (See the 2014 metadata report for a description of the Big Bay station).  It is located on the dock of the Edisto Interpretive Center at Edisto Beach State Park.  The Edisto Island station is bordered by Spartina alterniflora marsh, and American oyster (Crassostrea virginica) forms a reef along the creek bank.

St. Pierre Creek - approximately 9 kilometers from the weather station, GPS coordinates: 32.52800 N and -80.36144 W.  The station is surrounded by a wide expanse of Spartina alterniflora marsh.  Extensive mud flats and oyster reefs fringe the banks. Development in the immediate area is sparse, and this creek is subject to light boat traffic.

Fishing Creek – approximately 8 kilometers from the weather station, GPS coordinates: 32.63593 N and -80.36556 W.  The station is near Jehossee Island, a protected USFWS property, and is surrounded by extensive Spartina cynosuroides marsh and vast mud flats. The upland area is characterized by slash pine, live oak, and cabbage palmetto.

Mosquito Creek – approximately 2 kilometers from the weather station, GPS coordinates: 32.5558 N and -80.4380 W.  The station is surrounded by agriculture fields and impounded wetlands.  A public boat landing and a commercial seafood business with three commercial shrimp boats and a fueling station are located about 0.8 km (0.5 mi) downstream of the monitoring station.

Grove Plantation – approximately 12 kilometers from the weather station, GPS coordinates: 32.6637 N and -80.4130 W.  The station is in the South Edisto River located at the Grove Plantation unit of the ACE Basin National Wildlife Refuge, which is owned and managed by the US Fish and Wildlife Service (USFWS).  The station is surrounded by tidal freshwater fringe marsh and managed wetlands (aka “waterfowl impoundments”). The upland area is dominated by pine-mixed hardwood forest that is characterized by several species of oaks and pines.

Jehossee Island – approximately 9 kilometers from the weather station, GPS coordinates: 32.6209 N and -80.3965 W.  The station is in the South Edisto River proper.  It is located at the Jehossee Island unit of the ACE Basin National Wildlife Refuge, which is owned and managed by the US Fish and Wildlife Service (USFWS). The station is surrounded by Spartina cynosuroides fringe marsh and managed wetlands (aka “waterfowl impoundments”). The upland area is dominated by inland maritime forest that is characterized by slash pine, live oak, and cabbage palmetto.

Measurements for all sites are taken every fifteen minutes over an approximate two-week or four week collection period (depending upon fouling previously experienced and future fouling rate expected).

In July 1997, the Reserve staff initiated nutrient monitoring study.  The objective of the study is to ascertain the annual and tidal fluctuations in nutrient levels near our two data logger sites.  Nutrient levels are measured during a complete tidal cycle each month, and the samples are analyzed for ammonia, nitrite & nitrate, orthophosphate, and chlorophyll a concentrations.  In January of 2002, the nutrient monitoring protocol (NUT) was added to the NERR System Wide Monitoring Program (SWMP).

The results from these studies and additional studies conducted in the ACE Basin can be obtained by contacting the Reserve.

# II. Physical Structure Descriptors

9)  Sensor specifications –

Parameter: Temperature and Relative Humidity

Units: Celsius, %

Model #: HMP45C

Operating Temperature: -40 to +60 °C

Temperature Measurement Range: -40 to +60 °C

Temperature Accuracy: +/- 2% °C @ 20 °C

Relative Humidity Measurement Range: 0 - 100% non-condensing

RH Accuracy: +/- @% RH (0 - 90%) and +/- 3% (90 - 100%)

Uncertainty of calibration: +/- 1.2% RH

Serial Number: E3010019

Date of Last Calibration: 02/02/2016

Previous Calibrations: 05/03/2013, 01/28/2010

Dates of Sensor Use: 04/18/2016 – current as of 12/31/2016

Serial Number: D3520073

Date of last calibration: 03/13/2014

Date of sensor Use: 07/23/2014-04/18/2016

# Parameter: Barometric Pressure

# Units: millibars (mb)

Model #: PTB101B

Calibrated Range: 26" - 32" (Standard)

Supply Voltage: 12 VDC at 12 mA

Accuracy: +/- 0.7 of span

Operating Temperature Range: -22 to +55 degree C

Serial Number: A195003

Date of last calibration: 02/29/2016

Previous Calibrations: 02/27/2012, 10/16/2008

Date of sensor Use: 04/18/2016- current as of 12/31/2016

Serial Number: Y0930012

Date of last calibration: 05/13/2013

Previous Calibrations: 02/10/2010, 06/21/2007

Date of sensor Use: 02/27/2014-04/18/2016

Parameter: Wind speed

Units: meter per second (m/s)

Sensor type: 18 cm diameter 4-blade helicoids propeller molded of polypropylene

Parameter: Wind direction

Units: degrees

Sensor type: balanced vane, 38 cm turning radius

Model #: R.M. Young 05103 Wind Monitor (2) and 3001 MET One Wind Set (1)

Range:  0-60 m/s (134 mph); gust survival 100 m/s (220 mph)

Accuracy: +/- 0.3 m/s

Serial Number: 2365M3 (2)

Date of last calibration: 02/07/2016

Previous Calibrations: 03/02/2013, 10/31/2008

Date of sensor Use: 04/18/2016- current as of 12/31/2016

Serial Number: N/A (1)

Date of last calibration: 05/06/2013

Previous Calibrations: 02/04/2010, 06/21/2007

Date of sensor Use: 02/27/2014 - 04/18/2016

Parameter: Photosynthetic Active Radiation (PAR)

Units: mmoles m-2 (total flux)

LI-COR Quantum Sensor

Model #: LI-190

Serial #: Q31559

Stability: < +/- 2% change over a 1 year period

Operating Temperature: -20 to +65 °C

Sensitivity: Typically 30 µA per 100 klux

Light Spectrum Wavelength: 400 to 700 nanometers

Calibration Factor: 264.70 µmol m-2 s-1 per mV

Multiplier History and Date Changed:

Date of last calibration: 04/28/2016

- PAR sensor 1 Calibrated 09/2002 – Installed May 3, 2003 – multiplier 1.27

- PAR sensor 2 Calibrated 07/2005 – Installed August 22, 2005 – multiplier 1.36

- PAR sensor 1 Calibrated 06/2007 – Installed April 22, 2008 – multiplier 1.34

- PAR sensor 2 Calibrated 10/2008 – Installed Sept. 30, 2009 – multiplier 2.18

- PAR sensor 1 Calibrated 02/2010 – Installed Dec. 9, 2010 – multiplier 1.38

- PAR sensor 2 Calibrated 02/2012 – Installed Mar. 16, 2012 – multiplier 3.12

-SN Q35240

- PAR sensor 1 Calibrated 05/2013 – Installed Feb. 27, 2014 – multiplier 1.54

-SN Q31559

- PAR sensor 2 Retired 01/2016

-SN Q35240

- PAR sensor 1 Calibrated 04/2016 – Installed May 06, 2016- multiplier 1.54

-SN Q31559

# Parameter: Precipitation

# Units: millimeters (mm)

# Tipping Bucket Rain Gauge

Texas Electronics

Model #: TR-525I

SN: 23864-399

Calibration: 0.01 inch per tip

Accuracy: +/- 3% (Rates of 1 to 6 inches per hour)

Date of last calibration: 04/28/2016

Previous Calibrations: 02/27/2014, 12/09/2012, 06/07/2012, 05/24/2012, 02/20/2012, 02/07/2011, 03/08/2010, 07/18/2006

# Storage Module

Model # SM4M

Storage capacity: 2 million low-resolution data values

Program storage: stores up to 8 program with a total capacity of 128 KB

Processor: Hitachi H8S

Operating system: 64 KB, flash memory based, user downloadable

Operating range: Temp: -35 to +65 degree C

Baud rates: 9600, 76800

Power requirements: 5 +/- 0.3 VDC @ 100 mA

The CR1000 has 2 MB of Flash EEPROM that is used to store the Operating System.  Another 128 K Flash is used to store configuration settings.  A minimum of 2 MB SRAM is available for program storage (16K), operating system use, and data storage.  Additional storage is available by using a compact flash card in the optional CFM100 Compact Flash Module

Date CR1000 Installed: 06/29/2006, New CR1000 Installed 03/16/2012 SN 46152

CR1000 Firmware Version:  CR1000_OS_29

Last update: 05/16/2016

CR1000 Program Version: ACEBPMET_V5.5_050616

ACEBPMET_V5.5_022714

10)  Coded variable definitions -

Sampling station:		Sampling site code:	Station code:

Bennett's Point	             BP		   	acebpmet

11)  QAQC flag definitions –

QAQC flags provide documentation of the data and are applied to individual data points by insertion into the parameter’s associated flag column (header preceded by an F_).   During primary automated QAQC (performed by the CDMO), -5, -4, and -2 flags are applied automatically to indicate data that is above or below sensor range, or missing.  All remaining data are then flagged 0, as passing initial QAQC checks.   During secondary and tertiary QAQC 1, -3, and 5 flags may be used to note data as suspect, rejected due to QAQC, or corrected.

-5	Outside High Sensor Range

-4	Outside Low Sensor Range

-3		Data Rejected due to QAQC

-2		Missing Data

-1		Optional SWMP supported parameter

0		Passed Initial QAQC Checks

1		Suspect Data

2		Open - reserved for later flag

3		Open - reserved for later flag

4		Historical Data:  Pre-Auto QAQC

5		Corrected Data

12)  QAQC code definitions –.

QAQC codes are used in conjunction with QAQC flags to provide further documentation of the data and are also applied by insertion into the associated flag column.  There are three (3) different code categories, general, sensor, and comment.  General errors document general problems with the CR1000, sensor errors are sensor specific, and comment codes are used to further document conditions or a problem with the data.  Only one general or sensor error and one comment code can be applied to a particular data point, but some comment codes (marked with an * below)  can be applied to the entire record in the F_Record column.

General Errors

GIM	Instrument malfunction

GIT	Instrument recording error, recovered telemetry data

GMC	No instrument deployed due to maintenance/calibration

GMT 	Instrument maintenance

GPD	Power down

GPF	Power failure / Low battery

GPR	Program reload

GQR	Data rejected due to QA/QC checks

GSM	See metadata

Sensor Errors

SDG	Suspect due to sensor diagnostics

SIC		Incorrect calibration constant, multiplier or offset

SIW	Incorrect wiring

SMT	Sensor maintenance

SNV	Negative value

SOC	Out of calibration

SQR	Data rejected due to QAQC checks

SSD	Sensor drift

SSN	Not a number / unknown value

SSM	Sensor malfunction

SSR		Sensor removed

Comments

CAF	Acceptable calibration/accuracy error of sensor

CCU	Cause unknown

CDF	Data appear to fit conditions

CML	Snow melt from previous snowfall event

CRE*	Significant rain event

CSM*	See metadata

CVT*	Possible vandalism/tampering

CWE*	Significant weather event

13)  Other remarks/notes –

Data are missing due to equipment or associated specific sensors not being deployed, equipment failure, time of maintenance or calibration of equipment, or repair/replacement of a sampling station platform.  Any NANs in the dataset stand for “not a number” and are the result of low power, disconnected wires, or out of range readings.  If additional information on missing data is needed, contact the Research Coordinator at the reserve submitting the data.

Small negative PAR values are within range of the sensor and are due to normal errors in the sensor and the CR1000 data logger. The Maximum signal noise error for the Licor sensor is +/- 2.214 µmol/m2 over a 15 minute interval.

Relative Humidity data greater than 100 are within range of the sensor accuracy of +/-3%.

Please note that the 3001 MET One Wind Set has an offset of 0.2 and does not record values of 0.

Data recorded for all parameters (with the exception of cumulative precipitation) at the midnight timestamp (00:00) are the 15 minute averages and totals for the 23:45-23:59 time period of the previous day.  Cumulative precipitation data at the midnight timestamp (00:00) are the sum of raw (unrounded) precipitation data from 00:00 to 23:59 of the previous day.  Summing each individual 15-minute total precipitation value from the same period will result in small differences from cumulative precipitation due to rounding.  It is especially important to note how data at the midnight timestamp are recorded when using January 1st and December 31st data.  Note:  Cumulative precipitation is no longer available via export from the CDMO.  Please contact the Reserve or the CDMO for more information or to obtain these data.

All Parameters Blanket Statement for Bennett’s Point Weather Station

Table 1 list significant weather events from the National Hurricane Center administered by the National Oceanic and Atmospheric Administration (NOAA).  All events listed have impacted the South Carolina Coast in 2016.

Table 1:

| Significant Weather Event | Date Started | Date Ended | Max Wind Speed | Min. Pressure |
| --- | --- | --- | --- | --- |
| Hurricane Hermine AL092016 | August 28, 2016 | September 3, 2016 | 70 kt | 981 mb |
| Tropical Storm Julia AL112016 | September 13, 2016 | September 18, 2016 | 45 kt | 1007 mb |
| Hurricane Matthew AL142016 | September 28, 2016 | October 9, 2016 | 145 kt | 934 mb |

Hurricane Hermine:

Hurricane Hermine was a category 1 storm that made landfall along the Big Bend coast of Florida.  Hermine moved across Georgia, South Carolina, and North Carolina as a tropical storm and then turned off the mid-Atlantic coast as an extratropical low.  The impact of hurricane Hermine was observed at the Bennett’s Point meteorological station in Green Pond, South Carolina during the below times;

Data collected from 09/01/2016 at 00:00 to 09/03/2016 at 23:45 are flagged for a significant weather event.  All parameters were affected therefore; the F_Record column is flagged as {CWE}.

Tropical Storm Julia:

Tropical Storm Julia was a cyclone that formed near the coast of southeastern Florida.  The storm continued to move north-northwest along the eastern Florida and southeastern Georgia coasts.  Julia maintained intensity for a couple of days while it moved erratically eastward and southeastward off the South Carolina coast.  The impact of tropical storm Julia was observed at the Bennett’s Point meteorological station in Green Pond, South Carolina during the below times;

Data collected from 09/13/2016 at 00:00 to 09/15/2016 at 23:45 are flagged for a significant weather event.  All parameters were affected therefore; the F_Record column is flagged as {CWE}.

Hurricane Matthew:

Hurricane Mathew was a category 5 hurricane that moved north off the Florida east coast before weakening to a category 2 storm off the coast of Georgia.  Hurricane Matthew made land fall near Cape Romain, South Carolina as a category 1 hurricane.  Matthew maintained an east-northeastward motion until it lost its tropical characteristics.  The impact of hurricane Matthew was observed at the Bennett’s Point meteorological station in Green Pond, South Carolina during the below times;

Data collected from 10/04/2016 at 00:00 to 10/09/2016 at 23:45 are flagged for a significant weather event.  All parameters were affected therefore; the F_Record column is flagged as {CWE}.

# Rejected Data (Flag <-3>)

# Data are missing for all parameters from 05/06/2016 at 10:00 to 05/09/2016 at 10:30 due to instrument malfunction (see below).  The first record of the CR1000 retaining data (05/09/2016 at 10:45) were flagged as <-3> (CSM) due to uncertainty related to capturing all three of the 5 minute readings.

# Suspect Data (Flag <1>)

None

# Missing Data (Flag <-2>)

# Data are missing for all parameters from 05/06/2016 at 10:00 to 05/09/2016 at 10:30 due to instrument malfunction.  A new computer was being set up to communicate with the CR1000 and the MET program was not able to be loaded to the CR1000.  The problem was repaired on 05/09/2016, and repairs were completed by 05/09/2016 at 10:30.  Data are flagged as <-2> [GIM] (CSM).

# Temperature/Relative Humidity

# Rejected Data (Flag <-3>)

# None

# Suspect Data (Flag <1>)

# None

# Missing Data (Flag <-2>)

# None

# Passed Initial QAQC Checks (Flag <0>)

# During 2017, the CDMO discovered an incorrect line in the CR1000 programming.  If RHumidity>100 And RHumidity<108 Then RHumidity=100.  A decision was made by the DMC during 2006 to discontinue correcting >100 RH values to 100. This change was never made in our program and has remained in each updated version until it was removed during 2017.  By correcting all values >100 during data collection we may have missed erroneous values that could have indicated a problem with the RH sensor.  CSM coding was added to all RH data from 2007 until the programming change in 2017.

# Barometric Pressure

# Rejected Data (Flag <-3>)

# None

# Suspect Data (Flag <1>)

BP data collected on 04/18/2016 at 11:45 was considered suspect due to prior readings and after reading at ~1026 in comparison to flagged value at 1010.  This was two readings after the weather station was undergoing maintenance.  The data are flagged as <1> [GSM] .(CCU).

# Missing Data (Flag <-2>)

# None

# Wind Speed/Direction

# Rejected Data (Flag <-3>)

# None

# Suspect Data (Flag <1>)

Data collected from 02/27/2016 at 00:00 to 04/18/2016 at 11:15 were suspect due to sensor past calibration due date.  Values are flagged <1> [SOC] (CSM).

# Missing Data (Flag <-2>)

# None

# Total Precipitation

# Rejected Data (Flag <-3>)

# None

# Suspect Data (Flag <1>)

# Data collected from 03/31/2016 at 00:00 to 04/18/2016 at 11:15 were suspect due to sensor past calibration due date.  Values were flagged as <1> [SOC] (CSM).

# Missing Data (Flag <-2>)

# None

# Corrected Data (Flag <5>)

# The rain bucket accidentally tipped during site maintenance.  No rain fell while the technicians were onsite.  The data was corrected to a 0 reading for 09/01/2016 at 13:30.  The data was flagged <5> [SMT] (CSM).

# The rain bucket accidentally tipped during site maintenance.  No rain fell while the technicians were onsite.  The data was corrected to a 0 reading for 10/26/2016 at 11:15.  The data was flagged <5> [SMT] (CSM).

# Total PAR

# Rejected Data (Flag <-3>)

PAR data for 2016 dataset is commented due to elevated nighttime PAR readings.  Moisture seeping into the LI-COR sensor and/or sensor cable is thought to be the cause for the elevated nighttime PAR, however it is still under investigation.  Since the range of error for this sensor is ± 2.214 mmoles/m2 timepoints that fall between one half hour past sunset and one half hour before sunrise that have values greater than 2.214 mmoles/m2 are rejected.  This was done only when the elevated nighttime readings were adjacent to the daytime readings.  Data are flagged <-3> (CSM).  Note that http://www.sunrisesunset.com was used to determine sunrise and sunset times using Green Pond, South Carolina as a city of reference.

# PAR data collected from 4/18 at 11:30 to 5/6 at 09:45 were recorded as NAN.  Data are coded as <-3> [GCM] (CSM).

Suspect Data (Flag <1>)

PAR data for 2016 dataset is commented due to elevated nighttime PAR readings.  Moisture seeping into the LI-COR sensor and/or sensor cable is thought to be the cause for the elevated nighttime PAR, however it is still under investigation.  Since the range of error for this sensor is ± 2.214 mmoles/m2 timepoints that fall between one half hour before sunrise and one half hour past sunset that have values greater than 0 mmoles/m2 but less than 2.214 mmoles/m2, are marked as suspect.  This was done only when the elevated nighttime readings were adjacent to the daytime readings.  Data are flagged <1> (CSM).  Those data values appear satisfactory; however, data users should be aware of possible data inconsistencies.Note that http://www.sunrisesunset.com was used to determine sunrise and sunset times using Green Pond, South Carolina as a city of reference.

Negative PAR readings are commented and flagged as <1> (CAF).

# Missing Data (Flag <-2>)

# Data collected from 04/18/2016 at 11:30 to 05/06/2016 at 09:45 are missing due to PAR sensor not being deployed while being calibrated.  NAN values flagged as <-2> [GMC] (CSM).

# Data collected on 05/09/2017 at 10:45 is missing due to PAR sensor not being deployed while being calibrated.  NAN value flagged as <-2> [GMC] (CSM).

# Cumulative Precipitation

# Rejected Data (Flag <-3>

# None

# Suspect Data (Flag <1>)

# Data collected from 03/31/2016 at 00:00 to 04/18/2016 at 11:15 are suspect due to sensor past calibration due date.  Values are flagged <1> [SOC] (CSM).

# Missing Data (Flag <-2>)

# None

# Corrected Data (Flag <5>)

# The rain bucket accidentally tipped during site maintenance.  No rain fell while the technicians were onsite.  The data was corrected to by subtracting 1.0 from each reading for 09/01/2016 at 13:30 to 09/02/2016 at 00:00.  The data was flagged <5> [SMT] (CSM).

# The rain bucket accidentally tipped during site maintenance.  No rain fell while the technicians were onsite.  The data was corrected by subtracting 0.5 from each reading for 10/26/2016 at 11:15 to 10/27/2016 at 00:00.  The data was flagged <5> [SMT] (CSM).

14)  Station History

# Bennett’s Point:

2001 - Station became active with CR10X data logger. No data collected until March due to battery issues.  Data collection stopped in June.  PC208W software used to download data.  Parameters measured are: temperature, humidity, barometric pressure, rainfall, wind speed and direction and PAR.  Probes are located on 3-meter galvanized poles on top of control box.

2002 - Data collection resumed in January

2002 - NADP – National Atmospheric Deposition Program – began in January

2003 - All new probes installed on May 3 and old ones sent back to be calibrated.

2003 - A 15-foot galvanized tower is installed on which the probes are positioned

2005 - All newly calibrated probes installed on August 22

200 6- CR10X data logger replaced with the CR1000 data logger and PC208W software replaced with LoggerNet Software in June 29

2006 - NL115 Ethernet cable installed to allow for downloading of data and troubleshooting problems at Charleston Office.

2006 - GOES satellite telemetry equipment installed to view real-time weather data.

2006 - Rain tipping bucket calibrated on July 18

2006 - NL115 Ethernet cable cut by workers in December

2007- NL115 Ethernet cable replaced and squirrels chewed through wire.  Lost connection in September.

2007 - NADP – National Atmospheric Deposition Program – ended in December

2008 - New program upload with Cumulative Precipitation parameter added in March

2008 - All newly calibrated probes installed on April 22

2009 - All newly calibrated probes installed on September 30

2010 - Rain tipping bucket calibrated on January 1

2010 - Rain tipping bucket calibrated in March 18

2010 - Newly calibrated probes installed on December 9 with the exception of the temperature probe.

2012 - Newly calibrated probes installed on March 16

2012 - New CR1000 installed on March 16

2012 - Rain bucket calibrated on June 7

2012 - Rain bucket calibrated on December 19

2013 – Barometric pressure calibrated May 16

2013 - Newly calibrated Temp/RH sensor installed July 31

2014 - Newly calibrated probes installed February 27

2014 - Rain bucket calibrated on February 27

2014 - Rain bucket calibrated on March 31

2014 - Newly calibrated Temp/RH sensor installed July 23

2015 - New battery installed November 25

2016 - Weather sensors sent to be calibrated January 15

2016 - Newly calibrated sensors (all except PAR) installed April 18

2016 - Newly calibrated PAR sensor installed May 06