Grand Bay (GND) National Estuarine Research Reserve Meteorological Metadata

January 1 – December 31, 2012

Latest Update:  July 8, 2014

I. Data Set & Research Descriptors

1) Principal investigators & contact persons:

Address:		Grand Bay National Estuarine Research Reserve

6005 Bayou Heron Road

Moss Point, MS  39562

Phone:  (228) 475-7047

Fax:      (228) 475-8097

Contact Persons:	Dr. Mark Woodrey, Research Coordinator

E-mail:  msw103@ra.msstate.edu

(228) 475-7047, ext. 107

Kim Cressman, Research Technician

E-mail:  kimberly.cressman@dmr.ms.gov

(228) 475-7047, ext. 119

2) Entry verification:

The meteorological information is sampled every five seconds from each instrument on the weather station and stored on a Campbell Scientific CR1000 data logger.  Data are output to a file in one array: array 15 stores 15 minute average and maximum and minimum data.  The program run by the CR1000 is updated as necessary to reflect changes in PAR and Barometric Pressure sensor multipliers (see Section 13 below for additional program information).  The CR1000 interfaces with the LoggerNet software supplied by Campbell Scientific for data retrieval and conversion.

Data are uploaded from the CR1000 data logger to a Personal Computer (PC compatible).  Files are exported from LoggerNet in a comma-delimited format and uploaded to the CDMO where they undergo automated primary QAQC and become part of the CDMO’s online provisional database.  During primary QAQC, data are flagged if they are missing or out of sensor range.  The edited file is then returned to the Reserve where it is opened in Microsoft Excel and processed using the CDMO’s NERRQAQC Excel macro.  The macro inserts station codes, creates metadata worksheets for flagged data and summary statistics, and graphs the data for review.  It allows the user to apply QAQC flags and codes to the data, append files, and export the resulting data file to the CDMO for tertiary QAQC and assimilation into the CDMO’s authoritative online database.  For more information on QAQC flags and QAQC codes, see Sections 11 and 12.

Kim Cressman is responsible for data management.

3) Research objectives:

The National Estuarine Research Reserve (NERR) System-wide Monitoring Program (SWMP) was designed to fulfill two major overall goals:  1) to support state-specific non-point pollution control programs by establishing local networks of continuous water quality monitoring stations in representative protected estuarine ecosystems and 2) to develop a nation-wide database of baseline environmental conditions in the NERR system of estuaries.  The specific goal of SWMP is to identify and track short-term variability and long-term changes in the integrity and biodiversity of representative estuarine ecosystems and coastal watersheds for the purpose of contributing to effective national, regional, and site specific coastal zone management.  This comprehensive program consists of three phased components:  1) abiotic conditions such as water quality and meteorological monitoring; 2) biodiversity monitoring; and 3) habitat mapping and change analysis.  With the initial focus of phase 1, the NERR SWMP provides data necessary for intra- and inter- site baseline studies, trend analyses, and impact assessments.

The principal objective of meteorological monitoring is to record meteorological information for the Grand Bay NERR that can be used 1) as a vital reference of atmospheric data for various research projects at the reserve -- an integral part of the general NERR mission is to provide a platform for estuarine research, 2) to give meteorological context (atmospheric-forcing) for the half-hourly SWMP water quality data, and other long-term environmental monitoring programs at the reserve (including nutrients and shoreline change), 3) to observe and characterize important events, such as storms, heat and cold waves, droughts and heavy rainfalls, and 4) to detect trends and characterize climate variability over the long-term.

4) Research methods:

The Campbell Scientific weather station samples every five seconds continuously throughout the year.  These data are used by the CR1000 to produce 15 minute averages of air temperature, relative humidity, barometric pressure, rainfall, wind speed, and wind direction; maximums and minimums of air temperature and maximum wind speed and the time of occurrence for each; standard deviation of wind speed; total precipitation, photosynthetically active radiation, and daily cumulative precipitation totals.

The 15 minute Data are collected in the following formats for the CR1000:

Averages from 5-second data:

Air Temperature (°C), Relative Humidity (%), Barometric Pressure (mb), Wind Speed (m/s), Wind Direction (degrees), Battery Voltage (volts)

Maximum and Minimum Air Temperature (°C) and their times from 5-second data (these data are not included in the data set but are available from the Reserve)

Maximum Wind Speed (m/s) and time from 5-second data

Wind Direction Standard Deviation (degrees)

Totals:

Precipitation (mm), PAR (millimoles/m2), and Cumulative Precipitation (mm)

Sensors on the weather station are inspected monthly for damage and/or debris.  During these monthly observations, data taken from a nearby airport and a Kestrel 3000 are used to check the calibrations of the sensors.  Sensors are removed and calibrated on an annual or biannual basis depending on the particular sensor.

- Temperature/Humidity- annual recalibration

- Rain Gauge- annual recalibration

- Wind Speed/Direction- bi-annual recalibration

- Barometric Pressure- bi-annual recalibration

- PAR- bi-annual recalibration

Campbell Scientific data telemetry equipment was installed at the Crooked Bayou station on July 21, 2006 and transmits data to the NOAA GOES satellite, NESDIS ID #3B0190E2.  The transmissions are scheduled hourly and contain four (4) data sets reflecting fifteen minute data sampling intervals.  Upon receipt by the CDMO, the data undergoes the same automated primary QAQC process detailed in Section 2 above.  The near real-time telemetry data become part of the provisional dataset until undergoing secondary and tertiary QAQC and assimilation in the CDMO’s authoritative online database.  Provisional and authoritative data are available at http://cdmo.baruch.sc.edu.

5) Site location and character:

The Grand Bay National Estuarine Research Reserve (GBNERR) is a shallow, microtidal estuary located on the Gulf Coast in southeastern Jackson County, Mississippi.  It lies within the Mississippi Delta province of the Louisianian biogeographic region of the United States as identified by the National Oceanic and Atmospheric Administration.  Currently, no other reserves exist within the Mississippi Delta province of this biogeographic region.  The GBNERR encompasses approximately 7,450 ha (18,400 acres) of state-owned, federally-owned, and privately-owned lands and open water areas.  The primary land owners within the reserve boundaries are the State of Mississippi (3,684 ha; 9,100 ac) and the U.S. Fish and Wildlife Service (2,024 ha; 5,000 ac).  The balance of the property is privately owned (1,740 ha; 4,300 acres).

The Grand Bay meteorological station is located near the center of the Reserve (30° 21’ 33.12 N, 88° 25’ 12.00 W) on a 3m x 3m platform 10m away from a saltwater marsh at an elevation of 1.5m above mean high tide.  The platform is approximately 2m above the substrate.  One of four hydrological monitors is located in Point Aux Chenes Bay, 1183m south.  Another monitor is located on the east side of the reserve 6652m north/northeast of the station.  The other two monitors are located on the west side of the reserve 3141m north/northwest and 4127m due west of the station.  Although small clusters of dwarfed trees are present on nearby shell middens, they are not considered to serve as wind blocks, shades, or obstructions to the weather station due to the elevation of the station and the meteorological sensors.

The wind monitor and PAR sensor are mounted at a height of 3m above the platform onto a crossbar sensor mount oriented in an east/west direction.  The wind monitor is mounted on the west side of the crossbar, and the PAR sensor is mounted on the east side of the crossbar.  Lower on the tower, the temperature and relative humidity probe is mounted inside a 12-plate gill radiation shield attached to the tower leg facing the prevailing wind (E/SE) 1.7m (68”) above the tower base.  The barometric pressure sensor is mounted inside the control panel attached to two tower legs facing east/northeast 1.3m (50”) above the tower base.  The rain gauge is mounted at a height of 1.8m (72”) to the east post of the chain link fence surrounding the weather station.

6) Data Collection Period

The Grand Bay NERR weather station originally began collecting data on August 17, 2004.  On August 29, 2005, the platform and weather station were physically removed from the pilings by Hurricane Katrina.  The 10’ x 10’ platform and meteorological station were reconstructed and began collecting data on June 21, 2006.  This metadata is applicable only to data collected from January 1, 2012 00:00 to December 31, 2012 23:45.

7) Distribution

NOAA retains the right to analyze, synthesize and publish summaries of the NERRS System-wide Monitoring Program data.  The PI retains the right to be fully credited for having collected and processed the data.  The NERRS retains the right to be fully credited for having collected and process the data.  Following academic courtesy standards, the NERR site where the data were collected should be contacted and fully acknowledged in any subsequent publications in which any part of the data are used.  The data set enclosed within this package/transmission is only as good as the quality assurance and quality control procedures outlined by the enclosed metadata reporting statement.  The user bears all responsibility for its subsequent use/misuse in any further analyses or comparisons.  The Federal government does not assume liability to the Recipient or third persons, nor will the Federal government reimburse or indemnify the Recipient for its liability due to any losses resulting in any way from the use of this data.

Requested citation format:

National Estuarine Research Reserve System (NERRS). 2012.  System-wide Monitoring Program. Data accessed from the NOAA NERRS Centralized Data Management Office website: http://cdmo.baruch.sc.edu/; accessed 12 October 2012.

NERR water quality data and metadata can be obtained from the Research Coordinator at the individual NERR site (please see Principal Investigators and Contact Persons), from the Data Manager at the Centralized Data Management Office (please see personnel directory under the general information link on the CDMO home page) and online at the CDMO home page http://cdmo.baruch.sc.edu/.  Data are available in comma delimited format.

8) Associated researchers and projects:

Several research and monitoring projects are currently using the meteorological data from the Grand Bay NERR.  These projects will be listed in the annual metadata document.

- Effects of Landscape-Level Metrics on Marsh Bird Occupancy and Abundance within Tidal Marshes of the Northern Gulf of Mexico – Dr. Bob Cooper (University of Georgia), Ali Leggett (Mississippi Department of Marine Resources), Dr. Mark Woodrey (Grand Bay NERR/Mississippi State University)

- Shorebird and other waterbird use of wetlands and aquaculture ponds in the Mississippi Alluvial Valley and Gulf Coast region – Dr. Francisco Vilella (Mississippi State University)

- Testing habitat model assumptions for the Seaside Sparrow (Ammodramus maritimus) in Northern Gulf of Mexico Tidal Salt Marshes - Dr. Bob Cooper (University of Georgia), Dr. Mark Woodrey (Grand Bay NERR/Mississippi State University)

- Resolving Drivers of Variability in Estuarine Metabolism from Sustained Observations of Water Quality in the SE US - Dr. Jane Caffrey (University of West Florida)

- The Response of Phytoplankton Communities to Nutrient Enrichment Experiments in Three Northern Gulf of Mexico Estuaries - Kendra Straub (University of West Florida), Dr. Jane Caffrey (University of West Florida)

- Erosion Monitoring - 6 Monitoring Stations Used to Estimate the Rate of Erosion at Sites Representing Varying Degrees of Wave Exposure and Geological Substrates - Tom Strange (Grand Bay NERR), Will Underwood (Grand Bay NERR)

- Legacy Effects of Land-use Change and Nitrogen Source Shifts on a Benchmark System: Building Capacity for Collaborative Research Leadership at the Grand Bay Reserve - Lead PI Dr. Ruth Carmichael (Dauphin Island Sea Lab, University of South Alabama); collaborators: Elizabeth Condon (PhD student, Dauphin Island Sea Lab, University of South Alabama), Capt William Burkhardt (USFDA Office of Food Safety), Cdr Kevin Calci (USFDA Office of Food Safety), Dr. Wei Wu (Gulf Coast Research Laboratory, University of Southern Mississippi), Dave Ruple (Grand Bay NERR), Dr. William Walton (Auburn University Shellfish Laboratory)

- Monitoring breeding marsh bird populations at the Grand Bay NERR and Pascagoula River Marshes Coastal Preserve – Dr. Mark Woodrey (Grand Bay NERR/Mississippi State University) Dr. Bob Cooper (University of Georgia), Dr. Scott Rush (Mississippi State University)

- Degradation Following Watershed Land Use Changes - Dr. Mark S. Woodrey (Co-PI; coauthor of grant proposal with C.A. May, Grand Bay NERR (CO-PI), Dr. J. Cebrian, Dauphin Island Sea Lab (Co-PI), Dr. R.H. Beard, NOAA (Co-PI), and Dr. A.R. Parsons, NOAA (Co-PI);

- Developing Optimal Survey Techniques for Monitoring Population Status of Rails, Snipe, Coots, and Gallinules - Dr. Mark S. Woodrey (Co-PI; coauthor of grant proposal with Dr. C.J. Conway, U.S.G.S. (PI);

- Fish Communities of Nearshore Habitats within the Grand Bay NERR/NWR – Ron Cole (Grand Bay NERR), Brenna Ehmen (Grand Bay NERR), Jake Walker (Grand Bay NERR), Dr. Mark Woodrey (Grand Bay NERR/Mississippi State University);

- Nesting ecology of the diamondback terrapin (Malaclemys terrapin pileata) at the Grand Bay National Estuarine Research Reserve, Mississippi - Collaborators - Christina Mohrman (Florida A&M University/Grand Bay NERR) and Dr. Mark Woodrey (Mississippi State University/Grand Bay NERR);

- Distribution and Abundance of Submerged Aquatic Vegetation at the Grand Bay NERR - a project being conducted by Grand Bay NERR staff

- Atmospheric Mercury Deposition Monitoring at the Grand Bay NERR - collaborators: Dr. Winston Luke (Air Resources Lab - NOAA), Dr. Mark Cohen (Air Resources Lab - NOAA), Mr. Dave Ruple (Grand Bay NERR), Dr. Mark Woodrey (Grand Bay NERR/Mississippi State University).

In addition to meteorological data, the NERR SWMP program also generates water quality and nutrient data sets that are available for use.

II. Physical Structure Descriptors

9) Sensor specifications

Parameter: Temperature

Units: Celsius

Sensor type: Platinum resistance temperature detector (PRT)

Model #:  HMP45C Temperature and Relative Humidity Probe

Operating Temperature:  -40°C to +60°C

Range: -40°C to +60°C

Accuracy: ± 0.2 °C @ 20°C

Date of Last calibration:

Sensor installed 12/9/11: s/n C1430020, calibrated 9/28/11

Parameter: Relative Humidity

Units: Percent

Sensor type: Vaisala HUMICAP© 180 capacitive relative humidity sensor

Model #:  HMP45C Temperature and Relative Humidity Probe

Range:  0-100% non-condensing

Accuracy at 20°C:  +/- 2% RH (0-90%) and +/- 3% (90-100%)

Temperature dependence of RH measurement: +/- 0.05% RH/°C

Date of Last calibration:

Sensor installed 12/9/11: s/n C1430020, calibrated 9/28/11

Parameter: Barometric Pressure

Units: millibars (mb)

Sensor type: Vaisala Barocap © silicon capacitive pressure sensor

Model #: CS-105 OR CS-106

CS-105 specifications:

Operating Range:  Pressure: 600 to 1060 mb; Temperature:  -40°C to +60°C;

Humidity:  non-condensing

Accuracy: ± 0.5 mb @ 20°C; +/- 2 mb @ 0°C to 40°C; +/- 4 mb @ -20°C to 45°C; +/- 6 mb @ -40°C to 60°C

Stability: ± 0.1 mb per year

CS-106 specifications:

Operating Range:  Pressure: 500 to 1100 mb; Temperature:  -40°C to +60°C;

Humidity:  non-condensing

Accuracy: Accuracy: ±0.3 mb @ +20°C; ±0.6 mb @ 0° to 40°C; ±1.0 mb @ -20° to +45°C; ±1.5 mb @ -40° to +60°C

Stability: ± 0.1 mb per year

Date of Last calibration:

Sensor installed 12/9/11: CS-105, s/n Y0930011, calibrated 10/6/11

Parameter: Wind speed

Units: meters per second (m/s)

Sensor type: 18 cm diameter 4-blade helicoids propeller molded of polypropylene

Model #: R.M. Young 05103 Wind Monitor

Range:  0-60 m/s (130 mph); gust survival 100 m/s (220 mph)

Accuracy: +/- 2%

Date of last calibration:

Sensor installed 12/9/11: s/n 55459, calibrated 9/30/11

Parameter: Wind direction

Units: degrees

Sensor type: balanced vane, 38 cm turning radius

Model #: R.M. Young 05103 Wind Monitor

Range:  360° mechanical, 355° electrical (5° open)

Accuracy: +/- 5%

Date of last calibration:

Sensor installed 12/9/11: s/n 55459, calibrated 9/30/11

Parameter: Photosynthetically Active Radiation (PAR)

Units: mmoles m-2 (total flux)

Sensor type: LI-COR Quantum Sensor: High stability silicon photovoltaic detector (blue enhanced)

Model #: LI190SB

Light spectrum waveband:  400 to 700 nm

Temperature dependence: 0.15% per °C maximum

Stability:  <±2% change over 1 yr

Operating Temperature:  -40°C to 65°C; Humidity: 0 to 100%

Sensitivity:  typically 5 µA per 1000 µmoles s-1 m-2

Sensor installed 12/9/11: s/n Q35327, calibrated 10/4/11, multiplier 1.73

New sensor purchased 08/2012; installed 9/14/12: s/n Q47542, calibrated 4/5/12, multiplier 0.897

Parameter: Precipitation

Units: millimeters (mm)

Sensor type: Tipping Bucket Rain Gauge

Model #:  TE525

Rainfall per tip: 0.01 inch

Operating range: Temperature: 0° to +/- 50°C; Humidity: 0 to 100%

Accuracy:  +/- 1.0% up to 1 in./hr; +0, -3% from 1 to 2 in./hr; +0, -5% from 2 to 3 in./hr

s/n 36566-705

Date of Last calibration:  9/29/11; 11/1/12

The CR1000 has 2 MB Flash EEPROM that is used to store the Operating System. Another 128 K Flash is used to store configuration settings. A minimum of 2 MB SRAM is (4 MB optional) is available for program storage (16K), operating system use, and data storage. Additional storage is available by using a compact flash card in the optional CFM100 Compact Flash Module.

Date CR1000 installed:  June 21, 2006.  Water damage occurred in September of 2010 and a CR1000 borrowed from the CDMO was installed.

Repaired CR1000 installed 3/19/11.

CR1000 S/N: 5145

CR1000 Firmware Version: 2/17/2012 installed OS24; previously OS20

CR1000 Program Versions: GNDCRMET_6.0_112911.CR1, GNDCRMET_6.0_091312.CR1

10) Coded variable definitions:

Sampling station:		Sampling site code:		Station Code:

Crooked Bayou		CR				gndcrmet

11) QAQC flag definitions

QAQC flags provide documentation of the data and are applied to individual data points by insertion into the parameter’s associated flag column (header preceded by an F_).   During primary automated QAQC (performed by the CDMO), -5, -4, and -2 flags are applied automatically to indicate data that is above or below sensor range or missing.  All remaining data are then flagged 0, as passing initial QAQC checks.  During secondary and tertiary QAQC 1, -3, and 5 flags may be used to note data as suspect, rejected due to QAQC, or corrected.

-5	Outside High Sensor Range

-4	Outside Low Sensor Range

-3	Data Rejected due to QAQC

-2	Missing Data

-1	Optional SWMP supported parameter

0	Passed Initial QAQC Checks

1	Suspect Data

2	Open – reserved for later flag

3	Open – reserved for later flag

4	Historical Data:  Pre-Auto QAQC

5	Corrected Data

12)  QAQC code definitions

QAQC codes are used in conjunction with QAQC flags to provide further documentation of the data and are also applied by insertion into the associated flag column.  There are three (3) different code categories, general, sensor, and comment.  General errors document general problems with the CR1000, sensor errors are sensor specific, and comment codes are used to further document conditions or a problem with the data.  Only one general or sensor error and one comment code can be applied to a particular data point, but some comment codes (marked with an * below)  can be applied to the entire record in the F_Record column.

General Errors

GIM	Instrument Malfunction

GIT	Instrument Recording Error, Recovered Telemetry Data

GMC	No Instrument Deployed due to Maintenance/Calibration

GMT 	Instrument Maintenance

GPD	Power Down

GPF	Power Failure / Low Battery

GPR	Program Reload

GQR	Data Rejected Due to QA/QC Checks

GSM	See Metadata

Sensor Errors

SDG	Suspect due to sensor diagnostics

SIC	Incorrect Calibration Constant, Multiplier or Offset

SIW	Incorrect Wiring

SMT	Sensor Maintenance

SNV	Negative Value

SOC	Out of Calibration

SQR	Data rejected due to QAQC checks

SSN	Not a Number / Unknown Value

SSM	Sensor Malfunction

SSR	Sensor Removed

Comments

CAF	Acceptable Calibration/Accuracy Error of Sensor

CDF	Data Appear to Fit Conditions

CML	Snow melt from previous snowfall event

CRE*	Significant Rain Event

CSM*	See Metadata

CCU	Cause Unknown

CVT*	Possible Vandalism/Tampering

CWE*	Significant weather event

13) Other remarks/notes:

Data are missing due to equipment or associated specific sensors not being deployed, equipment failure, time of maintenance or calibration of equipment, or repair/replacement of a sampling station platform.  Any NANs in the dataset stand for “not a number” and are the result of low power, disconnected wires, or out of range readings.  If additional information on missing data is needed, contact the Research Coordinator at the reserve submitting the data.

Small negative PAR  values are within range of the sensor and are due to normal errors in the sensor and the CR1000 Datalogger. The Maximum signal noise error for the Licor sensor is +/- 2.214 mmoles/m2 over a 15 minute interval.

Relative Humidity data greater than 100 are within range of the sensor accuracy of +/-3%.  Values above 104% are rejected.

Data recorded for all parameters (with the exception of cumulative precipitation) at the midnight timestamp (00:00) are the 15 minute averages and totals for the 23:45-23:59 time period of the previous day.  Cumulative precipitation data at the midnight timestamp (00:00) are the sum of raw (unrounded) precipitation data from 00:00 to 23:59 of the previous day.  Summing each individual 15-minute total precipitation value from the same period will result in small differences from cumulative precipitation due to rounding.  It is especially important to note how data at the midnight timestamp are recorded when using January 1st and December 31st data.

Relative humidity values >103 started occurring soon after the sensor was calibrated and reinstalled.  We believe corrosion in the sensor head (discovered after sensor removal) caused these frequent high readings.  All readings >103 were rejected (<-3> SQR CSM).  Values 101-103 are flagged as <1> GSM CAF.  Because there is such a large quantity of values >103, we want researchers to be aware that these values should be accounted for in order to avoid skewing data and making our RH seem lower than it really is.  Based on spot checks with a Kestrel, we believe other readings from the sensor are correct.  All other RH data were coded with GSM or CSM.  Users should be aware that Campbell Scientific recommendations are less strict than what the NERRS uses as out of range cutoffs; they recommend that values 101-106 be corrected to 100, and higher values be discarded.

Aging PAR sensors tend to have high nighttime readings, possibly due to moisture intrusion into the sensor body.  Nighttime values higher than 0 but less than the signal noise error of 2.214 were flagged as suspect.  Nighttime values higher than 2.214 were rejected.

PAR sensors on the Grand Bay weather station used to show frequent strings of repeating values. It has been determined that this was caused by the programming of the CR1000 data logger.

old program command: VoltDiff(PAR,1,mV2500,3,True,0,_60Hz,0.897,0)

new program command: VoltDiff(PAR,1,mV25,3,True,0,_60Hz,0.897,0)

The program was updated and sent to the CR1000 on April 17, 2013. The mV25 setting allows measurements to be taken at a finer, more precise scale than mV2500, and has, for the most part, eliminated repeating values from the data set. While values before this program change are not as precise as they could have been, they are not erroneous. All valid PAR values have been flagged <0> (CSM) to draw attention to this issue.

2/17/12  13:20-13:35:  Upgraded CR1000 operating system to OS 24.  Data are missing 13:00 – 13:30.  Data were rejected at 13:45 because of a program update.  Precipitation data were not rejected since no rainfall occurred during these times.

7/18/12:  Weather station stopped transmitting.  On 7/23/12 this was determined to be a problem with the regulator; the CR1000 was not getting power.  Data are missing from 7/18/12 13:15 – 7/23/12 10:30.  Data at 7/23/12 10:45 were rejected since they were not a full 15 minutes of 5-second data.

7/23/12:  When logger started again, cumulative precip was 2.5mm.  This is believed to be carryover, as when it stopped logging on the 18th, that was the value.  Also, total precip once the logger started up again was 0.  Cumulative precip values were rejected for that day.

7/23/12 10:45 – 9/14/12 09:30:  All PAR readings were rejected.  It is believed that the same event that fried the regulator also fried the PAR sensor and the GPS antenna.

8/1/12 08:30: We believe that the data are a full 15 minutes of 5-second data, however, since data are missing at 08:15 and we do not know that for sure, the data are flagged as suspect.

8/13/12 13:15: Data are rejected due to more than likely not being a full 15 minutes of 5-second data.

8/13/12 13:30 – 8/14/12 00:00: Cumulative precipitation values are flagged as suspect due to missing data earlier in the day.  No precipitation was recorded at nearby stations, however, since data are missing, we are not confident that precipitation wasn’t missed.

8/27/12-8/30/12:  Hurricane Isaac and its effects.

9/14/12  09:40-11:45:  Installed new PAR sensor and GPS antenna.  Updated program with PAR multiplier of 0.897 and uploaded:  gndcrmet_v6.0_091312.

9/14/12  12:15 – 9/28/12 10:30: Data are missing, the CR1000 had not been logging since the new program was sent. 9/28/12 10:45 data are rejected because of a program upload.

9/28/12  10:30-11:00:  Installed loaner TX-312.  This did not affect data and is here for documentation only.

9/28/12 11:15-17:15: Telemetered data were used during these timestamps.  There are unexplained random gaps in the transmissions that resulted in missing data.  Data collected at 13:00, 17:00, and 17:30, immediately after the gaps, were rejected due to possibly not being a full 15 minutes of 5-second data.

9/30/12 17:45:  All data were rejected due to possibly not being a full 15 minutes of 5-second data following unexplained missing data.

9/30/12 18:00 – 10/1/12 00:00: Cumulative precipitation data were rejected due to missing data earlier in the day.  Rainfall occurred throughout the day and may have been missed during the gap in the data.

11/01/12  9:45-12:15:  Calibrated rain gauge. Precipitation data corrected to 0. Cumulative precipitation data for rest of day also corrected to 0.

11/16/12 14:00-14:20:  Reinstalled Grand Bay’s TX-312. This did not affect data and is here for documentation only.