Grand Bay (GND) National Estuarine Research Reserve Meteorological Metadata

January 1 – December 31, 2013

Latest Update:  March 28, 2014

Note: This is a provisional metadata document; it has not been authenticated as of its download date.  Contents of this document are subject to change throughout the QAQC process and it should not be considered a final record of data documentation until that process is complete.  Contact the CDMO (cdmosupport@belle.baruch.sc.edu) or Reserve with any additional questions.

I. Data Set & Research Descriptors

1) Principal investigators & contact persons:

Address:		Grand Bay National Estuarine Research Reserve

6005 Bayou Heron Road

Moss Point, MS  39562

Phone:  (228) 475-7047

Fax:      (228) 475-8097

Contact Persons:	Dr. Mark Woodrey, Research Coordinator

E-mail:  msw103@ra.msstate.edu

(228) 475-7047, ext. 107

Kim Cressman, Research Technician

E-mail:  kimberly.cressman@dmr.ms.gov

(228) 475-7047, ext. 119

2) Entry verification:

The meteorological information is sampled every five seconds from each instrument on the weather station and stored on a Campbell Scientific CR1000 data logger.  Data are output to a file in one array: array 15 stores 15 minute average and maximum and minimum data.  The program run by the CR1000 is updated as necessary to reflect changes in PAR and Barometric Pressure sensor multipliers (see Section 13 below for additional program information).  The CR1000 interfaces with the LoggerNet software supplied by Campbell Scientific for data retrieval and conversion.

Data are uploaded from the CR1000 data logger to a Personal Computer (PC compatible).  Files are exported from LoggerNet in a comma-delimited format and uploaded to the CDMO where they undergo automated primary QAQC and become part of the CDMO’s online provisional database.  During primary QAQC, data are flagged if they are missing or out of sensor range.  The edited file is then returned to the Reserve where it is opened in Microsoft Excel and processed using the CDMO’s NERRQAQC Excel macro.  The macro inserts station codes, creates metadata worksheets for flagged data and summary statistics, and graphs the data for review.  It allows the user to apply QAQC flags and codes to the data, append files, and export the resulting data file to the CDMO for tertiary QAQC and assimilation into the CDMO’s authoritative online database.  For more information on QAQC flags and QAQC codes, see Sections 11 and 12.

Kim Cressman is responsible for data management.

3) Research objectives:

The National Estuarine Research Reserve (NERR) System-wide Monitoring Program (SWMP) was designed to fulfill two major overall goals:  1) to support state-specific non-point pollution control programs by establishing local networks of continuous water quality monitoring stations in representative protected estuarine ecosystems and 2) to develop a nation-wide database of baseline environmental conditions in the NERR system of estuaries.  The specific goal of SWMP is to identify and track short-term variability and long-term changes in the integrity and biodiversity of representative estuarine ecosystems and coastal watersheds for the purpose of contributing to effective national, regional, and site specific coastal zone management.  This comprehensive program consists of three phased components:  1) abiotic conditions such as water quality and meteorological monitoring; 2) biodiversity monitoring; and 3) habitat mapping and change analysis.  With the initial focus of phase 1, the NERR SWMP provides data necessary for intra- and inter- site baseline studies, trend analyses, and impact assessments.

The principal objective of meteorological monitoring is to record meteorological information for the Grand Bay NERR that can be used 1) as a vital reference of atmospheric data for various research projects at the reserve -- an integral part of the general NERR mission is to provide a platform for estuarine research, 2) to give meteorological context (atmospheric-forcing) for the half-hourly SWMP water quality data, and other long-term environmental monitoring programs at the reserve (including nutrients and shoreline change), 3) to observe and characterize important events, such as storms, heat and cold waves, droughts and heavy rainfalls, and 4) to detect trends and characterize climate variability over the long-term.

4) Research methods:

The Campbell Scientific weather station samples every five seconds continuously throughout the year.  These data are used by the CR1000 to produce 15 minute averages of air temperature, relative humidity, barometric pressure, rainfall, wind speed, and wind direction; maximums and minimums of air temperature and wind speed and the time of occurrence for each; standard deviation of wind speed; total precipitation, photosynthetically active radiation, and daily cumulative precipitation totals.

The 15 minute Data are collected in the following formats for the CR1000:

Averages from 5-second data:

Air Temperature (°C), Relative Humidity (%), Barometric Pressure (mb), Wind Speed (m/s), Wind Direction (degrees), Battery Voltage (volts)

Maximum and Minimum Air Temperature (°C) and their times from 5-second data (these data are not included in the data set but are available from the Reserve)

Maximum Wind Speed (m/s) and time from 5-second data

Wind Direction Standard Deviation (degrees)

Totals:

Precipitation (mm), PAR (millimoles/m2), and Cumulative Precipitation (mm)

Sensors on the weather station are inspected monthly for damage and/or debris.  During these monthly observations, data taken from a nearby airport and a Kestrel 3000 are used to check the calibrations of the sensors.  Sensors are removed and calibrated on an annual or biannual basis depending on the particular sensor.

Recommended calibration frequency for the MET station sensors:

- Temperature/Humidity- annual recalibration

- Rain Gauge- annual recalibration

- Wind Speed/Direction- annual or bi-annual inspection (depending on the sensor)

- Barometric Pressure- bi-annual recalibration

- PAR- bi-annual recalibration

- CR1000-every 5 years (required beginning 2014, one year initial grace period)

Campbell Scientific data telemetry equipment was installed at the Crooked Bayou station on July 21, 2006 and transmits data to the NOAA GOES satellite, NESDIS ID #3B0190E2.  The transmissions are scheduled hourly and contain four (4) data sets reflecting fifteen minute data sampling intervals.  Upon receipt by the CDMO, the data undergoes the same automated primary QAQC process detailed in Section 2 above.  The near real-time telemetry data become part of the provisional dataset until undergoing secondary and tertiary QAQC and assimilation in the CDMO’s authoritative online database.  Provisional and authoritative data are available at http://cdmo.baruch.sc.edu.

5) Site location and character:

The Grand Bay National Estuarine Research Reserve (GBNERR) is a shallow, microtidal estuary located on the Gulf Coast in southeastern Jackson County, Mississippi.  It lies within the Mississippi Delta province of the Louisianian biogeographic region of the United States as identified by the National Oceanic and Atmospheric Administration.  Currently, no other reserves exist within the Mississippi Delta province of this biogeographic region.  The GBNERR encompasses approximately 7,450 ha (18,400 acres) of state-owned, federally-owned, and privately-owned lands and open water areas.  The primary land owners within the reserve boundaries are the State of Mississippi (3,684 ha; 9,100 ac) and the U.S. Fish and Wildlife Service (2,024 ha; 5,000 ac).  The balance of the property is privately owned (1,740 ha; 4,300 acres).

The Grand Bay meteorological station is located near the center of the Reserve (30° 21.551’N, 88° 25.202’W) on a 3m x 3m platform 10m away from a saltwater marsh at an elevation of 1.5m above mean high tide.  One of four hydrological monitors is located in Point Aux Chenes Bay, 1183m south.  Another monitor is located on the east side of the reserve 6652m north/northeast of the station.  The other two monitors are located on the west side of the reserve 3141m north/northwest and 4127m due west of the station.  Although small clusters of dwarfed trees are present on nearby shell middens, they are not considered to serve as wind blocks, shades, or obstructions to the weather station due to the elevation of the station and the meteorological sensors.

The wind monitor and PAR sensor are mounted onto a crossbar sensor mount oriented in an east/west direction on the top of a 3m triangular aluminum tower.  The wind monitor is mounted on the west side of the crossbar, and the quantum sensor is mounted on the east side of the crossbar.  Lower on the tower, the temperature and relative humidity probe is mounted inside a 12-plate gill radiation shield attached to the tower leg facing the prevailing wind (E/SE) 1.7m (68”) above the tower base.  The barometric pressure sensor is mounted inside the control panel attached to two tower legs facing east/northeast 1.3m (50”) above the tower base.  The rain gauge is mounted at a height of 1.8m (72”) to the east post of the chain link fence surrounding the weather station.

6) Data Collection Period

The Grand Bay NERR weather station originally began collecting data on August 17, 2004.  On August 29, 2005, the platform and weather station were physically removed from the pilings by Hurricane Katrina.  The 10’ x 10’ platform and meteorological station were reconstructed and began collecting data on June 21, 2006.  This metadata is applicable only to data collected from January 1, 2013 00:00 to December 31, 2013 23:45.

7) Distribution

NOAA/ERD retains the right to analyze, synthesize and publish summaries of the NERRS System-wide Monitoring Program data.  The PI retains the right to be fully credited for having collected and processed the data.  Following academic courtesy standards, the PI and NERR site where the data were collected will be contacted and fully acknowledged in any subsequent publications in which any part of the data are used.  Manuscripts resulting from this NOAA/OCRM supported research that are produced for publication in open literature, including refereed scientific journals, will acknowledge that the research was conducted under an award from the Estuarine Reserves Division, Office of Ocean and Coastal Resource Management, National Ocean Service, National Oceanic and Atmospheric Administration.  The data set enclosed within this package/transmission is only as good as the quality assurance and quality control procedures outlined by the enclosed metadata reporting statement.  The user bears all responsibility for its subsequent use/misuse in any further analyses or comparisons.  The Federal government does not assume liability to the Recipient or third persons, nor will the Federal government reimburse or indemnify the Recipient for its liability due to any losses resulting in any way from the use of this data.

NERR weather data and metadata can be obtained from the Research Coordinator at the individual NERR site (please see Principal investigators and contact persons), from the Data Manager at the Centralized Data Management Office (please see personnel directory under the general information link on the CDMO home page) and online at the CDMO home page http://cdmo.baruch.sc.edu/.  Data are available in comma separated format.

8) Associated researchers and projects:

Several research and monitoring projects are currently using the meteorological data from the Grand Bay NERR.

In addition to meteorological data, the NERR SWMP program also generates water quality and nutrient data sets that are available for use.

- Water Quality in Bangs Lake: effects of recurrent phosphate spills to a coastal estuary – Dr. Kevin Dillon (Gulf Coast Research Laboratory, University of Southern Mississippi), Dr. Mark Woodrey (Grand Bay NERR/Mississippi State University), Kim Cressman (Grand Bay NERR), Larissa Graham (Grand Bay NERR), Dr. Ruth Carmichael (Dauphin Island Sea Lab), Dr. Jane Caffrey (University of West Florida)

- Effects of Landscape-Level Metrics on Marsh Bird Occupancy and Abundance within Tidal Marshes of the Northern Gulf of Mexico – Dr. Bob Cooper (University of Georgia), Ali Leggett (Mississippi Department of Marine Resources), Dr. Mark Woodrey (Grand Bay NERR/Mississippi State University)

- Shorebird and other waterbird use of wetlands and aquaculture ponds in the Mississippi Alluvial Valley and Gulf Coast region – Dr. Francisco Vilella (Mississippi State University)

- Erosion Monitoring - 6 Monitoring Stations Used to Estimate the Rate of Erosion at Sites Representing Varying Degrees of Wave Exposure and Geological Substrates - Tom Strange (Grand Bay NERR), Jay McIlwain (Grand Bay NERR), Brenna Ehmen (Grand Bay NERR), Lindsay Spurrier (Grand Bay NERR), Will Underwood (Grand Bay NERR)

- Initial observations of Colored Dissolved Organic Matter (CDOM) in the Grand Bay National Estuarine Research Reserve (GBNERR), Southeastern Mississippi – Kim Cressman (Grand Bay NERR), Christina Mohrman (Florida A&M University/Grand Bay NERR)

- Spectroscopic Characterization of Chromophoric Dissolved Organic Matter (CDOM) in the Grand Bay National Estuarine Research Reserve – Dr. Daniel Osborne (NOAA Environmental Cooperative Science Center, Florida A&M University), Kim Cressman (Grand Bay NERR), Christina Mohrman (NOAA Environmental Cooperative Science Center, Florida A&M University / Grand Bay NERR), Dr. Charles Jagoe (NOAA Environmental Cooperative Science Center, Florida A&M University)

- Legacy Effects of Land-use Change and Nitrogen Source Shifts on a Benchmark System: Building Capacity for Collaborative Research Leadership at the Grand Bay Reserve - Lead PI Dr. Ruth Carmichael (Dauphin Island Sea Lab, University of South Alabama); collaborators: Elizabeth Darrow (PhD student, Dauphin Island Sea Lab, University of South Alabama), Capt William Burkhardt (USFDA Office of Food Safety), Cdr Kevin Calci (USFDA Office of Food Safety), Dr. Wei Wu (Gulf Coast Research Laboratory, University of Southern Mississippi), Dave Ruple (Grand Bay NERR), Dr. William Walton (Auburn University Shellfish Laboratory)

- Monitoring breeding marsh bird populations at the Grand Bay NERR and Pascagoula River Marshes Coastal Preserve – Dr. Mark Woodrey (Grand Bay NERR/Mississippi State University) Dr. Bob Cooper (University of Georgia), Dr. Scott Rush (Mississippi State University)

- Degradation Following Watershed Land Use Changes - Dr. Mark Woodrey (Grand Bay NERR), Brenna Ehmen (Grand Bay NERR),Dr. Just Cebrian (Dauphin Island Sea Lab), Daniel Grigas (University of Arkansas at Pine Bluff), Dr. Yushun Chen (University of Arkansas at Pine Bluff ), Dr. Russ Beard (NOAA), Dr. A.R. Parsons (NOAA)

- Fish Communities of Nearshore Habitats within the Grand Bay NERR/NWR – Ron Cole (Grand Bay NERR), Brenna Ehmen (Grand Bay NERR), Gretchen Grammer (Grand Bay NERR), Jake Walker (Grand Bay NERR), Dr. Mark Woodrey (Grand Bay NERR/Mississippi State University);

- Nesting ecology of the diamondback terrapin (Malaclemys terrapin pileata) at the Grand Bay National Estuarine Research Reserve, Mississippi - Collaborators - Christina Mohrman (Florida A&M University/Grand Bay NERR) and Dr. Mark Woodrey (Mississippi State University/Grand Bay NERR);

- Atmospheric Mercury Deposition Monitoring at the Grand Bay NERR - collaborators: Dr. Winston Luke (Air Resources Lab - NOAA), Dr. Mark Cohen (Air Resources Lab - NOAA), Mr. Dave Ruple (Grand Bay NERR), Dr. Mark Woodrey (Grand Bay NERR/Mississippi State University).

- Characterizing Marsh Soil Contaminants of Runoff and Watershed Loading Affecting the Ecological Systems of Bangs Lake, Grand Bay NERR – Shelton Clerk (Ph.D. student at Jackson State University; ECSC)

II. Physical Structure Descriptors

9) Sensor specifications

Parameter: Temperature

Units: Celsius

Sensor type: Platinum resistance temperature detector (PRT)

Model #:  HMP45C Temperature and Relative Humidity Probe

Operating Temperature:  -40°C to +60°C

Range: -40°C to +60°C

Accuracy: ± 0.2 °C @ 20°C

Date of Last calibration:

Sensor installed 12/9/11 – 12/3/13: s/n C1430020, calibrated 9/28/11

Sensor installed 12/3/13 – present as of 12/31/13: s/n A3510149, calibrated 10/29/13

Parameter: Relative Humidity

Units: Percent

Sensor type: Vaisala HUMICAP© 180 capacitive relative humidity sensor

Model #:  HMP45C Temperature and Relative Humidity Probe

Range:  0-100% non-condensing

Accuracy at 20°C:  +/- 2% RH (0-90%) and +/- 3% (90-100%)

Temperature dependence of RH measurement: +/- 0.05% RH/°C

Date of Last calibration:

Sensor installed 12/9/11 – 12/3/13: s/n C1430020, calibrated 9/28/11

Sensor installed 12/3/13 – present as of 12/31/13: s/n A3510149, calibrated 10/29/13

Parameter: Barometric Pressure

Units: millibars (mb)

Sensor type: Vaisala Barocap © silicon capacitive pressure sensor

Model #: CS-105 OR CS-106

CS-105 specifications:

Operating Range:  Pressure: 600 to 1060 mb; Temperature:  -40°C to +60°C;

Humidity:  non-condensing

Accuracy: ± 0.5 mb @ 20°C; +/- 2 mb @ 0°C to 40°C; +/- 4 mb @ -20°C to 45°C; +/- 6 mb @ -40°C to 60°C

Stability: ± 0.1 mb per year

CS-106 specifications:

Operating Range:  Pressure: 500 to 1100 mb; Temperature:  -40°C to +60°C;

Humidity:  non-condensing

Accuracy: Accuracy: ±0.3 mb @ +20°C; ±0.6 mb @ 0° to 40°C; ±1.0 mb @ -20° to +45°C; ±1.5 mb @ -40° to +60°C

Stability: ± 0.1 mb per year

Date of Last calibration:

Sensor installed 12/9/11 – 12/3/13: CS-105, s/n Y0930011, calibrated 10/6/11

Sensor installed 12/3/13 – present as of 12/31/13: CS-106, s/nF4030089, calibrated 11/4/13

Parameter: Wind speed

Units: meters per second (m/s)

Sensor type: 18 cm diameter 4-blade helicoids propeller molded of polypropylene

Model #: R.M. Young 05103 Wind Monitor

Range:  0-60 m/s (130 mph); gust survival 100 m/s (220 mph)

Accuracy: +/- 2%

Date of last calibration:

Sensor installed 12/9/11 – 12/3/13: s/n 55459, calibrated 9/30/11

Sensor installed 12/3/13 – present as of 12/31/13: s/n 68641, calibrated 11/14/13

Parameter: Wind direction

Units: degrees

Sensor type: balanced vane, 38 cm turning radius

Model #: R.M. Young 05103 Wind Monitor

Range:  360° mechanical, 355° electrical (5° open)

Accuracy: +/- 5%

Date of last calibration:

Sensor installed 12/9/11 – 12/3/13: s/n 55459, calibrated 9/30/11

Sensor installed 12/3/13 – present as of 12/31/13: s/n 68641, calibrated 11/14/13

Parameter: Photosynthetically Active Radiation (PAR)

Units: mmoles m-2 (total flux)

Sensor type: LI-COR Quantum Sensor: High stability silicon photovoltaic detector (blue enhanced)

Model #: LI190SB

Light spectrum waveband:  400 to 700 nm

Temperature dependence: 0.15% per °C maximum

Stability:  <±2% change over 1 yr

Operating Temperature:  -40°C to 65°C; Humidity: 0 to 100%

Sensitivity:  typically 5 µA per 1000 µmoles s-1 m-2

Sensor purchased 08/2012; installed 9/14/12 – present as of 12/31/13: s/n Q47542, calibrated 4/5/12, multiplier 0.897

Parameter: Precipitation

Units: millimeters (mm)

Sensor type: Tipping Bucket Rain Gauge

Model #:  TE525

Rainfall per tip: 0.01 inch

Operating range: Temperature: 0° to +/- 50°C; Humidity: 0 to 100%

Accuracy:  +/- 1.0% up to 1 in./hr; +0, -3% from 1 to 2 in./hr; +0, -5% from 2 to 3 in./hr

s/n 36566-705

Dates of calibration:  11/1/12; 10/10/13

The CR1000 has 2 MB Flash EEPROM that is used to store the Operating System. Another 128 K Flash is used to store configuration settings. A minimum of 2 MB SRAM is (4 MB optional) is available for program storage (16K), operating system use, and data storage. Additional storage is available by using a compact flash card in the optional CFM100 Compact Flash Module.

Date CR1000 originally installed:  June 21, 2006.  Water damage occurred in September of 2010 and a CR1000 borrowed from the CDMO was installed.

Repaired CR1000 installed 3/19/11.

CR1000 Firmware Version (s):  OS24 installed 2/17/2012; current as of 12/31/2013

CR1000 Program Version(s):

gndcrmet_v6.0_091312.cr1 run  9/14/12 - 1/8/13

gndcrmet_v6.1_010713.cr1 run  1/8/13 - 4/17/13

gndcrmet_v6_032813.cr1  run 4/17/13 - 10/10/13

gndcrmet_v6_070213.cr1 run 10/10/13 - 12/3/13

gndcrmet_6.0_111813.cr1 running 12/3/13 – present as of 12/31/2013

10) Coded variable definitions:

Sampling station:		Sampling site code:		Station Code:

Crooked Bayou		CR				gndcrmet

11) QAQC flag definitions

QAQC flags provide documentation of the data and are applied to individual data points by insertion into the parameter’s associated flag column (header preceded by an F_).   During primary automated QAQC (performed by the CDMO), -5, -4, and -2 flags are applied automatically to indicate data that is above or below sensor range or missing.  All remaining data are then flagged 0, as passing initial QAQC checks.  During secondary and tertiary QAQC 1, -3, and 5 flags may be used to note data as suspect, rejected due to QAQC, or corrected.

-5	Outside High Sensor Range

-4	Outside Low Sensor Range

-3	Data Rejected due to QAQC

-2	Missing Data

-1	Optional SWMP supported parameter

0	Passed Initial QAQC Checks

1	Suspect Data

2	Open – reserved for later flag

3	Open – reserved for later flag

4	Historical Data:  Pre-Auto QAQC

5	Corrected Data

12)  QAQC code definitions

QAQC codes are used in conjunction with QAQC flags to provide further documentation of the data and are also applied by insertion into the associated flag column.  There are three (3) different code categories, general, sensor, and comment.  General errors document general problems with the CR1000, sensor errors are sensor specific, and comment codes are used to further document conditions or a problem with the data.  Only one general or sensor error and one comment code can be applied to a particular data point, but some comment codes (marked with an * below)  can be applied to the entire record in the F_Record column.

General Errors

GIM	Instrument Malfunction

GIT	Instrument Recording Error, Recovered Telemetry Data

GMC	No Instrument Deployed due to Maintenance/Calibration

GMT 	Instrument Maintenance

GPD	Power Down

GPF	Power Failure / Low Battery

GPR	Program Reload

GQR	Data Rejected Due to QA/QC Checks

GSM	See Metadata

Sensor Errors

SDG	Suspect due to sensor diagnostics

SIC	Incorrect Calibration Constant, Multiplier or Offset

SIW	Incorrect Wiring

SMT	Sensor Maintenance

SNV	Negative Value

SOC	Out of Calibration

SQR	Data rejected due to QAQC checks

SSN	Not a Number / Unknown Value

SSM	Sensor Malfunction

SSR	Sensor Removed

Comments

CAF	Acceptable Calibration/Accuracy Error of Sensor

CDF	Data Appear to Fit Conditions

CML	Snow melt from previous snowfall event

CRE*	Significant Rain Event

CSM*	See Metadata

CCU	Cause Unknown

CVT*	Possible Vandalism/Tampering

CWE*	Significant weather event

13) Other remarks/notes:

Data are missing due to equipment or associated specific sensors not being deployed, equipment failure, time of maintenance or calibration of equipment, or repair/replacement of a sampling station platform.  Any NANs in the dataset stand for “not a number” and are the result of low power, disconnected wires, or out of range readings.  If additional information on missing data is needed, contact the Research Coordinator at the reserve submitting the data.

Small negative PAR  values are within range of the sensor and are due to normal errors in the sensor and the CR1000 Datalogger. The Maximum signal noise error for the Licor sensor is +/- 2.214 mmoles/m2 over a 15 minute interval.

Relative Humidity data greater than 100 are within range of the sensor accuracy of +/-3%.  Values above 104% are rejected.

Data recorded for all parameters (with the exception of cumulative precipitation) at the midnight timestamp (00:00) are the 15 minute averages and totals for the 23:45-23:59 time period of the previous day.  Cumulative precipitation data at the midnight timestamp (00:00) are the sum of raw (unrounded) precipitation data from 00:00 to 23:59 of the previous day.  Summing each individual 15-minute total precipitation value from the same period will result in small differences from cumulative precipitation due to rounding.  It is especially important to note how data at the midnight timestamp are recorded when using January 1st and December 31st data.

Aging PAR sensors tend to have high nighttime readings, possibly due to moisture intrusion into the sensor body.  Nighttime values higher than 0 but less than the signal noise error of 2.214 were flagged as suspect.  Nighttime values higher than 2.214 were rejected.

PAR sensors on the Grand Bay weather station used to show frequent strings of repeating values. It has been determined that this was caused by the programming of the CR1000 data logger.

old program command: VoltDiff(PAR,1,mV2500,3,True,0,_60Hz,0.897,0)

new program command: VoltDiff(PAR,1,mV25,3,True,0,_60Hz,0.897,0)

The program was updated and sent to the CR1000 on April 17, 2013. The mV25 setting allows measurements to be taken at a finer, more precise scale than mV2500, and has eliminated repeating values from the data set. While values before this program change are not as precise as they could have been, they are not erroneous. All valid PAR values have been flagged <0> (CSM) to draw attention to this issue.

1/8/2013 12:45 – new program uploaded: gndcrmet_v6_R.1_010713.cr1

This program has instructions for a GoesMetrics table, which logs communication attempts between the CR1000 logger and the TX312 transmitter. It will log on an hourly basis and store up to 300 records. It will overwrite starting at the oldest record. This table only needs to be downloaded when a problem is suspected.

4/17/2013 12:45 – new program uploaded: gndcrmet_v6_032813.cr1

This program changed a line in the PAR instructions from mV2500 to mV25. This resolved the repeating values problem.

10/10/2013 10:00-10:30 – Total PAR values rejected due to installation of Apogee PAR sensor adjacent to LiCor. The LiCor may have been shaded during this installation.

10/10/2013 10:45-11:00 – new program uploaded: gndcrmet_v6_07_02_13_APAR.cr1

An Apogee brand PAR sensor was installed on the weather station next to the LiCor PAR sensor in an effort to determine whether data from the two sensors is comparable. The new program makes a table that includes values from both sensors. Apogee data will not be included in the SWMP dataset.

- To install the Apogee adjacent to the LiCor without it being shaded by the GPS antenna, the GPS antenna was moved.

10/10/13 11:00 – 13:00 – calibrating rain gauge. Total Precip and Cum Precip (until 10/11 0:00) corrected to 0.

12/3/2013 12:41 – 13:28 – station was powered down for installation of fresh Temp/RH, BP, and Wind sensors. New program with code for CS-106 pressure sensor sent to station at 13:32.