Narragansett Bay (NAR) National Estuarine Research Reserve

Water Quality Metadata

January to June 2017

Last updated:  Friday, July 28, 2017

Note: This is a provisional metadata document; it has not been authenticated as of its download date.  Contents of this document are subject to change throughout the QAQC process and it should not be considered a final record of data documentation until that process is complete.  Contact the CDMO (cdmosupport@belle.baruch.sc.edu) or the Narragansett Bay NERR with any additional questions.

# I.  Data Set and Research Descriptors

## 1.  Principal investigators and contact persons

Dr. Kenneth B. Raposa

Research Coordinator

E-mail: kenny@nbnerr.org

Ph (401) 683-7849

Dr. Daisy Durant, Marine Research Specialist II

SWMP Coordinator and Data Manager

E-mail: daisy@nbnerr.org

Ph (401) 683-7368

Address:	Narragansett Bay NERR

55 South Reserve Drive

PO Box 151

Prudence Island, RI 02872

Phone:  (401) 683-6780

Fax:  (401) 682-7366

## 2.  Entry Verification

Deployment data are uploaded from the YSI data logger to a Personal Computer (IBM compatible).  Files are exported from EcoWatch in a comma-delimited format (.CDF), EcoWatch Lite in a comma separated file (CSV) or KOR Software in an Excel File (.XLS) and uploaded to the CDMO where they undergo automated primary QAQC; automated depth/level corrections for changes in barometric pressure (cDepth or cLevel parameters); and become part of the CDMO’s online provisional database.  All pre- and post-deployment data are removed from the file prior to upload.  During primary QAQC, data are flagged if they are missing or out of sensor range.  The edited file is then returned to the Reserve for secondary QAQC where it is opened in Microsoft Excel and processed using the CDMO’s NERRQAQC Excel macro.  The macro inserts station codes, creates metadata worksheets for flagged data and summary statistics, and graphs the data for review.  It allows the user to apply QAQC flags and codes to the data, remove any overlapping deployment data, append files, and export the resulting data file for upload to the CDMO.  Upload after secondary QAQC results in ingestion into the database as provisional plus data, recalculation of cDepth or cLevel parameters, and finally tertiary QAQC by the CDMO and assimilation into the CDMO’s authoritative online database.  Where deployment overlap occurs between files, the data produced by the newly calibrated sonde is generally accepted as being the most accurate.  For more information on QAQC flags and codes, see Sections 11 and 12.

Dr. Daisy Durant (Marine Research Specialist II) was responsible for compiling and error checking the January through March 2017 water quality data covered in this document.

## 3.  Research Objectives

YSI 6600-V2 and EXO2 data loggers are being deployed off Prudence Island in Narragansett Bay as part of the Narragansett Bay National Estuarine Research Reserve's (NBNERR) System-Wide Monitoring Program (SWMP).  The goal is to develop long-term data sets for representative estuarine systems in order to track changes in water quality over time.  Because Prudence Island is located in the geographic center of Narragansett Bay, it is an ideal location for monitoring the status and trends in water quality in the Bay over time.  One NBNERR water quality monitoring station has been established at Potter Cove since December 1995, on the Island's northeastern shore.  This area is impacted by boat traffic and storm runoff from mainland urban and residential areas.  The second NBNERR water quality monitoring station, T-Wharf, was established on September 1996 and is situated on the southeastern shore of the island, facing the open waters of Rhode Island Sound.  It is approximately 6 miles south of the Potter Cove site. Boat traffic is sparse at this site and storm runoff is less likely to have a significant impact on water quality.  A third monitoring site was added in March of 2002.  This monitoring site is located in Nag Creek, a salt marsh tidal creek that flows into the West Passage of Narragansett Bay.  The addition of this site completes our representation of dominant habitat types occurring in Narragansett Bay (i.e. marsh, cove, and open water).  In July of 2002, the T-Wharf monitoring station was replaced with two new monitoring sites located a short distance from the original T-Wharf location.  The new stations are situated on either side of a wharf support piling.  One data logger records water quality near the surface (approximately 1.0 m deep) while the second records water quality parameters approximately 1.0 m off the bottom.  This allows for the identification of both the frequency and duration of any stratification, which may occur in the open waters of Narragansett Bay.

## 4.  Research Methods

Calibrating procedures

One data logger (also known as sonde) is deployed at each permanent monitoring station at the Reserve on Prudence Island (see map on Section 5).  The instruments are deployed for approximately two to four weeks at a time (depending on the season), at the end of which are retrieved and newly calibrated sondes are deployed instead.  During these extended deployment, the sondes collect data on water temperature, specific conductance, dissolved oxygen, pH, depth, turbidity conditions, and chlorophyll fluorescence every 15 minutes.  Historically, data from all stations were collected at 30-minute sampling intervals.  In 2004, the NB NERR became involved with a statewide fixed site water quality monitoring program at which time data collection at all stations were changed to a 15-minute sampling interval in order to be compliant with this local monitoring effort.  Two years later, all Reserves within the Reserve System were required also required to collect at 15-minutes sampling intervals.

For each sampling period, each instrument is calibrated against known standards following the methods recommended in the YSI 6-Series Multi-Parameter Water Quality Monitoring Standard Operating Procedure manual and the EXO User Manual-Advanced Water Quality Monitoring Platform.  The specific conductivity probe is calibrated using conductivity calibrator 50,000 µS/cm; a three-point calibration is done for pH using pH buffer 4, 7, and 10; a two point calibration (0, 126 with YSI 6600-V2 sonde or 0, 124 with EXO2 sondes) is done for turbidity using deionized water and YSI 6073G turbidity standard; a one-point calibration is done in air-saturated water for dissolved oxygen (DO); and a one-point calibration (0.0) with deionized water is done for chlorophyll fluoresces (data available only upon request).  The sonde is programmed to begin taking measurements approximately 10 hours in advance of planned deployment, allowing the DO membrane to stabilize.  The file is checked for DO sensor drift before the DO probe is calibrated.  Calibration of the dissolved oxygen sensor is usually done within 2 hours of deployment.  The V2 and EXO2 sondes have optical DO probes (DO ROX) and membranes are changed annually or when readings become unstable; this sensor can be calibrated at the same time as the other sensors; however, it is usually done within 2 hours of deployment in air-saturated water.  At the end of each sampling period, the data loggers are retrieved and freshly calibrated instruments are deployed.  The retrieved data loggers are brought back to the laboratory, post-deployment evaluations of the instrument (Section 13) are carried out the same day of retrieval, and the device is cleaned and serviced by methods outlined in the service manuals mentioned above.

During every deployment, sea-truthing is conducted by measuring water quality parameters with an additional data logger taking real-time measurements at the time of deployment.  This information is used to compare with the data collected by the sonde to ensure accuracy of the readings.  The data is recorded using a hand-held YSI 556-MPS data logger.  Values for temperature, specific conductance (and salinity), dissolved oxygen (% saturation and mg/L), and pH are recorded real-time at either 0.5 or 1.0 meter intervals between the surface and bottom at each deployment site. The calibration procedure for the YSI 556 used in sea-truthing follows the same general procedure as that for the sondes used for extended deployments.  The YSI 556 specific conductivity and pH sensors are calibrated the day before a planned deployment; dissolved oxygen is calibrated the day of deployment in air-saturated water.  Data and calibrating logs for the YSI 556 data logger are kept at the Reserve and available upon request.

Site Infrastructure

The monitoring station at Potter Cove was originally constructed in 1995, and it consisted of a PVC pipe mounted vertically on a piling located approximately 6 feet west of a floating dock.  To facilitate water flow across the sensors, openings were cut into the PVC pipes.  The pipes were positioned to ensure that the sensors were less than 1 meter from the bottom.  In 2006, the infrastructure at Potter Cove was replaced.  It currently consists of a short PVC tube attached to an adjacent dock, with a line attached through the pipe to an anchor on the bottom.  The sonde is attached to the anchor via a shackle and hook set-up, and a float is attached to the top of the sonde to keep it approximately 0.75 m off the bottom. With this setup, the sonde is free and clear of any pipe affects since the pipe does not extend to the bottom and serves only as a “decoy pipe” to prevent vandalism.

In 1996, the original T-Wharf station was put into service using a PVC pipe mounted on a piling attached to the wharf.  In 2002, the two replacement T-Wharf sites (Surface and Bottom) were also deployed in PVC pipes mounted to a piling.  The sonde at the T-Wharf Surface station was maintained just below the surface by means of a buoy attached by rope to the adjacent wharf; the sonde at T-Wharf Bottom was maintained approximately 0.5 m off the bottom.  On 10/11/06, the PVC deployment pipes and securing structures were replaced.  The design was similar to the previous pipe design (following YSI recommendations) and consisted of holes drilled through the pipe and large slits at the bottom of the pipe to allow free water movement at the sonde.  The T-Wharf Bottom pipe extends approximately 6 m though the water to the bottom of the site, where the sonde is kept approximately 1.0 m off the bottom.  The T-Wharf Surface pipe extends approximately 2 m under the surface of the water (at low tide), where the sonde was attached to a float to keep it approximately 0.5 m below the surface of the water.  After Tropical Storm Irene in August of 2011, the deploying structure was damaged; the PVC pipe was lost and the piling where it was attached to was damaged beyond repair.  A temporary deploying structure was build consisting of a steel cable attached to the wharf with an anchor at the end to keep the cable straight and in place.  The Surface and Bottom sondes were attached to the cable for each deployment and lowered into the water with a security line.  The Reserve arranged to build a permanent deploying structure; a new piling was installed on February 2012, two 4” diameter PVC pipes were installed in June (following YSI recommendations), and the sondes were deployed in July in these new structures.  The Surface sonde was deployed attached to a float to be approximately 1.0 m below the surface at low tide, and the Bottom sonde 1.0 m above the bottom substrate.  Due to the narrow PVC pipes, the Surface sonde was getting jammed inside and not floating free; thus, the deployment was changed to a fixed one at approximately 0.5 m at low tide.  As standard procedure, the pipes have been cleaned monthly with a chimney brush as needed but mostly during the warm months (April to November) to reduce the impact of biofouling on the data.

The Nag Creek sonde was originally deployed in a metal cage, which was tethered to the bank.  A permanent deployment structure was installed 12/31/02, consisting of a 4” x 4” pressure-treated post with a hinged 2” x 4” horizontal arm.  On 09/08/06, the sonde deployment structure was redesigned and changed.  The new structure consists of an L-shape wooden structure that held in place in the sediment by a large metal tripod sunk into the mud. The sonde is extended from the arm into the water via a cleat, eye and line system and hangs approximately 0.30 m off the bottom of the creek.  In 2012, the deployment structure was changed for a wooden tripod sunken into the mud, which facilitated its removal in the winter when the low temperatures freeze up the creek.

Real-time data transmission

A Sutron Sat-Link2 transmitter was installed at the T-Wharf Bottom monitoring station on 07/27/06 and transmits data to the NOAA GOES satellite, NESDIS ID #3B0335EE.  The transmissions are scheduled hourly and contain four (4) data sets reflecting fifteen minute data sampling intervals.  Upon receipt by the CDMO, the data undergoes the same automated primary QAQC process detailed in Section 2 above.  The “real-time” telemetry data become part of the provisional dataset until undergoing secondary and tertiary QAQC and assimilation in the CDMO’s authoritative online database.  Provisional and authoritative data are available at http://cdmo.baruch.sc.edu.

## 5.  Site Location and Character

The NBNERR is located close to the geographic center of Narragansett Bay in Rhode Island.  It consists of approximately 1802 hectares (4453 acres: 2544 acres of land plus 1909 acres of surrounding estuarine waters, approximately) of diverse estuarine and terrestrial habitats ranging from deep water to salt marshes to forested uplands.  The land holdings include approximately 65% of Prudence Island, most of nearby Patience Island, and all of Hope and Dyer Islands (see map below).  The most recent land acquisition was the Eugene Chase Farm property, made official by the end of year 2015.

The Narragansett Bay watershed consists of nine subwatersheds draining an area of approximately 4,836 square km (Pilson, 1985) and numerous and substantial freshwater inputs to the Bay.  Approximately 39% of the watershed lies in Rhode Island and 61% in Massachusetts.  It is referred to as a shallow estuary; however, its water depth varies considerably.  Depth averages approximately 9.0 m throughout the Bay, but it is deeper in the East Passage (approximately 15.2 m) and shallower in the West Passage (approximately 7.5 m).  More information and a detailed description of the Narragansett Bay NERR and the Narragansett Bay watershed can be found in Raposa and Schwartz, available at http://nbnerr.org/research-and-monitoring/publications.

Specific characteristics of the Narragansett Bay National Estuarine Research Reserve

Location:	41°38’ 30” N, 71°20’30” W

Tidal range:	-0.2 to 1.7 meters MLW

Salinity:  	15 to 32 ppt

Temperature:  	-1.0 to 26 C

Province:  	North temperate, Virginian bioregion

Specific characteristics of the Nag Creek site are:

Location:  	41° 37’ 29.458” N, 71° 19’ 27.421” W

Depth:  	0.1 to 1.4 meters

Bottom habitat: 	Organic mud

Pollutants:  	Negligible

Watershed:  	Narragansett Bay, West Passage

Specific characteristics of the Potter Cove site are:

Location:  	41° 38’ 25.984” N, 71° 20’ 27.165” W

Depth:  	0.9 to 3.9 meters

Bottom habitat: 	Sand, silt, some organic mud

Pollutants:  	Boaters’ wastes, storm runoff from mainland urban areas

Watershed:  	Narragansett Bay, North Prudence (4801 km2)

Specific characteristics of the T-Wharf Surface site are:

Location:  	41° 34’ 42.099” N, 71° 19’ 16.049” W

Depth:  	0.2 to 0.9 meters

Bottom habitat: 	Sand, silt, some organic mud

Pollutants:  	Negligible

Watershed:  	Narragansett Bay, South Prudence

Specific characteristics of the T-Wharf Bottom site are:

Location:  	41° 34’ 42.099” N, 71° 19’ 16.049” W

Depth:  	4.6 to 6.9 meters

Bottom habitat: 	Sand, silt, some organic mud

Pollutants:  	Negligible

Watershed:  	Narragansett Bay, South Prudence

## 6.  Data collection period

Date and time of the first and last readings included in each data file downloaded from the sondes deployed at each of the four water quality stations on Prudence Island are shown below for 2017.

Nag Creek (Data collection ongoing since 2002).

| Deploy Date | Deploy Time | Retrieve Date | Retrieve Time | EXO 2 Sonde Serial Number |
| --- | --- | --- | --- | --- |
| 1 12/20/16 | 11:45 | 01/17/17 | 11:30 | 14A101166 |
| 01/17/17 | 11:45 | 02/14/17 | 11:45 | 14A101167 |
| 2 02/14/17 | 12:00 | 02/22/17 | 11:15 | 14A101166 |
| 02/22/17 | 11:30 | 03/08/17 | 09:45 | 14A101166 |
| 03/08/17 | 10:15 | 04/07/17 | 11:30 | 14A101167 |
| 3 04/07/17 | 12:00 | 04/11/17 | 13:30 |  |
| 04/11/17 | 13:45 | 05/04/17 | 09:15 | 14A101166 |
| 05/04/17 | 09:15 | 05/31/17 | 08:45 | 14A101167 |
| 05/31/17 | 09:00 | 06/22/17 | 08:45 | 14A101166 |
| 06/22/17 | 09:00 | 07/13/17 | 08:45 | 14A101167 |

Notes

1 This deployment started on 12/20/16 11:45, but only data from 2017 are included in this provisional metadata document and corresponding dataset file.

2 No sonde was deployed at Nag during this period because the creek was frozen.

3 No sonde was deployed during this time.

Potter Cove (Data collection ongoing since 1995).

| Deploy Date | Deploy Time | Retrieve Date | Retrieve Time | EXO 2 Sonde Serial Number |
| --- | --- | --- | --- | --- |
| 1 12/20/16 | 12:15 | 01/17/17 | 11:45 | 14A101165 |
| 01/17/17 | 12:15 | 02/15/17 | 11:00 | 14A101164 |
| 02/15/17 | 11:15 | 03/08/17 | 10:30 | 14A101165 |
| 03/08/17 | 10:45 | 04/06/17 | 07:00 | 14A101164 |
| 04/06/17 | 07:15 | 05/04/17 | 09:45 | 14A101165 |
| 05/04/17 | 10:15 | 05/31/17 | 09:15 | 14A101164 |
| 05/31/17 | 09:30 | 06/22/17 | 09:45 | 14A101165 |
| 06/20/17 | 10:15 | 07/13/17 | 08:00 | 14A101164 |

Notes

1 This deployment started on 12/20/16 12:15, but only data from 2017 are included in this provisional metadata document and corresponding dataset file.

T-Wharf Surface (Data collection ongoing since 2002).

| Deploy Date | Deploy Time | Retrieve Date | Retrieve Time | EXO 2 Sonde Serial Number |
| --- | --- | --- | --- | --- |
| 1 12/20/16 | 10:45 | 01/17/17 | 10:15 | 16J102354 |
| 01/17/17 | 10:30 | 02/15/17 | 10:00 | 16J102353 |
| 02/15/17 | 10:15 | 03/07/17 | 10:00 | 16J102354 |
| 03/07/17 | 10:30 | 04/06/17 | 06:15 | 16J102353 |
| 04/06/17 | 06:30 | 05/03/17 | 09:15 | 16J102354 |
| 05/03/17 | 09:45 | 06/01/17 | 08:45 | 16J102353 |
| 06/01/17 | 09:15 | 06/20/17 | 08:30 | 16J102354 |
| 2 06/20/17 | 23:45 | 07/12/17 | 10:00 | 16J102353 |

Notes

1 This deployment started on 12/20/16 10:45, but only data from 2017 are included in this provisional metadata document and corresponding dataset file.

2 The sonde was in the water on 06/20 08:45; however, no data collected from 06/20 08:45 to 23:15 due to a mistake in the sonde programming.

T-Wharf Bottom (Data collection ongoing since 2002).

| Deploy Date | Deploy Time | Retrieve Date | Retrieve Time | 6600V2 Sonde Serial Number |
| --- | --- | --- | --- | --- |
| 1 12/28/16 | 11:15 | 01/17/17 | 10:30 | 05E1065AE |
| 01/17/17 | 11:00 | 02/15/17 | 08:00 | 06C1207AA |
| 02/15/17 | 10:30 | 03/07/17 | 10:15 | 05E1065AE |
| 2 03/07/17 | 11:00 | 03/30/17 | 12:00 | 06C1207AA |
| 03/30/17 | 12:30 | 05/03/17 | 09:15 | 05E1065AE |
| 05/03/17 | 09:45 | 06/01/17 | 09:00 | 06C1207AA |
| 06/01/17 | 09:30 | 06/20/17 | 08:15 | 05E1065AE |
| 06/20/17 | 09:15 | 07/10/17 | 10:15 | 06C1207AA |

Notes

1 This deployment started on 12/28/16 11:15, but only data from 2017 are included in this provisional metadata document and corresponding dataset file.

2 A catastrophic temperature failure occurred during this deployment.

## 7.  Distribution

According to the Ocean and Coastal Resource Management Data Dissemination Policy for the NERRS System-wide Monitoring Program,

NOAA retains the right to analyze, synthesize and publish summaries of the NERRS System-wide Monitoring Program data.  The NERRS retains the right to be fully credited for having collected and process the data.  Following academic courtesy standards, the NERR site where the data were collected should be contacted and fully acknowledged in any subsequent publications in which any part of the data are used.  The data set enclosed within this package/transmission is only as good as the quality assurance and quality control procedures outlined by the enclosed metadata reporting statement.  The user bears all responsibility for its subsequent use/misuse in any further analyses or comparisons.  The Federal government does not assume liability to the Recipient or third persons, nor will the Federal government reimburse or indemnify the Recipient for its liability due to any losses resulting in any way from the use of this data.

Requested citation format:

NOAA National Estuarine Research Reserve System (NERRS).  System-wide Monitoring Program accessed from the NOAA NERRS Centralized Data Management Office website: http://www.nerrsdata.org/; accessed 12 October 2012.

NERR water quality data and metadata can be obtained from the Research Coordinator at the individual NERR site (please see Principal Investigators and Contact Persons), from the Data Manager at the Centralized Data Management Office (please see personnel directory under the general information link on the CDMO home page) and online at the CDMO home page www.nerrsdata.org.  Data are available in text tab-delimited format.

## 8.  Associated researchers and projects

Complementary to the existing long-term water quality monitoring program, the NERRS implemented a new nutrient and chlorophyll monitoring program in 2002.  The two sub-components of this program include monthly grab sampling at each of the four water quality stations, and diel sampling once a month at one site.  The grab sampling program requires the collection of duplicate water samples every month from each of the four long-term water quality monitoring stations with the purpose of quantify seasonal patterns of nutrient and chlorophyll concentrations in different estuarine habitats (marsh creek, cove, surface open water, bottom open water).

Reserve staff setting the ISCO at

Potter Cove.

Grab sampling done by Reserve staff.

The diel sampling program requires to collect a series of samples from one station over an approximately 24-hour period each month to examine how nutrient and chlorophyll concentrations change over diel and tidal cycles.  Previously (from 2002 to 2010) the diel station was located at T-Wharf.  However, after analyzing the historic data from the site, no significant trends or patterns were found over time.  Therefore, the diel station was moved to Potter Cove in January of 2011 in order to characterize nutrients and chlorophyll from this site.  All collected grab and diel samples are analyzed for concentrations of phosphates (PO4), ammonia (NH4), nitrite (NO2), nitrate (NO3), NO2+NO3, dissolved inorganic nitrogen (DIN), silicates (SiO4) and chlorophyll a.

Since 2001, meteorological data has been collected as part of the SWMP at the weather station (see image below) located on Prudence Island, approximately 389 m south of Potter Cove (41o 38’ 13.703” N, 71o 20’ 21.790” W, Trimble Geo XT, GeoExplorer 2008 Series; see map on section 5, Site Location and Character).  Data on air temperature, relative humidity, barometric pressure, wind speed and direction, photosynthetic active radiation, and precipitation are collected.  Meteorological data is continually used to complement the water quality, biological monitoring and scientific research efforts at the Reserve, as well as to assist stewardship, training, and educational activities around the Bay.

Weather station on Prudence Island.

All this information is available to any interested party through the CDMO www.nerrsdata.org, NBNERR http://nbnerr.org/research-and-monitoring/system-wide-monitoring/, or directly contacting the Research Coordinator or the Marine Research Specialist II.

In 2004, the NBNERR became involved in the Bay Window Monitoring Program (BWMP).  The BWMP housed several programs under different state and federal agencies to study Narragansett Bay’s fish and fisheries, sediment pollution, currents, and hydrography.  Even though Bay Window ended in 2010, some programs where able to keep their monitoring with other funding source.  Currently, NBNERR continues to be an essential part of the original network of fixed-sites recording water quality data in the Bay as part of the Rhode Island Department of Environmental Management and the Bay Assessment and Response Team –(RIDEM BART, (http://www.dem.ri.gov/programs/emergencyresponse/bart/index.php) Our two water quality sites, T-Wharf Surface and T-Wharf Bottom, are part of to the Rhode Island Department of Environmental Management Narragansett Bay Fixed-Site Monitoring Network (RIDEM NBFSMN) nested under the RIDEM Bay Assessment Response Team (BART). The RIDEM BART Program was created for a rapid and effective response to environmental incidents in Narragansett Bay using the water quality data collected by the NBFSMN.  NBNFSM is a collaborative monitoring effort of several agencies and organizations (NBNERR, Office of Water Resources, Narragansett Bay Commission, URI Graduate School of Oceanography, Roger Williams University, Narragansett Bay Estuary Program, and URI Coastal Institute) in Narragansett Bay, to monitor water quality with over a dozen fixed-site monitoring network in the Bay, which in turn includes T-Wharf Surface and T-wharf Bottom stations on Prudence Island.  The NBNERR SWMP stations provided the only year-round water quality data in the Bay network, and without these stations, the program would have a large gap in the mid/lower East Passage of the Bay.  Data from all stations can be found at http://www.dem.ri.gov/programs/emergencyresponse/bart/netdata.php.

Dr. Candace Oviatt at URI-GSO is using NAR NERR chlorophyll data to track the timing and magnitude of winter/spring blooms in Narragansett Bay.

Mrs. Heather Stoffel et al. at URI-GSO compiles NAR NERR dissolved oxygen data from T-Wharf Surface, T-Wharf Bottom stations together with data from the Bay Assessment and Response Team (BART, http://www.dem.ri.gov/bart/latest.htm) to study hypoxia in the Bay.  The peer reviewed articles is listed below:

Codiga, D. L., H. E. Stoffel, C. F. Deacutis, S. Kiernan, and C. A. Oviatt. 2009. Narragansett Bay Hypoxic Event Characteristics Based on Fixed-Site Monitoring Network Time Series: Intermittency, Geographic Distribution, Spatial Synchronicity, and Interannual Variability. Estuaries and Coasts 32: 621-641.  Available at http://www.gso.uri.edu/merl/merl_pdfs/Codiga_etal_2009.pdf.

# II. Physical Structure Descriptors

## 9.  Sensor specifications

From January to March of 2017, NAR NERR deployed 6600V2 and EXO2 data loggers.  EXO2 sondes were deployed exclusively at Nag Creek, Potter Cove and T-Wharf Surface during this period.  At T-Wharf Bottom, YSI 6600V2 sondes were deployed during this period.  All the sondes deployed at all sites were equipped with optical DO ROX probes.

YSI 6600 V2 data sonde configurations

Parameter: 	Temperature

Units: 	Celsius - oC

Sensor Type: 	Thermistor

Model #: 	6560

Range: 	-5 to 50 °C

Accuracy: 	± 0.15 °C

Resolution: 	0.01 °C

Parameter: 	Conductivity

Units: 	milli-Siemens per cm (mS/cm)

Sensor Type: 	4-electrode cell with autoranging

Model #: 	6560

Range: 	0 to 100 mS/cm

Accuracy: 	± 0.5% of reading + 0.001 mS/cm

Resolution: 	0.001 mS/cm to 0.1 mS/cm (range dependent)

Parameter: 	Salinity

Units: 	Parts per thousand (ppt)

Sensor Type: 	Calculated from conductivity and temperature

Range: 	0 to 70 ppt

Accuracy: 	± 1.0% of reading or 0.1 ppt, whichever is greater

Resolution: 	0.01 ppt

Sensor Type: 	Dissolved Oxygen

Units:	Percent saturation

Model #: 	6150 ROX optical probe with mechanical cleaning

Range: 	0 to 500% air saturation

Accuracy: 	0-200% air saturation: ± 1% of the reading or 1% air saturation, whichever is

greater; 200 500% air saturation: ± 15% of reading

Resolution: 	0.1% air saturation

Sensor Type: 	Dissolved Oxygen

Units: 	milligrams per Liter (mg/L)

Model #: 	6150 ROX

Range: 	0 to 50 mg/L

Accuracy: 	0 - 20 mg/L: ± 0.1% mg/L or 1% of the reading, whichever is greater; 20 to 50

mg/L: ± 15% of the reading

Resolution: 	0.01 mg/L

Parameter: 	Non-Vented Level – Shallow (Depth)

Units: 	meters (m)

Sensor Type: 	Stainless steel strain gauge

Range: 	0 to 30 ft. (9.1 m)

Accuracy: 	± 0.06 ft. (0.018 m)

Resolution: 	0.001 ft. (0.001 m)

Parameter: 	pH (bulb probe)

Units: 	pH units

Sensor Type: 	Glass combination electrode

Model #: 	6561

Range: 	0 to 14 units

Accuracy: 	± 0.2 units

Resolution: 	0.01 units

Parameter: 	Turbidity

Units: 	Nephelometric turbidity units (NTU)

Sensor Type: 	Optical, 90° scatter, with mechanical cleaning

Model #: 	6136

Range: 	0 to 1000 NTU

Accuracy: 	± 2 % reading or 0.3 NTU (whichever is greater)

Resolution: 	0.1 NTU

Parameter: 	Chlorophyll Fluorescence

Units: 	micrograms/Liter

Sensor Type: 	Optical probe with mechanical cleaning

Model#: 	6025

Range: 	0 to 400 µg/Liter

Accuracy: 	Dependent on methodology

Resolution: 	0.1 µg /L chl a, 0.1% FS

YSI EXO2 data sondes configuration

Parameter:	Temperature

Units:	Celsius (C)

Sensor Type:	Thermistor

Model#:	599870-01

Range:	-5 to 50 C

Accuracy:	-5 to -35: ± 0.01, -35 to -50: ± .005

Resolution:	0.01 C

Parameter:	Conductivity

Units:	milli-Siemens per cm (mS/cm)

Sensor Type:	4-electrode cell with autoranging

Model#:	599870-01

Range:	0 to 200 mS/cm

Accuracy:	0 to 100: ± 0.5% of reading or 0.001 mS/cm; 100 to 200: ± 1% of reading

Resolution:	0.001 mS/cm to 0.1 mS/cm (range dependent)

Parameter:	Salinity

Units:	practical salinity units (psu)/parts per thousand (ppt)

Sensor Type:	Calculated from conductivity and temperature

Range:	0 to 70 psu

Accuracy:	± 1.0% of reading pr 0.1 ppt, whichever is greater

Resolution:	0.01 psu

Parameter:	Dissolved Oxygen

Units:	Percent saturation

Sensor Type:	Optical probe w/ mechanical cleaning

Model#:	599100-01

Range:	0 to 500% air saturation

Accuracy:	0-200% air saturation: ± 1% of the reading or 1% air saturation, whichever is

greater 200-500% air saturation: ± 5% or reading

Resolution:	0.1% air saturation

Parameter:	Dissolved Oxygen (Calculated from % air saturation, temperature, and salinity)

Units:	milligrams/Liter (mg/L)

Sensor Type:	Optical probe w/ mechanical cleaning

Model#:	599100-01

Range:	0 to 50 mg/L

Accuracy:	0-20 mg/L: ± 0.1 mg/l or 1% of the reading, whichever is greater 20 to 50 mg/L:

± 5% of the reading

Resolution:	0.01 mg/L

Parameter:	Non-vented Level - Shallow (Depth)

Units:	feet or meters (ft. or m)

Sensor Type:	Stainless steel strain gauge

Range:	0 to 33 ft. (10 m)

Accuracy:	± 0.013 ft. (0.04 m)

Resolution:	0.001 ft. (0.001 m)

Parameter:	pH

Units:	pH units

Sensor Type:	Glass combination electrode

Model#:	599701(guarded) or 599702(wiped)

Range:	0 to 14 units

Accuracy:	± 0.01 units within ± 10° of calibration temperature, ± 0.02 units for entire

temperature range

Resolution:	0.01 units

Parameter:	Turbidity

Units:	formazin nephelometric units (FNU)

Sensor Type:	Optical, 90 degree scatter

Model#:	599101-01

Range:	0 to 4000 FNU

Accuracy:	0 to 999 FNU: 0.3 FNU or ± 2% of reading (whichever is greater); 1000 to 4000 FNU

± 5% of reading

Resolution:	0 to 999 FNU: 0.01 FNU, 1000 to 4000 FNU: 0.1 FNU

Parameter:	Chlorophyll Fluorescence

Units:	micrograms/Liter

Sensor Type:	Optical probe

Model#:	599102-01

Range:	0 to 400 µg/Liter

Accuracy:	Dependent on methodology

Resolution:	0.1 µg/L chl a, 0.1% FS

### Dissolved Oxygen Qualifier (Rapid Pulse / Clark type sensor):

The reliability of dissolved oxygen (DO) data collected with the rapid pulse / Clark type sensor after 96 hours post-deployment for non-EDS (Extended Deployment System) data sondes may be problematic due to fouling which forms on the DO probe membrane during some deployments (Wenner et al. 2001).  Some Reserves utilize the YSI 6600 EDS data sondes, which increase DO accuracy and longevity by reducing the environmental effects of fouling.  Optical DO probes have further improved data reliability.  The user is therefore advised to consult the metadata for sensor type information and to exercise caution when utilizing rapid pulse / Clark type sensor DO data beyond the initial 96-hour time period.  Potential drift is not always problematic for some uses of the data, i.e. periodicity analysis.  It should also be noted that the amount of fouling is very site specific and that not all data are affected.  If there are concerns about fouling impacts on DO data beyond any information documented in the metadata and/or QAQC flags/codes, please contact the Research Coordinator at the specific NERR site regarding site and seasonal variation in fouling of the DO sensor.

### Depth Qualifier:

The NERR System-Wide Monitoring Program utilizes YSI data sondes that can be equipped with either vented or non-vented depth/level sensors.  Readings for both vented and non-vented sensors are automatically compensated for water density change due to variations in temperature and salinity; but for all non-vented depth measurements, changes in atmospheric pressure between calibrations appear as changes in water depth.  The error is equal to approximately 1.02 cm for every 1 millibar change in atmospheric pressure, and is eliminated for vented sensors because they are vented to the atmosphere throughout the deployment time interval.

Beginning in 2006, NERR SWMP standard calibration protocol calls for all non-vented depth sensors to read 0 meters at a (local) barometric pressure of 1013.25 mb (760 mm/hg).  To achieve this, each site calibrates their depth sensor with a depth offset number, which is calculated using the actual atmospheric pressure at the time of calibration and the equation provided in the SWMP calibration sheet or digital calibration log.  This offset procedure standardizes each depth calibration for the entire NERR System.  If accurate atmospheric pressure data are available, non-vented sensor depth measurements at any NERR can be corrected.

Beginning in 2006, NERR SWMP standard calibration protocol calls for all non-vented depth sensors to read 0 meters at a (local) barometric pressure of 1013.25 mb (760 mm/hg).  To achieve this, each site calibrates their depth sensor with a depth-offset number, which is calculated using the actual atmospheric pressure at the time of calibration and the equation provided in the SWMP calibration sheet or digital calibration log.  This offset procedure standardizes each depth calibration for the entire NERR System.  If accurate atmospheric pressure data are available, non-vented sensor depth measurements at any NERR can be corrected.

In 2010, the CDMO began automatically correcting depth/level data for changes in barometric pressure as measured by the Reserve’s associated meteorological station during data ingestion.  These corrected depth/level data are reported as cDepth and cLevel, and are assigned QAQC flags and codes based on QAQC protocols.  Please see sections 11 and 12 for QAQC flag and code definitions.

NOTE:  older depth data cannot be corrected without verifying that the depth offset was in place and whether a vented or non-vented depth sensor was in use.  No SWMP data prior to 2006 can be corrected using this method.  The following equation is used for corrected depth/level data provided by the CDMO beginning in 2010:

((1013-BP)*0.0102)+Depth/Level = cDepth/cLevel.

### Salinity Units Qualifier:

In 2013, EXO sondes were approved for SWMP use and began to be utilized by Reserves. While the 6600 series sondes report salinity in parts per thousand (ppt) units, the EXO sondes report practical salinity units (psu).  These units are essentially the same and for SWMP purposes are understood to be equivalent, however, psu is considered the more appropriate designation.  Moving forward the NERR System will assign psu salinity units for all data regardless of sonde type.

### Turbidity Qualifier:

In 2013, EXO sondes were approved for SWMP use and began to be utilized by Reserves. While the 6600 series sondes report turbidity in nephelometric turbidity units (NTU), the EXO sondes use formazin nephelometric units (FNU).  These units are essentially the same but indicate a difference in sensor methodology, for SWMP purposes they will be considered equivalent.  Moving forward, the NERR System will use FNU/NTU as the designated units for all turbidity data regardless of sonde type.  If turbidity units and sensor methodology are of concern, please see the Sensor Specifications portion of the metadata.

### Chlorophyll Fluorescence Disclaimer:

YSI chlorophyll sensors (6025 or 599102-01) are designed to serve as a proxy for chlorophyll concentrations in the field for monitoring applications and complement traditional lab extraction methods; therefore, there are accuracy limitations associated with the data that are detailed in the YSI manual including interference from other fluorescent species, differences in calibration method, and effects of cell structure, particle size, organism type, temperature, and light on sensor measurements.

## 10.  Coded variable definitions

| Water Quality Sampling station | Sampling site code | Station code |
| --- | --- | --- |
| Nag Creek | NC | narncwq |
| Potter Cove | PC | narpcwq |
| T-Wharf Surface | TS | nartswq |
| T-Wharf Bottom | TB | nartbwq |

## 11.  QAQC flag definitions

QAQC flags provide documentation of the data and are applied to individual data points by insertion into the parameter’s associated flag column (header preceded by an F_).  During primary automated QAQC (performed by the CDMO), -5, -4, and -2 flags are applied automatically to indicate data that is missing and above or below sensor range.  All remaining data are then flagged 0, passing initial QAQC checks.   During secondary and tertiary QAQC 1, -3, and 5 flags may be used to note data as suspect, rejected due to QAQC, or corrected.

| Flag | Description |
| --- | --- |
| -5 | Outside High Sensor Range |
| -4 | Outside Low Sensor Range |
| -3 | Data Rejected due to QAQC |
| -2 | Missing Data |
| -1 | Optional SWMP Supported Parameter |
| 0 | Data Passed Initial QAQC Checks |
| 1 | Suspect Data |
| 2 | Open - reserved for later flag |
| 3 | Calculated data: non-vented depth/level sensor correction for changes in barometric pressure |
| 4 | Historical Data: Pre-Auto QAQC |
| 5 | Corrected Data |

## 12.  QAQC code definitions

QAQC codes are used in conjunction with QAQC flags to provide further documentation of the data and are also applied by insertion into the associated flag column.  There are three (3) different code categories, general, sensor, and comment.  General errors document general problems with the deployment or YSI datasonde, sensor errors are sensor specific, and comment codes are used to further document conditions or a problem with the data.  Only one general or sensor error and one comment code can be applied to a particular data point, but some comment codes (marked with an * below) can be applied to the entire record in the F_Record column.

General Errors

GIC	No instrument deployed due to ice

GIM	Instrument malfunction

GIT	Instrument recording error; recovered telemetry data

GMC	No instrument deployed due to maintenance/calibration

GNF	Deployment tube clogged / no flow

GOW	Out of water event

GPF	Power failure / low battery

GQR	Data rejected due to QA/QC checks

GSM	See metadata

Corrected Depth/Level Data Codes

GCC	Calculated with data that were corrected during QA/QC

GCM	Calculated value could not be determined due to missing data

GCR	Calculated value could not be determined due to rejected data

GCS	Calculated value suspect due to questionable data

GCU	Calculated value could not be determined due to unavailable data

Sensor Errors

SBO	Blocked optic

SCF	Conductivity sensor failure

SCS	Chlorophyll spike

SDF	Depth port frozen

SDG	Suspect due to sensor diagnostics

SDO	DO suspect

SDP	DO membrane puncture

SIC	Incorrect calibration / contaminated standard

SNV	Negative value

SOW	Sensor out of water

SPC	Post calibration out of range

SQR	Data rejected due to QAQC checks

SSD	Sensor drift

SSM	Sensor malfunction

SSR	Sensor removed / not deployed

STF	Catastrophic temperature sensor failure

STS	Turbidity spike

SWM	Wiper malfunction / loss

Comments

CAB*	Algal bloom

CAF	Acceptable calibration/accuracy error of sensor

CAP	Depth sensor in water, affected by atmospheric pressure

CBF	Biofouling

CCU	Cause unknown

CDA*	DO hypoxia (<3 mg/L)

CDB*	Disturbed bottom

CDF	Data appear to fit conditions

CFK*	Fish kill

CIP*	Surface ice present at sample station

CLT*	Low tide

CMC*	In field maintenance/cleaning

CMD*	Mud in probe guard

CND	New deployment begins

CRE*	Significant rain event

CSM*	See metadata

CTS	Turbidity spike

CVT*	Possible vandalism/tampering

CWD*	Data collected at wrong depth

CWE*	Significant weather event

## 13.  Post deployment information

Post-deployment readings of sondes deployed at each site.  The readings are made using calibration standards the same day the sonde comes back from the field.  Header abbreviations: Retr=retrived, PDR=post-deployment readings, Cal=calibration, Sp Cond=specific conductivity, Chl Fl=chlorophyll fluorescence.

Nag Creek

| Deploy Date | Retr./ PDR Date | EXO Sonde | Sp Cond. | Dissolved Oxygen |  | pH |  |  | Turbidity |  | Depth | Chl Fl |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | mS/cm | Percent saturation |  | units |  |  | NTU |  | Reading (Offset) | µg L-1 |
| mm/dd/yy | mm/dd/yy | Serial No. | 50.0 | ODO 1 | ODO 2 | 7.00 | 10.00 | 4.00 | 0.0 | 124.0 | (m) | 0.00 |
| 12/20/16 | 01/17/17 | 14A101166 | 51.9 | 100.0 | 100.0 | 7.07 | 10.14 | 4.01 | 0.0 | 123.8 | -0.006 (-0.005) | -0.04 |
| 01/17/17 | 02/14/17 | 14A101167 | 50.1 | 99.2 | 99.2 | 7.04 | 10.06 | 3.94 | 0.1 | 124.1 | -0.060 (-0.054) | -0.29 |
| 02/22/17 | 03/08/17 | 14A101166 | 50.1 | 100.4 | 100.4 | 7.00 | 10.03 | 3.90 | 0.1 | 122.4 | -0.009 (-0.011) | 0.03 |
| 1 03/08/17 | 04/11/17 | 14A101167 | 49.9 | 100.9 | 100.9 | 7.09 | 10.09 | 4.00 | 0.2 | 123.8 | 0.064 (0.065) | -0.25 |
| 04/11/17 | 05/04/17 | 14A101166 | 50.4 | 101.0 | 101.0 | 7.08 | 10.20 | 4.01 | 0.0 | 123.0 | 0.116 (0.117) | -0.03 |
| 05/04/17 | 05/31/17 | 14A101167 | 50.0 | 99.5 | 99.5 | 7.01 | 10.05 | 3.94 | -0.3 | 124.8 | 0.034 (0.035) | 0.20 |
| 05/31/17 | 06/22/17 | 14A101166 | 50.2 | 99.6 | 99.6 | 7.09 | 10.08 | 3.99 | 0.0 | 122.8 | 0.063 (0.000) | -0.02 |
| 06/22/17 | 07/13/17 | 14A101167 | 50.2 | 99.6 | 99.7 | 6.95 | 9.94 | 3.90 | 0.4 | 123.1 | 0.012 (0.011) | 0.18 |

Notes

1 This sonde was lost in the marsh due to a winter blizzard passing through the Northeast on 03/14/17.  It was found on 04/10/17 and post deployment readings were done the next day.  The post-deployment readings and techs from YSI seems to agree that the sonde was not adversely impacted by being out of the water for almost four weeks.

Potter Cove

| Deploy Date | Retr./ PDR Date | EXO 2 Sonde | Sp Cond. | Dissolved Oxygen |  | pH |  |  | Turbidity |  | Depth | Chl Fl |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | mS/cm | Percent saturation |  | units |  |  | NTU |  | Reading (Offset) | µg L-1 |
| mm/dd/yy | mm/dd/yy | Serial No. | 50.0 | ODO 1 | ODO 2 | 7.00 | 10.00 | 4.00 | 0.0 | 124.0 | (m) | 0.00 |
| 12/20/16 | 01/17/17 | 14A101165 | 49.9 | 100.4 | 100.4 | 7.02 | 10.09 | 3.97 | -0.01 | 124.0 | 0.000 (0.000) | -0.05 |
| 01/17/17 | 02/15/17 | 14A101164 | 50.0 | 99.0 | 99.0 | 6.99 | 10.02 | 3.91 | -0.01 | 124.5 | -0.184 (-0.186) | 0.04 |
| 02/15/17 | 03/08/17 | 14A101165 | 50.1 | 100.0 | 100.0 | 6.91 | 9.97 | 3.82 | 0.3 | 122.5 | -0.020 (-0.016) | -0.02 |
| 03/08/17 | 04/06/17 | 14A101164 | 49.8 | 99.8 | 99.8 | 7.08 | 10.12 | 3.98 | 0.1 | 123.9 | -0.056 (-0.050) | 0.15 |
| 04/06/17 | 05/04/17 | 14A101165 | 50.0 | 100.8 | 100.8 | 7.16 | 10.19 | 4.10 | 0.2 | 123.6 | 0.116 (0.114) | -0.03 |
| 05/04/17 | 05/31/17 | 14A101164 | 49.9 | 100.6 | 100.6 | 7.02 | 10.04 | 3.97 | 0.1 | 123.6 | 0.019 (0.024) | -0.09 |
| 05/31/17 | 06/20/17 | 14A101165 | 49.0 | 99.8 | 99.8 | 7.00 | 10.03 | 3.93 | 0.4 | 123.5 | -0.009 (0.000) | -0.16 |
| 06/20/17 | 07/13/17 | 14A101164 | 49.7 | 100.0 | 100.0 | 7.03 | 10.03 | 3.98 | 0.9 | 123.4 | -0.007 (0.008) | -0.16 |

T-Wharf Surface

| Deploy Date | Retr./ PDR Date | EXO 2 Sonde | Sp Cond. | Dissolved Oxygen |  | pH |  |  | Turbidity |  | Depth | Chl Fl |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | mS/cm | Percent saturation |  | units |  |  | NTU |  | Reading (Offset) | µg L-1 |
| mm/dd/yy | mm/dd/yy | Serial No. | 50.0 | Reading 1 | & 2 | 7.00 | 10.00 | 4.00 | 0.00 | 124.0 | (m) | 0.00 |
| 12/20/16 | 01/17/17 | 16J102354 | 49.8 | 101.0 | 101.0 | 7.06 | 10.10 | 3.97 | -0.08 | 125.0 | 0.072 (0.073) | 0.01 |
| 01/17/17 | 02/15/17 | 16J102353 | 50.2 | 99.0 | 99.0 | 7.06 | 10.08 | 3.96 | -0.01 | 123.3 | -0.175 (-0.179) | 0.03 |
| 02/15/17 | 03/07/17 | 16J102354 | 50.1 | 100.8 | 100.8 | 7.01 | 10.07 | 3.90 | -0.02 | 120.9 | 0.136 (0.137) | -0.06 |
| 03/07/17 | 04/06/17 | 16J102353 | 49.9 | 100.2 | 100.2 | 7.02 | 10.06 | 3.94 | 0.36 | 125.0 | -0.023 (-0.023) | 0.03 |
| 04/06/17 | 05/03/17 | 16J102354 | 49.3 | 99.2 | 99.2 | 7.08 | 10.10 | 3.98 | 0.06 | 124.0 | -0.046 (-0.046) | 0.08 |
| 05/03/17 | 06/01/17 | 16J102353 | 50.1 | 99.6 | 99.6 | 7.04 | 10.07 | 3.98 | -0.22 | 123.2 | -0.015 (-0.014) | 0.46 |
| 06/01/17 | 06/20/17 | 16J102354 | 49.9 | 99.2 | 99.2 | 7.06 | 10.09 | 3.97 | 0.11 | 122.9 | -0.021 (-0.020) | -0.10 |
| 06/20/17 | 07/12/17 | 16J102353 | 49.9 | 100.1 | 100.0 | 7.06 | 10.08 | 4.03 | -0.17 | 123.9 | 0.009 (0.005) | 0.22 |

T-Wharf Bottom

| Deploy Date | Retr./ PDR Date | 6600V2 | Sp Cond. | Dissolved Oxygen |  | pH |  |  | Turbidity |  | Depth | Chl Fl |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | mS/cm | Percent saturation |  | units |  |  | NTU |  | Reading (Offset) | µg L-1 |
| mm/dd/yy | mm/dd/yy | Serial No. | 50.0 | Reading 1 | & 2 | 7.00 | 10.00 | 4.00 | 0.0 | 126.0 | (m) | 0.00 |
| 12/28/16 | 01/17/17 | 05E1065AE | 50.3 | 101.0 | 101.0 | 7.10 | 10.13 | 4.05 | 0.1 | 125.4 | 0.075 (0.077) | -0.20 |
| 01/17/17 | 02/15/17 | 06C1207AA | 49.8 | 98.7 | 98.7 | 7.07 | 10.10 | 3.99 | -0.70 | 127.9 | -0.184 (-0.193) | 0.30 |
| 02/15/17 | 03/07/17 | 05E1065AE | 50.3 | 101.5 | 101.5 | 7.14 | 10.22 | 4.04 | -0.20 | 124.9 | 0.110 (0.120) | -0.50 |
| 1 03/07/17 | 03/30/17 | 06C1207AA |  |  |  |  |  |  |  |  |  |  |
| 03/30/17 | 05/03/17 | 05E1065AE | 50.3 | 102.0 | 102.0 | 7.13 | 10.18 | 4.06 | 1.0 | 127.0 | -0.053 (-0.049) | 0.30 |
| 05/03/17 | 06/01/17 | 06C1207AA | 50.1 | 100.6 | 100.6 | 7.06 | 10.09 | 4.04 | -0.1 | 127.4 | -0.020 (-0.018) | 0.30 |
| 06/01/17 | 06/20/17 | 05E1065AE | 50.6 | 99.8 | 99.8 | 7.16 | 10.20 | 4.10 | 0.8 | 127.8 | -0.023 (-0.019) | 0.10 |
| 06/20/17 | 07/10/17 | 06C1207AA | 50.8 | 100.2 | 100.2 | 7.05 | 10.05 | 4.01 | -1.1 | 125.1 | 0.037 (0.038) | 0.00 |

Note

1 No post-calibration readings were done after retrieving this sonde on 03/30 because it had a catastrophic temperature failure.  Since the readings of all water quality parameters are temperature compensated for the 6600V2 sondes, if any post-deployment readings were made, they would be inaccurate.

## 14.  Other Remarks / Notes

Data are missing due to equipment or associated specific probes not being deployed, equipment failure, time of maintenance or calibration of equipment, or repair/replacement of a sampling station platform.  Any NANs in the dataset stand for “not a number” and are the result of low power, disconnected wires, or out of range readings.  If additional information on missing data is needed, contact the Research Coordinator at the reserve submitting the data.

Chlorophyll and fluorescence data were collected along with the data presented in this document, but because chlorophyll and fluorescence are not part of the NERR SWMP Water Quality Program, the data are not reported within this dataset.  Chlorophyll fluorescence data are available upon request by contacting the Reserve.

Slight depth anomalies frequently occur at the Nag Creek deployment location.  Depth at this site may have been influenced to some extent by changes in atmospheric pressure because the sonde was situated in very shallow waters.  On occasion, the combination of low atmospheric pressure and shallow water resulted in negative depth values.  Weather station data appear to support this theory.  It is also suspected that due to its shallow location and location in a salt marsh, freshwater runoff from the island may affect specific conductivity/salinity readings several days after a rain event has occurred.  Specific conductivity, salinity, and dissolved oxygen can also vary greatly in Nag Creek over tidal and diel cycles.  Nag creek is a very shallow site and is highly affected by tides and rain events.  On some occasions, dissolved oxygen (DO) is recorded as negative values at the Nag Creek deployment location.  Dissolved oxygen frequently ranges between supersaturation and complete anoxic conditions (DO = 0 % saturation) at this site during the summer months.  Ice is also common in Nag Creek.  There have been occasions where ice has affected all sonde parameters and are marked in the data as “see metadata”.

It is likely that some of the variability in specific conductivity at Nag Creek is due to freshwater entering the marsh from the red maple swamp abutting it to the east (across the road).  The amount of fresh groundwater inflow has not been quantified, but evidence for it comes in the form of 1) the presence of brackish vegetation species along the upland borders of the marsh in areas abutting the swamp, and 2) new data from a salinity mapping project conducted by Dr. Rick McKinney (EPA Atlantic Ecology Division) that quantifies lower salinities along those same borders.  The variability is surely exacerbated due to the shallow nature of the creek in which the sonde is located.  After rain events, salinities are likely to be relatively high at high tide due to the influx of Bay water, and then correspondingly low at low tide when the surface freshwater predominates.

The following are description of different events that happened during deployment and explanations to the CSM (Comment-See Metadata) code for the four water quality station data files from January to March of 2017.

Nag Creek

Throughout the year there are small turbidity spikes marked 0 GSM CTS and 1 STS CSM. Due to the sonde being deployed in a horizontal position, there is, at times, drift algae that gets caught up on the station and affects the turbidity data.

We believe that the on and off formation of ice in the creek affected the tidal exchange for a few days during the following deployments.  We flagged and coded depth as 1 GSM CWD.  The other water quality parameters were flagged and coded as 0 GSM CWD since they seem not to be affected.

- 01/07 00:45 – 01/11 05:25, except turbidity data was rejected because it showed an increase and then decrease which might imply a problem or debris caught by the sonde that resolved on its own.

- 02/10 05:15 – 02/12 18:15

- 03/04 03:00 –.  03/07 16:15

Deployment period 02/14 – 02/22: No data was collected during this period because the sonde was retrieved from the creek because it was frozen.  This period was flagged and coded as -2 GIC CSM.

Deployment period 03/14 11:00 - 04/11 11:30.  The sonde was dragged out of the water onto the marsh by strong winds and tidal current due to the winter blizzard of 03/14/17.  We made several trips to the marsh and creek to try to find the sonde with no success until 04/11/17 at the high marsh among the bushes about 600 ft. from the deploying site in the creek. The data collected during this period were rejected because the sonde was out of the water; the data were flagged and coded as -3 GOW CSM.  The data missing was flagged and coded as -2 GOW CSM.

Potter Cove

Deployment period 06/08 15:00 – 06/22 09:45.  The depth data collected during this period were flagged and coded 1 GSM CWD.  Depth data collected were slightly off because it seems the line securing the sonde got tangled with the floating dock and was lifted during high tides.  This issue showed as truncated curves representing the high tide when plotting the data (see below).  The problem persisted until the end of the deployment but much less pronounced.  Since this is a shallow, well-mixed site, the other water quality parameters seem not to be affected and were flagged and coded as 0 GSM CDW.

T-Wharf Bottom

Deployment periods 03/02 13:15 – 03/07 10:15; 04/03 10:15 - 05/03 09:15.  The turbidity data collected during this period was flagged and coded 1 GSM CBF and 0 GSM CBF.  While retrieving the sonde and more up close in the lab, we found skeleton shrimp all over the sonde and probes; we also found amphipods, to a lesser extent.  These invertebrate might have caused intermittent high turbidity readings, which we flagged and coded as suspect; it seems that by the end of the deployment the turbidity data went back to what is historically typical readings for the site.

Winter Blizzard March 14, 2017

Potter Cove, T-Wharf Surface and T-Wharf Bottom

Deployment period 03/14 11:30 – 18:15.  The depth data collected at Potter Cove, T-Wharf Surface and T-Wharf Bottom was slightly different due to high winds from the East during the winter blizzard this day.  It seems the wind pushed the tide and did not let it go out.  We flagged and coded depth as 1 GSM CWE.  The other water quality parameters were not affected and were flagged and coded 0 GSM CWE.  Wind information was obtained from our weather station on Prudence Island, about 160 m south of Potter Cove.

Pilson, M.E.Q. 1985. On the residence time of water in Narragansett Bay. Estuaries 8:2–14.

Narragansett Bay National Estuarine Research Reserve.  2007.  An Ecological Profile of the Narragansett Bay National Estuarine Research Reserve.  K.B. Raposa and M.L. Schwartz (eds.), Rhode Island Sea Grant, Narragansett, R.I. 176pp.

Wenner, E.L., A.F. Holland, M.D. Arendt, Y. Chen, D. Edwards, C. Miller, M. Meece, and J. Caffrey.  2001.  A Synthesis of Water Quality Data from the National Estuarine Research Reserve’s System-Wide Monitoring Program. NOAA Grant NA97OR0209, MRD Contribution No. 459, 276p.