Rookery Bay (RKB) National Estuarine Research Reserve Meteorological Metadata

January – December 2008

Latest Update:  July 13, 2017

I. Data Set & Research Descriptors

1) Principal investigator(s) & contact persons

Principal Investigator:

Christina Ottman, Research Coordinator

christina.ottman@dep.state.fl.us

Contact Person/Weather Data Manager:

Christina Panko Graff, Water Quality Program Manager

christina.pankograff@dep.state.fl.us

Address:

Rookery Bay NERR

300 Tower Rd.

Naples, FL 34113

Phone: 239-417-6310 ext 403

Fax: 239-417-6315

2) Entry verification

Christina Panko Graff, Water Quality Program Manager, was responsible for sensor calibration, monthly downloads and weather data management.  The CR1000 programs rkbv4.0062107.csi (Jan 1, 2007 - March 18, 2008), rkbv5.0031808.csi (March 18, 2008 - April 1, 2008), and rkbv5.5032608.csi (April 1, 2008 - December 31, 2008) controlled the sensors and data collection schedule. Fifteen minute meteorological data were downloaded from each instrument on the weather station to a CR1000 datalogger.  The CR1000 then interfaced with Loggernet software supplied by Campbell Scientific. This software was located on a laptop computer to which the data were downloaded approximately every month in a comma-delimited format.  The data were saved as a monthly raw data file (.dat) onto a separate hard drive and backed up onto the Rookery Bay NERR server.  The file was uploaded to the CDMO where they underwent automated primary QAQC and became part of the CDMO’s online provisional database.  During primary QAQC, data were flagged if they were missing or out of sensor range.  The edited file was then returned to the Reserve where it was opened in Microsoft Excel and processed using the CDMO’s NERRQAQC Excel macro.  The macro inserted station codes, created metadata worksheets for flagged data, and graphed the data for review.  It allowed the user to apply QAQC flags and codes to the data, append files, and export the resulting data file to the CDMO for tertiary QAQC and assimilation into the CDMO’s authoritative online database.  For more information on QAQC flags and QAQC codes, see Sections 11 and 12.

3) Research objectives (Campbell Weather Station):

As part of the National Estuarine Research Reserve’s (NERR) System-Wide Monitoring Program (SWMP), the principal objective was to record long-term meteorological data at Rookery Bay in order to observe any environmental changes or trends over time. Secondary objectives included using the meteorological data to help document periodic biological and environmental events such as:  fish kills, red tide events, wildfires, prescribed burns, and hurricanes/tropical storms.  It also supported all research conducted at the reserve including providing valuable ancillary data for the reserve's water quality monitoring programs.

4) Research methods:

From January 1, 2008 - December 31, 2008 the Campbell Scientific weather station (operated by the CR1000 datalogger) collected data every 15 minutes in the following format:

Air Temperature (°C), Relative Humidity (%), Barometric Pressure (mb), Wind Speed (m/s), Wind Direction (degrees), and Battery Voltage (volts) based on averages from 5-second data.

Maximum Air Temperature (°C), Minimum Air Temperature (°C), and their times from 5-second data.  (these are not included in the data, but are available from the RKB NERR)

Maximum Wind Speed (m/s) and its time from 5-second data (began collecting as a standard parameter at RKB on 4/1/2008).

Wind Direction Standard Deviation (degrees) from 5-second data (began collecting as a standard parameter at RKB on 12/2/2008).

Precipitation (mm), Cumulative Precipitation (mm) (began collecting as a standard parameter at RKB on 12/2/2008), and Photosynthetic Active Radiation (PAR; millimoles/m2) based on totals of 5-second data.

Following the NERR SWMP Campbell Scientific CR1000 Meteorological Monitoring Station Standard Operating Procedure (SOP), every month the weather station was inspected and a field log was completed to document the condition of the weather station.  The CR1000 enclosure was examined, the wiring panel was checked for loose wires, the desiccant packs were replaced, and the barometric pressure vent was inspected for obstructions.  Additionally, each of the sensors, the solar panel, and the cable connections on the weather station were examined for damage or debris.  If any was found, it was repaired and/or cleaned.  To determine the accuracy of the sensors, real-time data from the weather station was recorded in the field log and compared to data collected either with a Kestrel instrument or from a another local weather station.  Sensors were removed and sent back to Campbell Scientific or Vaisala for recalibration as recommended by the NERR SWMP Campbell Scientific CR1000 Meteorological Monitoring Station SOP.

After inspection of the weather station, the data was downloaded from the CR1000 to a laptop.  After submitting the files to CDMO for automated primary QAQC as described in Section 2, the flagged data and graphs were reviewed within one week of the data download.  Data were further flagged and coded based on the field notes and review of the graphs.  There were no other analyses done on the meteorological data.  If erroneous data were found the sensor(s) underwent troubleshooting to determine if there was a problem that warranted removal and repair.

Campbell Scientific data telemetry equipment was used at this station to transmit to the NOAA GOES satellite, NESDIS ID #3B01D3E8.   The transmissions were scheduled hourly and contained four (4) datasets reflecting the fifteen minute data sampling intervals. Upon receipt by the CDMO, the data underwent the same automated primary QAQC process detailed in Section 2 above.  The “real-time” telemetry data became part of the provisional dataset until undergoing secondary and tertiary QAQC and assimilation in the CDMO’s authoritative online database.  Provisional and authoritative data are available at http://cdmo.baruch.sc.edu.

5) Site location and character:

Rookery Bay is located in the Tropical Rainy climatic group of Koppen (1931), i.e. the area below a west-east line extending from Ft. Myers to Melbourne, and where the mean temperature does not fall below 17.7°C (64°F) in the coolest month of the year.  Owing to the influence of the warm-water Florida Current, the seasonal effects from the Gulf of Mexico Loop Current, and its geographical position at 26° N latitude, the average annual temperature in the Rookery Bay  area is about 24°C (75°F).  Winter temperatures range from -1°C (ca. 30°F) to about 26°C (75°F), with cooler days and nights (10-15°C) in the months of January and February.  Warming trends in April and May are frequently modified by blustery winds from the southwest off the Gulf of Mexico, and by late season cold fronts with northerly breezes.  Summer high temperatures approach 35°C (95°F) or higher on occasion (Thomas 1974).

Rookery Bay and vicinity have an annual rainfall of 50-55 inches (127-140cm) per year (Thomas 1974).  The heaviest average monthly rainfall, 8-9 inches per month, occurs from June through September.  Lowest average rainfall, 1-2 inches per month, occurs from November through March.  Approximately 66% of the total yearly rainfall occurs between the months of June and October.  Southwest Florida lies in the seasonal tropical weather belt that channels hurricanes toward or along the coast.  Historically, the area is fortunate in that few severe hurricanes have come ashore.

The weather station is located at Rookery Bay's Environmental Learning Center (ELC) located at 300 Tower Road, Naples, FL, which is adjacent to Henderson Creek (Latitude: 26° 02.983' N/ Longitude: 81° 42.059' W).  The weather station is located approximately 1-2 miles upstream from the water quality site located in lower Henderson Creek.  The CR1000 enclosure housing the GOES transmitter, battery, and the barometric pressure sensor are on the bottom (approx. 1.3 m above the ground) of a 20 m aluminum tower adjacent to the ELC.  The base of the tower is approximately 3 m above sea level.  The temperature/relative humidity sensors are on the north side of the tower opposite the LiCor sensor (south side) about 10 m up on the tower.  The tipping bucket rain gauge is 5 m up on the tower on the south side.  There are no shading or obstruction effects on the temp/relative humidity sensor, LiCor sensor, or the tipping bucket.  From 1/1/2008 until 2/25/2008 the ultrasonic wind sensor was recording data at a lower position on the tower (12 m).  It is possible that there was some influence of the building on the ultrasonic wind sensor at the lower location that may be reflected in the data beginning on 1/1/2008 00:00.   On 2/25/2008 the WS425 ultrasonic wind sensor was moved to the north side of the tower at about 18 m up on the tower (about 21.4 m above sea level) and positioned for true north rather than magnetic north.  Sensors were wired according to NERR SWMP Campbell Scientific CR1000 Meteorological Monitoring Station SOP.

6) Data collection period:

Weather data have been collected at the Henderson Creek/Shell Island Road site since 1996.  The current SWMP weather station began collecting data on January 7, 2004 at 13:15. The data collection for 2008 began on January 1, 2008 at 00:00 and continued through 23:45 of December 31, 2008.

7) Distribution:

According to the Ocean and Coastal Resource Management Data Dissemination Policy for the NERRS System-wide Monitoring Program,NOAA/ERD retains the right to analyze, synthesize and publish summaries of the NERRS System-wide Monitoring Program data.  The PI retains the right to be fully credited for having collected and processed the data.  Following academic courtesy standards, the PI and NERR site where the data were collected will be contacted and fully acknowledged in any subsequent publications in which any part of the data are used.  Manuscripts resulting from the NOAA/OCRM supported research that are produced for publication in open literature, including refereed scientific journals, will acknowledge that the research was conducted under an award from the Estuarine Reserves Division, Office of Ocean and Coastal Resource Management, National Ocean Service, National Oceanic and Atmospheric Administration.  The data set enclosed within this package/transmission is only as good as the quality assurance/quality control procedures outlined by the enclosed metadata reporting statement.  The user bears all responsibility for its subsequent use/misuse in any further analyses or comparisons.  The Federal government does not assume liability to the Recipient or third persons, nor will the Federal government reimburse or indemnify the Recipient for its liability due to any losses resulting in any way from the use of this data.

NERR weather data and metadata can be obtained from the Research Coordinator at the individual NERR site (please see Section 1 Principal investigators and contact persons), from the Data Manager at the Centralized Data Management Office (please see personnel directory under the general information link on the CDMO home page) and online at the CDMO home page http://cdmo.baruch.sc.edu/.  Data are available in comma separated format.

8) Associated researchers and projects:

Rookery Bay participates in the NERR SWMP for water quality and has four monitoring locations within the Reserve boundary.  The principal objective of this study is to record long-term water quality data within Rookery Bay in order to observe any physical changes or trends in water quality over time.  The four sites were selected to represent various degrees of watershed hydrologic alteration.  The Henderson Creek water quality monitoring station is located approximately 1-2 miles downstream from this weather station on upper Henderson Creek.  Measurements are taken every 15 minutes including temperature, salinity, conductivity, pH, turbidity, dissolved oxygen and depth.

Additionally Rookery Bay participates in the SWMP for long-term nutrient monitoring.  This project was in collaboration with Dr. Joseph N. Boyer at the Southeast Environmental Research Center, Florida International University, Miami, FL.  This long-term monitoring study was initiated in 1991 to characterize status and trends in water quality of Florida Bay and has since expanded to cover much of the South Florida region, including several sites within the Rookery Bay NERR.  The parameters included in this project are soluble reactive phosphorus (filtered), total phosphorus, ammonium (filtered), nitrite (filtered), nitrate (filtered), nitrite + nitrate (filtered), dissolved inorganic nitrogen (calculated), total nitrogen, total organic nitrogen (calculated), chlorophyll a, total organic carbon, and silicates.

Both water quality and nutrient data generated by Rookery Bay are being used to analyze restoration targets established for the Picayune Strand Restoration Project (PSRP; formerly known as Southern Golden Gate Estates) which is a portion of the Comprehensive Everglades Restoration Plan (CERP). Additional datasets used in this analysis include a long-term fisheries survey (July 1998 to the present), a shark demographics survey (May 2000 to the present), and an oyster reef/benthic crab survey (1999 to 2008). The fisheries data are obtained through monthly trawls in the bays corresponding with the SWMP water quality sites to document the population dynamics in a variety of fish species, as well as commercially important  invertebrates such as stone crabs, blue crabs and pink shrimp. Shark demographic data are also collected monthly from the reference bays downstream of the PSRP through long-line and gillnet ‘tag, measure and release’ surveys. Benthic crabs are collected from oyster reefs using Hester-Dendy collection substrates at the four SWMP water quality stations.

II. Physical Structure Descriptors

9) Sensor specifications:

Parameter: Temperature

Units: Celsius

Sensor type: Platinum resistance temperature detector (PRT)

Model #: HMP35C Temperature and Relative Humidity Probe

Operating Temperature:  -40 to 60°C

Range: -40 to 60°C

Accuracy: ± 0.2 °C @ 20°C

Date of last calibration:  2/2008

Date of deployment: 3/18/2008

Parameter: Relative Humidity

Units: Percent

Sensor type: Vaisala HUMICAP© 180 capacitive relative humidity sensor

Model #: HMP35C Temperature and Relative Humidity Probe

Range:  0-100% non-condensing

Accuracy @ 20°C:  +/-2% RH (0-90%) and +/-3% (90-100%)

Temperature dependence of RH measurement: +/- 0.05% RH/°C

Uncertainty of calibration: ± 1.2% RH

Date of last calibration:  2/2008

Date of deployment: 3/18/2008

Parameter: Barometric Pressure

Units: millibars (mb)

Sensor type: Vaisala Barocap © silicon capacitive pressure sensor

Model #: CS-105

Operating Range: Pressure: 600 to 1060 mb; Temperature:  -40°C to +60°C;

Humidity: non-condensing

Accuracy: ± 0.5 mb @ 20°C; +/- 2 mb @ 0°C to 40°C; +/- 4 mb @ -20°C to 45°C; +/- 6 mb @ -40°C to 60°C

Stability: ± 0.1 mb per year

Date of last calibration:  2/2008

Date of deployment: 3/17/2008

Parameter: Wind Speed

Units: meters per second (m/s)

Sensor type: Vaisala ultrasonic wind sentry

Model #: WS425

Range: 0-65 m/s

Accuracy: ± 0.135 m/s or 3% of the reading

Operating Temperature:-40 to 55 °C

Date of last calibration: 4/2006

Date began using data: 7/5/2007

Parameter: Wind Direction

Units: degrees

Sensor type: Vaisala ultrasonic wind sentry

Model #: WS425

Range: 0-360°

Accuracy: +/- 2°

Operating Temperature: -40 to 55 °C

Date of last calibration: 4/2006

Date began using data: 7/5/2007

Parameter: Photosynthetic Active Radiation (PAR)

Units: mmoles/m2 (total flux)

Sensor type: LiCor Quantum Sensor- High stability silicon photovoltaic detector (blue enhanced)

Model #: LI190SB

Light spectrum wavelength:  400 to 700 nm

Temperature dependence: 0.15% per °C maximum

Stability:  < ±2% change over 1 yr

Operating Temperature:  -40 to 65°C; Humidity: 0 to 100%

Sensitivity:  typically 5 µA per 1000µmoles s-1 m-2

Q22437

Date of last calibration:  1/2007

Dates deployed: 1/1/2008-3/24/2008; 4/1/2008-4/14/2008

Multiplier:  1.28

Q34382

Date of last calibration: 1/2004; recalibrated 4/25/2008

Dates deployed: 3/24/2008-4/1/2008; 7/1/2008

Multiplier: 1.28 (1/2004); 1.24 (4/25/2008)

CDMO loaner Q34301

Date of last calibration: 12/2004

Dates deployed: 4/14/2008-7/1/2008

Multiplier: 1.00

Parameter: Precipitation

Units: millimeters (mm)

Sensor Type: Tipping Bucket Rain Gauge

Model # TR-525

Rainfall per tip: 0.01 inch

Operating range: Temperature: 0° to 50°C; Humidity: 0 to 100%

Range: 0.1 mm

Accuracy: +/- 1.0% up to 1 in./hr; +0, -3% from 1 to 2 in./hr; +0, -5% from 2 to 3 in./hr

Date of last calibration: 6/2008

Date of deployment: 6/29/2008

Storage Module

CR1000

Operating range:  Temperature= -25o to 50oC

Baud rates:  300 bsp to 115.2 kbps

Power requirements:  9.6-16 VDC

Date of installation: 5/13/2006

The CR1000 has two MB Flash EEPROM that are used to store the Operating System. Another 128 K Flash is used to store configuration settings. A minimum of 2 MB SRAM (4 MB optional) is available for program storage (16K), operating system use, and data storage.  Additional storage is available by using a compact flash card in the optional CFM100 Compact Flash Module.

10) Coded variable indicator and variable code definitions:

Sampling station:			Sampling site code:		Station code:

Upper Henderson Creek			UH			rkbuhmet

11)  QAQC flag definitions

QAQC flags provide documentation of the data and are applied to individual data points by insertion into the parameter’s associated flag column (header preceded by an F_).   During primary automated QAQC (performed by the CDMO), -5, -4, and -2 flags are applied automatically to indicate data that is above or below sensor range, or missing.  All remaining data are then flagged 0, as passing initial QAQC checks.   During secondary and tertiary QAQC 1, -3, and 5 flags may be used to note data as suspect, rejected due to QAQC, or corrected.

-5	Outside High Sensor Range

-4	Outside Low Sensor Range

-3		Data Rejected due to QAQC

-2		Missing Data

-1		Optional SWMP supported parameter

0		Passed Initial QAQC Checks

1		Suspect Data

2		Open - reserved for later flag

3		Open - reserved for later flag

4		Historical Data:  Pre-Auto QAQC

5		Corrected Data

12)  QAQC code definitions

QAQC codes are used in conjunction with QAQC flags to provide further documentation of the data and are also applied by insertion into the associated flag column.  There are three (3) different code categories, general, sensor, and comment.  General errors document general problems with the station, sensor errors are sensor specific, and comment codes are used to further document conditions or a problem with the data.  Only one general or sensor error and one comment code can be applied to a particular data point.

General Errors

GIM	Instrument Malfunction

GIT	Instrument Recording Error, Recovered Telemetry Data

GMC	No Instrument Deployed due to Maintenance/Calibration

GMT 	Instrument Maintenance

GPD	Power Down

GPF	Power Failure / Low Battery

GPR	Program Reload

GQR	Data Rejected Due to QA/QC Checks

GSM	See Metadata

Sensor Errors

SIC		Incorrect Calibration Constant, Multiplier or Offset

SNV	Negative Value

SOC	Out of Calibration

SSN	Not a Number / Unknown Value

SSM	Sensor Malfunction

SSR	Sensor Removed

Comments

CAF	Acceptable Calibration/Accuracy Error of Sensor

CDF	Data Appear to Fit Conditions

CRE	Significant Rain Event

CSM	See Metadata

CVT	Possible Vandalism/Tampering

13)  Other remarks/notes

Data are missing due to equipment or associated specific sensors not being deployed, equipment failure, time of maintenance or calibration of equipment, or repair/replacement of a sampling station platform.  Any NANs in the dataset stand for “not a number” and are the result of low power, disconnected wires, or out of range readings.  If additional information on missing data is needed, contact the Research Coordinator at the reserve submitting the data.

Small negative PAR values are within range of the sensor and are due to normal errors in the sensor and the CR1000 Datalogger. The maximum signal noise error for the Licor sensor is +/- 2.214 mmoles/m2 over a 15 minute interval.

Relative humidity data greater than 100 are within range of the sensor accuracy of +/-3%.

Cumulative precipitation data are recorded from 00:00 to 23:59 with the daily total recorded at the midnight mark (00:00). The midnight CumPrcp value is actually the total from the previous day.

As of April 1, 2008, all Reserves were required to align their wind direction sensor to True North.  RKB wind sensor was aligned on 2/22/2008 (see below for specific times).  Prior to this, the RKB wind sensor was aligned to magnetic north.

During 2017 the CDMO discovered an incorrect line in the CR1000 programming.  If RHumidity>100 And RHumidity<108 Then RHumidity=100.  A decision was made by the DMC during 2006 to discontinue correcting >100 RH values to 100. This change was never made in our program and has remained in each updated version until it was removed during 2017.  By correcting all values >100 during data collection we may have missed erroneous values that could have indicated a problem with the RH sensor.  CSM coding was added to all RH data from 2007 until the programming change in 2017.

Storm systems:  On 2/12/2008-2/13/2008 a strong front brought a severe storm resulting in significant rainfall and wind gusts; tropical storm Fay, August 19, 2008; hurricane Ike, September 9-10, 2008 (Passed through the Gulf of Mexico)

Program Failure:

From 13:15 on 3/17/2008 until 9:45 on 3/18/2008:  All sensors’ data were missing due to a program failure which did not allow retrieval of the data.

Temperature/Relative Humidity Sensor:

On 2/22/2008 from 11:30 until 12:15:  The temperature and relative humidity data were rejected because the 12 V wire for the Temp/RH probe was accidentally removed while removing wires for the wind sensor.

The temperature/relative humidity sensor was swapped with a freshly calibrated sensor on 3/18/2008 at 8:11.

On 4/14/2008 from 13:30-14:30:  Temperature and relative humidity data were rejected because a few of the cables had to be switched and rewired to fit a Li-Cor sensor loaned from CDMO.

Barometric Pressure Sensor:

The barometric pressure sensor was swapped with a freshly calibrated sensor on 3/17/2008.

On 3/18/2008 at 12:15: The barometric pressure data were rejected because it was abnormally low in the data trend.

From 16:15 on 8/18/09 until 13:45 on 8/19/09:  Tropical storm Fay passed through the Reserve resulting in a major pressure drop.  These data are considered suspect because of the possibility of a wire popping out due to the significant pressure drop when TS Fay passed through the area.

From 14:00 on 8/19/09 until 12:15 on 12/8/2008:  The barometric pressure data were rejected due to a loose wire that may have been causing the intermittently lower than usual readings. The wire may have popped out during the low pressure event with tropical storm Fay.

Wind Sensor:

On 2/22/2008 from 12:45 until 13:45 on 3/12/2008: The wind speed and wind direction data were rejected because of maintenance (assembled a new mount for the ultrasonic wind sensor, positioned it higher on the tower, and positioned it for true north) and associated problems that occurred.  The sensor was not wired in from 12:45 on 2/22/2008 until 14:30 on 2/25/2008 (weekend) while the mount was assembled and other adjustments were made.  On 2/25/2008 after mounting the sensor at the higher location and rewiring it (rewired at 14:45), the sensor was not recording properly.  Attempts over the course of several days to troubleshoot the problem were without success.  On 3/10/2008 removed the sensor at 10:00.  Broke the cable splice and discovered that the splice in the cable was damaged.  Tested the original cable at 14:10.  On 3/12/2008 the cable was re-spliced and then tested at 12:35. The sensor was mounted and rewired at 13:35 on 3/12/2008.

On 03/24/2008 from 15:45-16:00:  The wind speed and wind direction data were considered suspect because

the 12V wire was accidently pulled out while rewiring the PAR sensor.

Maximum wind speed and wind speed data were rejected on the following dates and times, potentially due to intermittent obstructions of the wind sensor or a program error resulting in a maximum wind speed reading of 447.0:  3/29/2008 at 13:30, 4/28/2008 at 11:15, 5/28/2008 at 6:45, 5/31/2008 at 6:45, 6/12/2008 at 16:00 and 16:15, 6/24/2008 at 16:15 and 16:30, 7/14/2008 at 6:00 and 9:00, 7/16/2008 at 11:45, 7/18/2008 at 14:45, 7/30/2008 at 6:15, 8/17/2008 at 17:15, 8/19/2008 at 3:45, 9/6/2008 at 4:15, 9/16/2008 at 6:30, 9/22/2008 at 14:00 and 14:15, 10/13/2008 at 14:00, 10/19/2008 at 7:00, 10/22/2008 at 8:45, 12:30, and 13:30, 10/30/2008 at 9:15, 10/31/2008 at 7:15, 11/12/2008 at 16:15, 11/13/2008 at 8:30 and 9:00, 11/14/2008 at 16:15, 11/15/2008 at 6:45 and 7:00, 11/16/2008 at 17:30, 11/25/2008 at 8:00, 11/30/2008 at 7:45, 8:15, 9:15, 12:45, 13:00, 13:45, 14:15, and 16:15, 12/1/2008 at 9:15, 12/2/2008 at 7:30, 12/7/2008 at 6:45 and 7:00.

On 11/19/2008 at 8:30 and 12/7/2008 at 10:15: the maximum wind speed and wind speed data were rejected due to erroneous maximum wind speed readings.

PAR (Li-Cor sensor):

From 00:15 on 1/1/2008 until 14:15 on 3/24/2008 (sensor Q22437):  All of the PAR data should be considered suspect due to inconsistent sensor drift at night (sunset to sunrise based on a national weather service local station) which exceeded the maximum signal noise error (+/- 2.214 mmoles/m2 over a 15 minute interval).  Since the error was inconsistent, all of the data may have been inaccurate and considered suspect.  All night-time PAR data that exceeded the maximum signal noise error at night were rejected.  The PAR data were also corrected for a miscalculated PAR multiplier value.  The data were corrected by taking the incorrect PAR data divided by the incorrect PAR multiplier multiplied by the correct PAR multiplier.

On 3/24/2008 from 14:30-15:45(sensor Q34382): PAR data was rejected because the Li-Cor sensor Q 22437 was switched with Q34382 to compare data temporarily. Sensor Q34382 was in use until 4/1/2008 9:30.

From 11:00 on 4/1/2008 until 13:00 on 4/14/2008 (sensor Q22437):  All of the PAR data should be considered suspect due to inconsistent sensor drift at night which exceeded the maximum signal noise error (+/- 2.214 mmoles/m2 over a 15 minute interval).  Since the error was inconsistent, all of the data may have been inaccurate and considered suspect.  All night-time PAR data that exceeded the maximum signal noise error at night were rejected.

On 4/14/2008 from 13:15-14:30: PAR data were rejected because sensor Q22437 was removed and replaced with a Li-Cor sensor loaned from CDMO.  Q22437 was recording values outside the maximum signal noise error at night even though still within calibration period.  Therefore, sensor Q22437 was sent back to Li-Cor to be evaluated for potential problems.  The loaner PAR sensor was installed until 7/1/2008 9:00 when it was replaced by sensor Q34382.

From 14:45 on 4/14/2008 until 9:45 on 5/12/2008 (sensor loaned from CDMO): PAR data were corrected for a miscalculated PAR multiplier value.  The data were corrected by taking the incorrect PAR data divided by the incorrect PAR multiplier multiplied by the correct PAR multiplier.  On 5/12/2008 at 9:40 the PAR multiplier was reset. The 9:45 PAR data were rejected because of the multiplier change at 9:40.

Concluding comments: Although Rookery Bay’s Li-Cor sensors were deployed within the average calibration time period (up to 2 years) other than during trouble-shooting, the sensors repeatedly recorded values outside the maximum signal noise error at night.  One possible explanation may be the location of the sensor, which is adjacent to the Environmental Learning Center building and potential ambient light sources.  All PAR data should be considered suspect.  All night-time PAR data that exceeded the maximum signal noise error at night were rejected.

Precipitation Sensor:

From 11:30 on 6/30/2008 until 16:00 7/17/2008:  TotPrcp and CumPrcp data were considered suspect because

of a wiring mistake.