# Rookery Bay (RKB) NERR Water Quality Metadata

January - December, 2007

Latest Update: May 10, 2012

I.  Dataset and Research Descriptors

Principal Investigators and Contact Persons

a) Reserve Contacts

Christina Gwaltney, Research Coordinator

Rookery Bay National Estuarine Research Reserve

300 Tower Rd.

Naples, FL 34113

Tel:  (239) 417-6310 ext. 402

Fax: (239) 417-6315

Christina.gwaltney@dep.state.fl.us

b) Laboratory Contact

Vicki K. McGee, Water Quality Program Manager

Rookery Bay National Estuarine Research Reserve

300 Tower Road

Naples, FL 34113

Tel: (239) 417-6310 ext 403

Fax: (239) 417-6315

vicki.mcgee@dep.state.fl.us

2) Entry verification

Deployment data are uploaded from the YSI data logger to a Personal Computer (IBM compatible).  Files are exported from EcoWatch in a comma-delimited format (.CDF) and uploaded to the CDMO where they undergo automated primary QAQC and become part of the CDMO’s online provisional database.  Excessive pre- and post-deployment data are removed from the file prior to upload with up to 2 hours of pre- and post-deployment data retained to assist in data management.  During primary QAQC, data are flagged if they are missing or out of sensor range.  The edited file is then returned to the Reserve where it is opened in Microsoft Excel and processed using the CDMO’s NERRQAQC Excel macro.  The macro inserts station codes, creates metadata worksheets for flagged data, and graphs the data for review.  It allows the user to apply QAQC flags and codes to the data, remove remaining pre- and post-deployment data, append files, and export the resulting data file to the CDMO for tertiary QAQC and assimilation into the CDMO’s authoritative online database.  Where deployment overlap occurs between files, the data produced by the newly calibrated sonde is generally accepted as being the most accurate.  For more information on QAQC flags and codes, see Sections 11 and 12.

Data management and analysis and reporting are performed by Vicki McGee, Water Quality Program Manager.

3) Research objectives

Beginning January 1, 2001, two stations, lower Henderson Creek and Middle Blackwater River were designated as the Reserve’s permanent SWMP (System-wide monitoring program) sites.  With the exception of turbidity, the dataset at these stations is nearly continuous beginning in April of 1996.  The dataset from these monitoring stations represents the longest record of continuous water quality measurements of all the Rookery Bay NERR water quality stations.  In January 2002 two additional sites, Faka Union Bay and Fakahatchee Bay, were designated as SWMP water quality stations.  All stations currently log temperature, turbidity, salinity, pH, dissolved oxygen and depth at 15-minute intervals

The four stations are in estuaries with watersheds having different patterns of landuse.  Their placement addresses priority resource management issues that are identified in the Reserve’s management plan.  Specifically, the data from these stations are providing valuable information concerning the affects of landuse activities on the quantity, quality and timing of freshwater inflow into the Reserve.

The lower Henderson Creek station receives most of its freshwater from a canal system that drains a watershed area with approximately 50% developed versus natural landscape.  Landuse in the developed area is divided equally between residential and agricultural activities.  A weir controls most of the freshwater flowing into Henderson Creek.  The Reserve obtained grant funding to upgrade the existing weir structure in May 2002 to allow for an automated control of water release.  The dataset being obtained from the lower Henderson Creek station, along with biological monitoring, and hydrological modeling is being used to develop a water release strategy for Henderson Creek.

Relative to Henderson Creek, the middle Blackwater River station has a watershed that has 75% natural landscape.  Agriculture dominates the land-use activities within the developed area of Blackwater River watershed.  Based on modeling of historic flow-way patterns, the watershed draining into Blackwater River has been significantly reduced resulting in higher than normal salinity values.

Faka Union Bay represents the most extreme condition of altered freshwater inflows of the four Rookery Bay NERR monitoring sites.  Forty –eight miles of canals drain 94 square miles of watershed intercepting freshwater which historically distributed to four Bays.  Restoration work is currently underway to restore historical sheet flow to the watershed.

Fakahatchee Bay is a relatively undisturbed system. The Fakahatchee Bay watershed is primarily under Preserve status and has been relatively undisturbed from the hydrological alteration of the SGGE (Southern Golden Gate Estates) canal system.  The water quality data collected from this station serves as a reference for assessing the effectiveness of watershed habitat restoration projects in southwest Florida.

4) Research methods

In January 2003 all data-loggers were upgraded from the YSI 6600 model to the Extended Deployment System (EDS); the Henderson Creek site was upgraded to the V2 model for 2007.  Prior to deployment, the sondes are calibrated for pH, specific conductivity, turbidity and dissolved oxygen following the procedures outlined in the YSI Operating and Service Manual, with two exceptions.  Both depth and dissolved oxygen were calibrated using the 760 mm Hg default barometric pressure value rather than the local barometric pressure.  As a result, error relative to the difference between the local barometric pressure and 760mm Hg was introduced to the dissolved oxygen calibration and depth offset. The depth data cannot be corrected for changes in atmospheric pressure.  A two-point calibration is used for pH (Fisher buffers 7& 10) and turbidity (0 & 123 NTU YSI, Inc.).  A 0.5M KCL solution (YSI conductivity calibrator) is used to calibrate specific conductivity.  Dissolved oxygen is calibrated in water, using a bucket and an air-stone.  In 2008 ROX optical DO probes were deployed at all four permanent SWMP sites.  Sondes are deployed for up to 2 weeks at a time. After approximately 2 weeks time the fouling on the probes becomes critical.  The sampling period is set for 15-minute intervals (readings are made every 15 minutes on the half-hour). The following physical water quality parameters are recorded: temperature, specific conductivity, salinity, dissolved oxygen (mg/l), dissolved oxygen (% saturation), depth, pH and turbidity.  A post-deployment calibration is done after the sonde is taken out of the water using the post-calibration protocol.  Field measurements are performed upon retrieval of the loggers as a “check” on the accuracy of the loggers. This is done using a handheld YSI 85 meter.

NOTE:  All data within this metadata file and the original data files are reported in Eastern Standard Time (EST).

All sondes are housed within 4-inch diameter PVC pipes, which are attached to “aid-to-navigation” pilings (Lower Henderson Creek, Middle Blackwater River, and Faka Union Bay), or attached to a 6 inch PVC pipe hydro-jetted into the substrate (Fakahatchee Bay).  The pipes are oriented vertically and are attached tightly with stainless steel hose clamps.  The bottom portion of the pipe is open and sits 6 inches above the bottom.  Probe readings are measured approximately 6-8 inches above the substrate.  Twenty 2-inch diameter holes are drilled circumferentially around the lower half of the pipe to ensure adequate water flow around the probes. The bottom two meters of the PVC housing pipe are painted with anti-fouling paint to protect against fouling, and the PVC pipes are changed out twice a year.  Sondes are hung by plastic coated steel cable inside the pipes; the cable is bolted to the piling.  A ‘kryptonite’ bicycle lock is used to secure the sonde inside the housing.

The sondes are further protected from crabs and other live organisms by restricting the openings within the sonde guard with plastic screening. The plastic mesh (with 1/8 in diamond-shaped holes) is attached on the outside of the sonde guards’ circumference using low-profile zip ties. Furthermore, the PVC is serviced on a minimum of a monthly basis, and replaced bi-annually.

A Sutron Sat-Link2 transmitter was installed at the Lower Henderson Creek site on 05/09/06 at 12:26 pm (EST), and transmits data to the NOAA GOES Satellite, NESDIS  ID#3B02E17C.  The transmissions are scheduled hourly and contain four (4) data sets reflecting fifteen minute data sampling intervals.  Upon receipt by the CDMO, the data undergoes the same automated primary QAQC process detailed in Section 2 above. The “real-time” telemetry data become part of the provisional dataset until undergoing secondary and tertiary QAQC and assimilation in the CDMO’s authoritative online database. Provisional and authoritative data are available at http://cdmo.baruch.sc.edu.

5) Site location & character

Lower Henderson Creek (rkblhwq):

The “Lower Henderson” site, is located at the mouth of Henderson Creek.  The piling (manatee caution sign) is located right of center (while facing downstream) of the creek channel, approximately 100 meters from RKBNERR’s boat dock. The Lower Henderson Creek labeling is to clarify the site from other historical water quality stations.

Tides at Lower Henderson Creek are mixed and range from 0 m to 2.76 m (average 1.06 m).  The monitoring site is located in lower Henderson Creek, approximately 5 km downstream of a four-lane highway (SR 951) that crosses the creek.  The creek is 5.8 km long (mainstream linear dimension), has an average mid-channel depth of approximately 2 meters at MHW, and an average width of 239 meters.  At the sampling site, the depth is 2 meters at MHW and the width is 600 meters.  Salinities at this site range from 0 to 39.7 ppt. Creek bottom habitats are predominantly fine sand and there is no bottom vegetation.  The dominant marsh vegetation near the sampling site is red mangrove.  The dominant natural vegetation of the watershed is hydric pine and cypress.  Upland land use near the sampling site includes residential areas with septic systems.

Watershed activities that potentially impact the site include non-point source pollution from road runoff, drift of mosquito control pesticides, runoff from upstream agricultural areas and leachate from nearby residential septic systems and a weir structure located at SR 41. The amount of water released from this weir can sometimes mask natural tidal salinity patterns.  The historic Henderson Creek watershed is approximately 50% under State ownership and much of this protected area has intact cypress sloughs and other wetland vegetation.  Canals and water use for agriculture and human consumption have altered the hydroperiod of this watershed. Consequently, the Henderson creek watershed receives non-point source pollutant runoff from a variety of sources.

UTM GPS coordinates (using UTM zone 17, datum NAD-83):

E	  426633

N	2878731

Lat/Long (Decimal Degrees):

26.0257   N

81.7332   W

Middle Blackwater River (rkbmbwq):

The “ Middle Blackwater River” sonde, is located at the mouth of the river.  The sonde is affixed to navigational marker #17 within the river channel.  The average depth at this marker is approximately 2 meters at MHW.  The Middle Blackwater River labeling is to clarify the site from other historical water quality stations.

The substrate within the channel is a mixture of sand and silt with oyster shell with some organic matter mixed in.  Mature red mangrove forests dominate the banks of the river.  Upstream influences consist of the Collier-Seminole State Park boat basin and upstream agricultural fields adjacent to Blackwater River’s main feeder canal (SR 41 canal).  In addition, canals and roads built during the 1950’s have caused significant disruptions to overland sheet-flow reducing the amounts of freshwater flowing to this estuary.  Despite these alterations, the salinity fluctuations of this site suggest that seasonal fluctuations in salinity are more closely correlated to watershed rainfall patterns than salinities of estuaries with water control structures, such as Henderson Creek,    Blackwater salinities range from 0 to 42.1 ppt and fluctuate with the tides and watershed rainfall.  The tide range for Middle Blackwater River varies between 0.2 and 1.8 meters.

UTM GPS coordinates (using UTM zone 17, datum NAD-83):

E 	440363

N	2868541

Lat/Long (Decimal Degrees):

25.9343   N

81.5946   W

To reference these sites on a nautical chart, see NOAA chart #11430 – “Lostman’s River to Wiggins Pass”.

Faka Union Bay (rkbfuwq):

The “ Faka Union Bay” sonde, site #3, is located at the mouth of the Faka Union Canal.  The sonde is affixed to a manatee speed zone sign within the main channel. The average depth at this site is approximately 2 meters at MHW.

The substrate within the channel is a mixture of sand and silt with some organic matter.  Mature red mangrove forests and spoil islands dominate the banks of the canal.  Upstream influences consist of the Port of the Islands development and marina. The upstream water flow consists of an elaborate canal system which has altered natural water drainage patterns into Faka Union Bay.  Salinities range from 0 to 42.3 ppt and fluctuate daily based on water management of the water control structures.  The tide range for Faka Union Bay varies between 0.2 and 1.6meters.

UTM GPS coordinates (using UTM zone 17, datum NAD-83):

E  448330

N  2864765

Lat/Long:

25.9005  N

81.5159  W

Fakahatchee Bay (rkbfbwq):

The Fakahatchee Bay sonde is located at the mouth of two rivers, Fakahatchee and East River.  The sonde is placed in a 4” PVC  housing secured to a 6” PVC pipe approximately 6 inches off the bottom.  The sonde is located within the bay adjacent to one of the Reserve’s biomonitoring stations. The average depth at MHW is approximately 2 meters.

The substrate within the channel is a mixture of sand, silt and some organic matter.  Mature red mangrove forests dominate the banks of the rivers.  Upstream there are minimal influences from the Picayune Strand State Forest with non-point source pollutants possible from the culverts under I-75 and US 41.  Fakahatchee Strand State Preserve and Big Cypress National Park manage the headwaters of Fakahatchee Bay.  Salinities range from 0.7 – 42.6 ppt and fluctuate daily with the tides and seasonal rainfall.  The tide range for Fakahatchee varies between 0.2 and 1.8 meters.

UTM GPS coordinates (using UTM zone 17, datum NAD-83):

E  2863829

N  452224

Lat/Long:

25.8922  N

81.4770  W

To reference these sites on a nautical chart, see NOAA chart #11430 – “Lostman’s River to Wiggins Pass”.

6) Data collection period

Between January & December 2007,  23 sonde deployments occurred at Lower Henderson Creek; 23 deployments occurred at Middle Blackwater River;  23 deployments in Faka Union Bay; and  23 deployments in Fakahatchee Bay.

Lower Henderson Creek (rkblhwq):

1st Reading:						Last Reading:

01/11/07	07:30	01/29/07	09:00

02/08/07	12:30	02/22/07	07:15

02/22/07	07:30	03/08/07	08:15

03/08/07	08:30	03/22/07	11:45

03/22/07	12:15	04/09/07	08:15

04/09/07	08:31*	04/19/07	06:46*

04/19/07	07:45	05/10/07	04:45

05/10/07	see section 14

05/24/07	05:46*	06/07/07	05:31*

06/07/07	see section 14	06/21/07	06:30	07/03/07	07:30

07/12/07	06:15	07/26/07	05:45

07/26/07	06:00	08/09/07	06:15

08/09/07	06:45**	08/23/07	6:15

8/23/2007 06:30	 09/06/07 	06:00

09/06/07	06:30**	09/20/07	06:15

- 09/20/07 06:45	-10/04/07  06:15

10/04/07	06:30**	10/18/07	06:30

10/18/07	06:44*, **	11/01/07	05:44*

11/01/07	05:59*, **	11/14/07	06:44*

11/14/07	06:59*, **	12/05/07	09:14*

12/05/07	09:29*, **	12/20/07	10:44*

12/20/07	11:00**	01/03/07	00:15

Middle Blackwater River (rkbmbwq):

1st Reading:					Last Reading:

01/11/07	10:00	02/08/07	09:30

02/08/07	09:45	02/22/07	08:45

02/22/07	09:00	03/08/07	09:15

03/08/07	09:30	03/22/07	09:00

03/22/07	09:15	04/05/07	08:30

04/05/07	08:45	04/19/07	10:00

04/19/07	10:15	05/10/07	08:15

05/10/07	08:31*	05/24/07	07:31

05/24/07	07:45	06/07/07	08:15

06/07/07	08:30	06/21/07	08:15

06/21/07	08:30**	07/12/07	08:15

07/12/07	08:45	07/26/07	07:15

07/26/07	07:29*, **	08/09/07	08:14*

08/09/07	08:30	08/23/07	08:15

08/23/07	08:29*, **	09/06/07	07:44*

09/06/07	08:01*	09/20/07	08:29*

09/20/07	08:29*, **	10/04/07	08:29*

10/04/07	08:45	10/18/07	08:45

10/18/07	09:00	11/01/07	07:45

11/01/07	08:00	11/14/07	08:45

11/14/07	09:00	12/05/07	10:00

12/05/07	10:16*	12/20/07	11:01*

12/20/07	11:15	01/06/08	11:45

Fakahatchee Bay (rkbfbwq):

1st Reading:					Last Reading:

01/11/07	10:30	02/08/07	10:30

02/08/07	10:45	02/2207	09:45

02/22/07	10:00	03/08/07	10:45

03/08/07	11:00	03/23/07	07:30

03/23/07	07:45	04/05/07	09:45

04/05/07	10:00	04/19/07	11:15

04/19/07	11:30	05/10/07	09:30

05/10/07	09:46*	05/24/07	08:31*

05/24/07	08:46*	06/07/07	09:31*

06/07/07	09:45	06/21/07	09:30

06/21/07	09:45	07/12/07	09:15

07/12/07	09:30	07/26/07	08:15

07/26/07	08:30	08/09/07	09:15

08/09/07	09:45	08/23/07	09:15

08/23/07	09:30	09/06/07	08:45

09/06/07	09:00	09/20/07	09:15

09/20/07	09:30	10/04/07	09:30

10/04/07	09:45	10/18/07	09:45

10/18/07	09:45	11/01/07	08:45

11/01/07	09:00**	11/14/07	09:30

11/14/07	09:59*, **	12/05/07	11:29*

12/05/07	12:00**	12/20/07	12:00

12/20/07	12:30**	01/06/08	13:30

Faka Union Bay (rkbfuwq):

1st Reading:					Last Reading:

01/11/07	10:30	02/08/07	10:00

02/08/07	10:30	02/22/07	09:30

02/22/07	09:45	03/08/07	10:15

03/08/07	10: 46*	03/23/07	07:46*

03/23/07	08:00	04/04/07	17:30

04/05/07	09:45	04/19/07	11:00

04/19/07	11:16*	05/10/07	09:01*

05/10/07	09:15	05/24/07	08:15

05/24/07	08:30	06/07/07	09:00

06/07/07	09:15	06/21/07	09:15

06/21/07	10:01*	07/12/07	09:01*

07/12/07	09:15	07/26/07	08:00

07/26/07	08:15	08/09/07	09:00

08/09/07	09:30	08/23/07	09:00

08/23/07	09:15	09/06/07	08:30

09/06/07	08:45	09/20/07	09:00

09/20/07	09:15	10/04/07	09:15

10/04/07	09:30	10/18/07	09:30

10/18/07	09:45	10/31/07	19:30

11/01/07	08:46*	11/14/07	09:31*

11/14/07	09:45	12/05/07	11:15

12/05/07	11:30	12/20/07	11:45

12/20/07	12:00	01/06/07	13:15

*miss-timed

** DO data collected using ROX probe

7. Distribution

According to the Ocean and Coastal Resource Management Data Dissemination Policy for the NERRS System-wide Monitoring Program, NOAA/ERD retains the right to analyze, synthesize and publish summaries of the NERRS System-wide Monitoring Program Data.  The PI retains the right to be fully credited for having collected and processed the data.  Following academic courtesy standards, the PI and NERR site where the data were collected will be contacted and fully acknowledged in any subsequent publications in which any part of the data are used. Manuscripts resulting from this NOAA/OCRM supported research that are produced for publication in open literature, including refereed scientific journals, will acknowledge that the research was conducted under an award from the Estuarine Reserves Division, Office of Ocean and Coastal Resources Management, National Ocean Services, National Oceanic and Atmospheric Administration.  The data set enclosed within this package/transmission is only as good as the quality assurance and quality control procedures outlined by the enclosed metadata reporting statement.  The user bears all responsibility for its subsequent use/misuse in any further analysis or comparisons. The Federal government does not assume liability to the Recipient or third person, nor will the Federal government reimburse or indemnify the Recipient for its liability due to any losses resulting in any way from the use of this data.

NERR water quality data and metadata can be obtained from the Research Coordinator at the individual NERR site (please see Section 2.2-1 Principal investigators and contact persons), from the Data Manager at the Centralized Data Management Office (please see personnel directory under the general information link on the CDMO home page) and online at the CDMO home page http://cdmo.baruch.sc.edu/ .  Data are available in text tab-delimited format.

8. Associated researchers & projects

The RKB NERR water quality monitoring project has been incorporated into the Picayune Strand Restoration Project’s performance measures. The data is also currently being used by several visiting investigators to support their work within the reserve. An annual summary report on the existing monitoring stations within the Reserve boundaries is also available upon request. In addition to this water quality dataset, the Reserve also participates in the SWMP long-term meteorological and nutrient monitoring data sets.

Greg Ellis, University of South Florida, Tampa, FL

Carbon and Nitrogen Isotopic Linkages Between Water-Column Particulates and Benthic Primary Consumers as a Means of Assessing Current and Historical Estuarine Trophic Conditions in Southwest Florida.   This study looks at the isotopic composition of particulate organic matter of primary consumers (Crassostrea virginica) to examine dominant sources of primary production at various points along a salinity gradient within Rookery Bay estuary.

Jane Caffrey, University of West Florida, Pensacola, FL

Ammonia Oxidation and Ammonia-oxidizing Bacteria and Archaea populations from estuaries with differing histories of hypoxia.  This study compares the abundance of ammonia oxidizing archaea (AOA) and ammonia oxidizing bacteria (AOB) amoA genes, along with potential nitrification rates, salinity, dissolved oxygen history, sediment chlorophylla, pore ammonium and sulfide concentrations to identify factors influencing the abundance and activity of nitrifiers in sediments from 6 different estuarine reserves.

Jian Geng – University of South Florida, Tampa, FL.

Hydrodynamic Modeling of Rookery Bay.  This study utilizes existing water quality data and new flow data to model the circulation of Rookery Bay driven by rivers, tides and winds, with the emphasis on the salinity response to freshwater inflows.

II.  Physical Structure Descriptors

9. Sensor specifications

YSI 6600EDS data logger (deployed at MB, FU, FH)

YSI 6600 – V2 EDS data logger (deployed at LH)

Parameter: Temperature

Units: Celsius (C)

Sensor Type: Thermistor

Model#: 6560

Range: -5 to 45 ºC

Accuracy: +/-0.15 ºC

Resolution: 0.01 ºC

Parameter: Conductivity

Units: milli-Siemens per cm (mS/cm)

Sensor Type: 4 –electrode cell with autoranging

Model#:6560

Range: 0 to 100mS/cm

Accuracy: +/-0.5% of reading + 0.001 mS/cm

Resolution: 0.001 mS/cm to 0.1 mS/cm (range dependant)

Parameter: Salinity

Units: parts per thousand (ppt)

Sensor Type: Calculated from conductivity and temperature

Range: 0 to 70 ppt

Accuracy: +/-1.0% of reading or 0.1 ppt, whichever is greater

Resolution: 0.01 ppt

Parameter: Dissolved Oxygen % saturation

Units: percent air saturation (%)

Sensor Type: Rapid Pulse – Clark type, polargraphic

Model#: 6562

Range: 0 to 500% air saturation

Accuracy: 0-200% air saturation: +/-2% of the reading or 2% air saturation, whichever is greater, 200 – 500% air saturation: +/-6% of the reading

Resolution: 0.1% air saturation

or

Parameter: Dissolved Oxygen % Saturation

Sensor Type: Optical probe w/ mechanical cleaning

Model#: 6150 ROX

Range: 0 to 500% air saturation

Accuracy: 0-200% air saturation: +/-1% of the reading or 1% air saturation, whichever is greater 200-500% air saturation: +/-15% of reading

Resolution: 0.1% air saturation

Parameter: Dissolved Oxygen mg/L (Calculated from % air saturation, temperature, and salinity)

Units: milligrams/Liter (mg/L)

Sensor Type: Rapid Pulse - Clark type, polargraphic

Model#: 6562

Range: 0 to 50 mg/L

Accuracy: 0-20 mg/L: +/- 2% of the reading or 0.2 mg/L, whichever is greater

20 to 50 mg/L: +/- 6% of the reading

Resolution: 0.01 mg/L

or

Units: milligrams/Liter (mg/L)

Sensor Type: Optical probe w/ mechanical cleaning

Model#: 6150 ROX

Range: 0 to 50 mg/L

Accuracy: 0-20 mg/L: +/-0.1 mg/l or 1% of the reading, whichever is greater

20 to 50 mg/L: +/- 15% of the reading

Resolution: 0.01 mg/L

Parameter: Non-vented Level – Shallow Depth

Units: meters  (m)

Sensor Type: Stainless steel strain gauge

Range: 0 – 30 ft (9.1 m)

Accuracy: +/-0.006 ft (0.018 m)

Resolution: 0.001 ft or (0.001 m)

Parameter: pH (bulb type)

Units: units

Sensor Type: Glass combination electrode

Model#: 6561

Range: 0 to 14 units

Accuracy: +/-0.2 units

Resolution: 0.01 units

Parameter: Turbidity

Units: nephelometric turbidity units (NTU)

Sensor Type: Optical, 90º scatter, with mechanical cleaning

Model#: 6136

Range: 0 to 1000 NTU

Accuracy: +/-2% reading or 0.3 NTU (whichever is greater)

Resolution: 0.1 NTU

Dissolved Oxygen Qualifier:

The reliability of the dissolved oxygen (DO) data after 96 hours post-deployment for non-EDS (Extended Deployment System) data sondes may be problematic due to fouling which forms on the DO probe membrane during some deployments (Wenner et al. 2001).  Many reserves have upgraded to the YSI 6600 EDS data sondes, which increases DO accuracy and longevity by reducing the environmental effects of fouling.  The user is therefore advised to consult the metadata and to exercise caution when utilizing the DO data beyond the initial 96-hour time period.  However, this potential drift is not always problematic for some uses of the data, i.e. periodicity analysis.   It should also be noted that the amount of fouling is very site specific and that not all data are affected.  The Research Coordinator at the specific NERR site should be contacted concerning the reliability of the DO data because of the site and seasonal variation in the fouling of the DO sensor. Only EDS sondes were in use during 2007, but both rapid-pulse (membrane) and ROX (optical) type DO sensors were used.  This qualifier refers to data collected with the rapid-pulse type sensor.  Please see the data collection period section for more information on the sensor used for each deployment.

Depth Qualifier:

The NERR System-Wide Monitoring Program utilizes YSI data sondes that can be equipped with either depth or water level sensors.  Both sensors measure water depth, but by convention, level sensors refer to atmospherically vented measurements and depth refers to non-vented measurements.  Readings for both vented and non-vented sensors are automatically compensated for water density change due to variations in temperature and salinity; but for all non-vented depth measurements, changes in atmospheric pressure between calibrations appear as changes in water depth.  The error is equal to approximately 1.03 cm for every 1 millibar change in atmospheric pressure, and is eliminated for level sensors because they are vented to the atmosphere throughout the deployment time interval.

10. Coded variable code definitions.

Sampling Station		Sampling site code		Station code

Lower Henderson Creek		LH			rkblhwq

Middle Blackwater River		MB			rkbmbwq

Faka Union Bay			FU			rkbfuwq

Fakahatchee Bay			FB			rkbfbwq

11. QAQC flag definitions:

QAQC flags provide documentation of the data and are applied to individual data points by insertion into the parameter’s associated flag column (header preceded by an F_).   During primary automated QAQC (performed by the CDMO), -5, -4, and -2 flags are applied automatically to indicate data that is missing and above or below sensor range.  All remaining data are then flagged 0, passing initial QAQC checks.   During secondary and tertiary QAQC 1, -3, and 5 flags may be used to note data as suspect, rejected due to QAQC, or corrected.

-5	Outside High Sensor Range

-4	Outside Low Sensor Range

-3		Data Rejected due to QAQC

-2		Missing Data

-1		Optional SWMP Supported Parameter

0		Data Passed Initial QAQC Checks

1		Suspect Data

2		Open - reserved for later flag

3		Open - reserved for later flag

4		Historical Data:  Pre-Auto QAQC

5		Corrected Data

12. QAQC Code definitions:

QAQC codes are used in conjunction with QAQC flags to provide further documentation of the data and are also applied by insertion into the associated flag column.  There are three (3) different code categories, general, sensor, and comment.  General errors document general problems with the deployment of YSI data sonde, sensor errors are sensor specific, and comment codes are used to further document conditions or a problem with the data.  Only one general or sensor error and one comment code can be applied to a particular data point.

General Errors

GIC	No Instrument Deployed Due to Ice

GIM	Instrument Malfunction

GIT	Instrument Recording Error; Recovered Telemetry Data

GMC 	No Instrument Deployed Due to Maintenance/Calibration

GNF	Deployment Tube Clogged / No Flow

GOW	Out of Water Event

GPF	Power Failure / Low Battery

GQR	Data Rejected Due to QA/QC Checks

GSM	See Metadata

Sensor Errors

SBO	Blocked Optic

SCF	Conductivity Sensor Failure

SDF	Depth Port Frozen

SDO	DO Suspect

SDP	DO Membrane Puncture

SIC		Incorrect Calibration / Contaminated Standard

SNV	Negative Value

SOW	Sensor Out of Water

SPC	Post Calibration Out of Range

SSD	Sensor Drift

SSM	Sensor Malfunction

SSR	Sensor Removed / Not Deployed

STF	Catastrophic Temperature Sensor Failure

STS	Turbidity Spike

SWM	Wiper Malfunction / Loss

Comments

CAB	Algal Bloom

CAF	Acceptable Calibration/Accuracy Error of Sensor

CAP	Depth Sensor in Water, Affected by Atmospheric Pressure

CBF	Biofouling

CCU	Cause Unknown

CDA	DO Hypoxia (<3 mg/L)

CDB	Disturbed Bottom

CDF	Data Appear to Fit Conditions

CFK	Fish Kill

CIP		Surface Ice Present at Sample Station

CLT	Low Tide

CMC	In Field Maintenance/Cleaning

CMD	Mud in Probe Guard

CND	New Deployment Begins

CRE	Significant Rain Event

CSM	See Metadata

CTS	Turbidity Spike

CVT	Possible Vandalism/Tampering

CWD	Data Collected at Wrong Depth

13. Post deployment information:

| Post deployment information |  |  |  |  | pH |  | Turbidity |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | DO% | Depth | SpCond | 7 | 10 | 0 | 123 |  |
| File name | Date | % sat | meters | ms/cm | SU | SU | NTU | NTU | Comm |
| LH011107 | 1/29/2007 | 100.2 | -0.013 | 52.32 | 6.85 | 9.45 | -0.6 | 120.4 |  |
| LH020807 | 2/22/2007 | 71.0 | -0.118 | 49.43 | 7.08 | 10.07 | 0.3 | 109.4 | failed post-cal |
| LH022207 | 3/8/2007 | 100.6 | 0.021 | 49.64 | 7.22 | 10.27 | 0.6 | 127.2 |  |
| LH030807 | 3/22/2007 | 97.4 | -0.030 | 48.81 | 7.15 | 10.24 | 0.9 | 125.2 |  |
| LH032207 | 4/9/2007 | 101.0 | -0.061 | 51.91 | 6.93 | 9.96 | 0.2 | 123.3 |  |
| LH040907 |  |  |  |  |  |  |  |  | no post-cal see metadata |
| LH041907 | 5/10/2007 | 98.6 | 0.044 | 51.57 | 7.09 | 10.17 | 1.1 | 125.6 |  |
| HC051007 | 5/24/2007 | 99.3 | 0.065 | 50.77 | 7.08 | 10.07 | 1.2 | 119.5 | no file program error |
| LH052407 | 6/7/2007 | 101.3 | 0.014 | 52.42 | 7.17 | 10.23 | 1.2 | 126.9 |  |
| LH060707 | 6/21/2007 | 70.6 | -0.018 | 49.98 | 7.08 | 9.89 | 0.6 | 122.7 | no file program error |
| LH062107 | 7/12/2007 | 100.1 | 0.003 | 52.23 | 6.90 | 9.78 |  |  | Cable comunication problem |
| LH071207 | 7/26/2007 | 76.2 | -0.009 | 50.30 | 7.03 | 10.01 | 1.5 | 123.7 | DO failed post-cal |
| LH072607 | 8/9/2007 | 118.1 | -0.034 | 50.87 | 6.70 | 9.63 | 0.9 | 126.2 |  |
| LH080907 | 8/23/2007 | 99.8 | 0.004 | 50.20 | 6.81 | 9.81 | 1.1 | 133.9 | communication problems |
| LH082307 | 9/6/2007 | 101.4 | 0.009 | 50.82 | 7.22 | 10.34 | 2.3 | 121.7 |  |
| LH090607 | 9/20/2007 | 94.9 | -0.055 | 50.21 | 6.71 | 9.70 | -0.2 |  | out of 123 NTU standard –one point cal |
| LH092007 | 10/4/2007 | 102.7 | 0.002 | 48.88 | 6.68 | 9.75 | 0.9 |  | out of 123 NTU standard – one point cal |
| LH100407 | 10/18/2007 | 98.4 | 0.018 | 48.38 | 7.08 | 10.08 | 0.4 | 122.5 |  |
| LH101807 | 11/1/2007 | 102.9 | -0.087 | 50.33 | 7.06 | 10.31 | 1.9 | 124.1 |  |
| LH110107 | 11/14/2007 | 93.9 | 0.016 | 50.06 | 7.21 | 10.02 | 0.8 | 122.2 |  |
| LH111407 | 12/5/2007 | 98.5 | -0.054 | 49.87 | 7.41 | 9.47 | 1.4 | 121.9 |  |
| LH120507 | 12/20/2007 | 101.0 | 0.016 | 49.43 | 6.99 | 9.97 | 0.3 | 125.7 |  |
| LH122007 | 1/6/2008 | 101.8 | 0.085 | 50.60 | 6.97 | 9.95 | 1.7 | 129.0 |  |

| Post deployment information |  |  |  |  | pH |  | Turbidity |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | DO% | Depth | SpCond | 7 | 10 | 0 | 123 |  |
| File name | Date | % sat | meters | ms/cm | SU | SU | NTU | NTU | Comm |
| MB011107 | 2/8/2007 | 90.2 | -0.056 | 49.92 | 6.94 | 9.83 | 0.6 | 119.4 | DO did not go hi-lo |
| MB020807 | 02/220/07 | 99.4 | -0.052 | 50.13 | 7.07 | 10.04 | 3.3 | 123.8 |  |
| MB022207 | 3/8/2007 | 96.2 | -0.021 | 50.47 | 6.96 | 9.97 | 1.6 | 125.0 |  |
| MB030807 | 3/22/2007 | 96.8 | -0.033 | 49.13 | 7.02 | 10.11 | 1.5 | 123.3 |  |
| MB032207 | 4/5/2007 | 92.7 | -0.119 | 50.54 | 7.14 | 10.07 | 0.3 | 124.1 |  |
| MB040507 | 4/19/2007 | 97.3 | -0.096 | 51.89 | 7.11 | 10.16 | 2.8 | 123.6 |  |
| MB041907 | 5/10/2007 | 104.1 | 0.024 | 48.60 | 6.89 | 9.44 | -0.6 | 122.3 | DO did not go hi-lo |
| MB051007 | 5/24/2007 | 71.9 | 0.060 | 50.96 | 7.03 | 10.11 | 2.5 | 120.7 | DO failed post-cal |
| MB052407 | 6/7/2007 | 1.6 | 0.000 | 48.00 | 6.97 | 10.02 | 2.1 | 124.8 | DO electrode was black |
| MB060707 | 6/21/2007 | 99.9 | -0.032 | 51.00 | 7.01 | 10.03 | 1.9 | 120.5 |  |
| MB062107 | 7/12/2007 | 91.2 | 0.010 | 38.48 | 7.04 | 8.70 | 1.0 | 124.2 | Failed post-cal |
| MB071207 | 7/26/2007 | 101.7 | -0.011 | 50.27 | 7.11 | 10.06 | 1.7 | 123.1 |  |
| MB072607 | 8/9/2007 | 100.0 | -0.041 | 45.60 | 7.00 | 9.84 | 0.7 | 124.0 |  |
| MB080907 | 8/23/2007 | 99.5 | 0.006 | 50.00 | 7.06 | 10.14 | 0.8 | 122.0 |  |
| MB082307 | 9/6/2007 | 96.0 | 0.002 | 45.96 | 6.94 | 9.94 | 1.8 | 120.1 |  |
| MB090607 | 9/20/2007 | 84.9 | -0.065 | 50.55 | 7.15 | 9.99 | -0.6 |  | DO failed post-cal |
| MB092007 | 10/4/2007 | 98.2 | 0.009 | 39.97 | 7.15 | 10.11 | 0.8 |  |  |
| MB100407 | 10/18/2007 | 86.6 | 0.002 | 50.09 | 7.09 | 9.94 | 0.8 | 124.4 | DO failed post-cal |
| MB101807 | 11/1/2007 | 102.6 | -0.075 | 49.91 | 6.99 | 10.04 | 1.6 | 121.1 | DO did not go hi-lo |
| MB110107 | 11/14/2007 | 98.7 | 0.015 | 50.45 | 6.85 | 9.49 | 0.4 | 121.5 |  |
| MB111407 | 12/5/2007 | 100.7 | -0.056 | 49.96 | 6.92 | 10.18 | 1.3 | 123.0 |  |
| MB120507 | 12/20/2007 | 93.9 | -0.005 | 50.57 | 6.94 | 9.74 | 0.1 | 125.3 |  |
| MB122007 | 1/6/2007 | 100.6 | -0.010 | 50.21 | 6.87 | 9.89 | 0.2 | 126.9 | Turbidity std bad |

| Post deployment information |  |  |  |  | pH |  | Turbidity |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | DO% | Depth | SpCond | 7 | 10 | 0 | 123 |  |
| File name | Date | % sat | meters | ms/cm | SU | SU | NTU | NTU | Comm |
| FU011107 | 2/8/2007 | 96.3 | -0.045 | 49.5 | 6.97 | 9.74 | 1.2 | 118.2 | DO had trouble stabilizing |
| FU020807 | 2/22/2007 | 101.8 | -0.016 | 52.49 | 6.97 | 10.01 | 2.1 | 111.9 | Turbidity looked off |
| FU022207 | 3/8/2007 | 98.9 | -0.023 | 50.15 | 7.12 | 9.98 | 1.6 | 123.1 |  |
| FU030807 | 3/23/2007 | 98.0 | -0.014 | 51.48 | 6.98 | 10.06 | 0.8 | 125.0 |  |
| FU032207 | 4/5/2007 | 101.1 | -0.119 | 50.28 | 7.01 | 9.84 | 4.0 | 67.0 | Turbidity failed post-cal |
| FU040507 | 4/19/2007 | 96.7 | -0.094 | 51.86 | 7.01 | 10.12 | 0.7 | 123.7 |  |
| FU041907 | 5/10/2007 | 102.4 | 0.021 | 49.63 | 7.05 | 9.92 | 2.1 | 132.7 | turbidity probe malfunction |
| FU051007 | 5/24/2007 | 100.6 | 0.065 | 52.49 | 7.03 | 9.92 | 3.8 | 121.3 |  |
| FU052407 | 6/7/2007 | 102.3 | 0.004 | 49.41 | 7.05 | 9.99 | 1.8 | 123.1 |  |
| FU060707 | 6/21/2007 | 98.7 | -0.019 | 51.36 | 7.06 | 10.05 | 1.1 | 121.4 |  |
| FU062107 | 7/12/2007 | 104.9 | 0.011 | 49.36 | 6.99 | 9.64 | 1.5 | 124.1 |  |
| FU071207 | 7/26/2007 | 88.8 | -0.008 | 49.36 | 7.17 | 10.13 | 1.1 | 124.2 |  |
| FU072607 | 8/9/2007 | 102.0 | -0.042 | 50.14 | 6.95 | 9.88 | -2.4 | 72.8 | wiper fell off |
| FU082307 | 9/6/2007 | 101.8 | 0.001 | 50.1 | 6.91 | 9.81 | 1.9 | 123.3 |  |
| FU080907 | 8/23/2007 | 95.8 | 0.007 | 50.02 | 7.01 | 9.98 | 1.1 | 121.9 |  |
| FU090607 | 9/20/2007 | 98.1 | -0.053 | 50.35 | 6.97 | 10.08 | 0.3 |  |  |
| FU092007 | 10/4/2007 | 98.1 | -0.053 | 50.35 | 6.97 | 10.08 | 0.3 |  |  |
| FU100407 | 10/18/2007 | 99.8 |  | 49.11 | 6.89 | 9.75 | -0.1 | 123.2 |  |
| FU101807 | 11/1/2007 | 97.0 | -0.104 | 48.58 | 7.06 | 10.08 | 2.8 | 123.1 |  |
| FU110107 | 11/14/2007 | 100.6 | 0.015 | 51.68 | 6.81 | 9.70 | 0.6 | 122.3 |  |
| FU111407 | 12/5/2007 | 111.0 | -0.078 | 49.56 | 7.06 | 10.26 | 2.5 | 122.6 | O-ring off DO probe |
| FU120507 | 12/20/2007 | 99.6 | -0.003 | 49.89 | 6.78 | 10 | 0.2 | 126.7 |  |
| FU122007 | 1/6/2008 | 99.4 | -0.004 | 50.53 | 7.09 | 10.15 | -9.1 | 113.3 | Turbidity standard bad |

| Post deployment information |  |  |  |  | pH |  | Turbidity |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | DO% | Depth | SpCond | 7 | 10 | 0 | 123 |  |
| File name | Date | % sat | meters | ms/cm | SU | SU | NTU | NTU | Comm |
| FB011107 | 2/8/2007 | 100.1 | -0.049 | 48.81 | 6.95 | 9.88 | 0.9 | 120.9 | DO didn't go hi-lo |
| FB020807 | 2/22/2007 | 100.7 | -0.012 | 50.10 | 7.13 | 10.08 | 1.7 | 122.4 |  |
| FB022207 | 3/8/2007 | 116.6 | -0.002 | 50.93 | 6.97 | 9.90 | 0.6 | 123.3 |  |
| FB030807 | 3/23/2007 | 97.5 | -0.030 | 48.94 | 6.99 | 10.13 | -0.5 | 111.9 |  |
| FB032207 | 4/5/2007 | 99.6 | -0.121 | 50.45 | 6.91 | 9.79 | -0.2 | 122.0 |  |
| FB040507 | 4/19/2007 | 98.6 | -0.093 | 49.50 | 7.05 | 9.95 | 2.4 | 124.2 |  |
| FB041907 | 5/10/2007 | 95.5 | 0.015 | 50.20 | 6.98 | 9.97 | 1.6 | 123.1 |  |
| FB051007 | 5/24/2007 | 98.9 | 0.062 | 50.89 | 6.90 | 9.84 | 4.4 | 121.1 | DO didn't go hi-lo |
| FB052407 | 6/7/2007 | 103.3 | -0.005 | 49.83 | 7.08 | 10.03 | 2.2 | 123.4 |  |
| FB060707 | 6/21/2007 | 102.3 | -0.027 | 50.14 | 7.01 | 9.99 | 3.1 | 120.3 |  |
| FB062107 | 7/12/2007 | 102.8 | 0.007 | 50.62 | 6.99 | 9.57 | 0.8 | 124.5 |  |
| FB071207 | 7/26/2007 | 94.1 | -0.004 | 51.03 | 7.19 | 10.15 | -0.6 | 123.0 |  |
| FB072607 | 8/9/2007 | 92.8 | -0.040 | 50.81 | 7.08 | 10.12 | 0.9 | 125.2 |  |
| FB080907 | 8/23/2007 | 103.3 | 0.013 | 49.64 | 7.10 | 10.18 | 0.1 | 120.2 |  |
| FB082307 | 9/6/2007 | 101.3 | 0.025 | 49.53 | 7.09 | 10.06 | 1.7 | 122.0 |  |
| FB090607 | 9/20/2007 | 107.1 | -0.046 | 51.07 | 7.11 | 10.07 | 0.3 |  | ran out of 123NTU standard |
| FB092007 | 10/4/2007 | 101.2 | 0.000 | 49.21 | 7.12 | 10.02 | 1.4 |  | ran out of 123NTU standard |
| FB100407 | 10/18/2007 | 101.9 | 0.000 | 48.32 | 7.13 | 10.10 | 2.5 | 120.8 |  |
| FB101807 | 11/1/2007 | 93.9 | -0.099 | 51.50 | 7.08 | 10.21 | 1.0 | 122.4 |  |
| FB110107 | 11/14/2007 | 95.7 | 0.008 | 54.30 | 6.74 | 9.63 | 0.1 | 122.3 |  |
| FB111407 | 12/5/2007 | 100.7 | -0.061 | 49.99 | 7.12 | 10.36 | 1.8 | 121.2 |  |
| FB120507 | 12/20/2007 | 97.7 | -0.031 | 50.62 | 6.94 | 9.90 | 0.3 | 125.7 |  |
| FB122007 | 1/6/2008 | 100.8 | -0.017 | 49.63 | 7.00 | 9.94 | 1.9 | 126.3 |  |

14. Other Remarks/ Notes:

Data are missing due to equipment or associated specific probes not being deployed, equipment failure, time of maintenance or calibration of equipment, or repair/replacement of a sampling station platform.  Any NANs in the dataset stand for “not a number” and are the result of low power, disconnected wires, or out of range readings.  If additional information on missing data is needed, contact the Research Coordinator at the reserve submitting the data.

All stations, all data:

RKB depth calibrations for all sample sites were not standardized using the local atmospheric pressure at the time of calibration, as is SWMP protocol.  Depth sensors were calibrated using the 760 mm Hg barometric pressure default value.  Therefore, RKB depth data cannot be corrected for changes in atmospheric pressure and should be considered suspect.

RKB dissolved oxygen calibrations for all sample sites did not take into account the local barometric pressure at the time of calibration.  Therefore, a margin of error dependent on the difference in the barometric pressure from the default calibration value of 760 mm Hg has been introduced.  As a result, all dissolved oxygen data should be considered suspect.

Lower Henderson Creek (rkblhwq):

01/29/07 09:15 – 02/08/07 12:15- The data is missing because the sonde was retrieved in order to investigate why the data stopped transmitting on 1/21/07. The field cable that connected to the data transmitter was an old model causing water intrusion into the connector on the sonde and pin corrosion. The sonde was not re-deployed.

02/08/07 12:30 – 02/22/07 07:15 – The DO data was considered suspect because the DO wouldn’t stabilize at post-deployment calibration and read 71%, which is outside the acceptable range. This may have been due to heavy fouling.

02/27/07 12:15- The sonde was pulled out of the water to show a reporter and camera crew.

05/10/07 05:00 – 05/24/07 05:30 – The data is missing due to a possible program error.

06/07/07 05:45 – 06/21/07 06:15 – The data is missing  due to a program error; the  sample interval was set at 15 seconds rather than 15 minutes.

07/12/07 06:15 – 07/26/07 05:45 –The DO was considered suspect because it read 76.2% at post-calibration, which is outsideof the  acceptable range. This is possibly due to fouling.

07/26/07 06:00 – 08/09/07 06:15 – The DO was considered suspect because the DO charge (sensor diagnostic) was very high and out of range throughout deployment. The DO also read high (118.1%) at post-deployment calibration.

08/23/2007 06:45 - 08/23/2007 09:45- The data is missing possibly due to a telemetry system voltage problem or program error.

Data from 8/23 6:30 to 10/4 6:15 may be off by as much as 12 hours due to the sonde’s time being set incorrectly before deployment. These datetimestamps were corroborated with other metadata and corrected to the best of our ability, but some assumptions were made and there is a possibility that an error was made in this process.

Middle Blackwater River (rkbmbwq):

01/01/07 00:00 – 02/08/07 09:15 – The data flagged as missing are due to the sampling interval being set at 30 minutes rather than 15 minutes.

01/26/07 01:00 – 02/08/07 09:30- The DO data was considered suspect because the DO failed the post-deployment calibration diagnostic test (DO warm up test) and the DO charge (sensor diagnostic) fell below the acceptable range during deployment (< 25).

05/10/07 08:30 – 05/24/07 07:30 – The DO data was rejected because there was a tear in the DO membrane upon retrieval.

05/24/07 07:45 – 06/07/07 08:15 – The DO data was rejected because the DO failed the post-deployment calibration diagnostic test (DO warm up test), the DO charge (sensor diagnostic) was low and out of range, and the DO read 1.6%, which is outside the acceptable range. The probe had one black electrode and the membrane was possibly punctured,

On the dates and times listed below, the specific conductivity and salinity data should be considered suspect because this probe repeatedly had low post-deployment calibration values and low, out of range cell constant values (sensor diagnostic):

06/21/07 8:30 - 07/09/07 16:15

07/26/07 07:30 - 08/09/07 08:15

08/23/07 08:30 - 09/06/07 07:45

09/20/07 08:00 – 10/04/07 08:30

From 07/09/07 16:30until 07/12/07 08:15, the conductivity and salinity data were rejected because there was a clear downward trend in the data with the last reading being significantly disjunctive with the first reading of the newly deployed sonde. Also the last reading did not match well with the spot check values. The specific conductivity also failed at post-deployment calibration with a reading of 38.48 mS/cm in the 50 mS/cm standard, which is outside the acceptable range.

06/21/07 08:30 – 07/12/07 08:15 The DO data was considered suspect because there was wiper rotation failure and severe fouling.

09/06/07 08:00 – 09/20/07 08:15 – The DO data was considered suspect because the DO failed the post-deployment calibration diagnostic test (DO warm up test) and read 84.9% saturation, which isoutside the acceptable range.

10/04/07 09:00 – 10/18/07 08:45 – The turbidity probe was calibrated using a one- point calibration method (only the 0 NTU standard). Data passed QAQC.

Faka Union Bay (rkbfuwq):

01/01/07 00:00 – 02/08/07 10:15 –The data flagged as missing are due to the sample interval beingset at 30 minutes rather than 15 minutes.

03/23/07 17:45 – 04/04/07 17:30 – The DO data was considered suspect because the DO failed the post-deployment calibration diagnostic test (DO warm up test).  The turbidity failed at post–deployment calibration; the turbidity probe read 67 NTU in the 123 NTU standard, which is outside the acceptable range.

This was probably due to heavy fouling.

04/04/07 17:45 – 04/05/07 09:30 – Missing data, due to power failure during deployment.

04/19/07 11:30 – 05/10/07 09:00 – The DO data was considered suspect because the DO failed the post-deployment calibration diagnostic test (DO warm up test). The turbidity failed at post-deployment calibration; the turbidity probe read 132.7 NTU in the 123 NTU standard, which is outside the acceptable range.

07/12/07 9:15 - 07/26/07 08:00 – The DO data was considered suspect because the DO failed the post-deployment calibration diagnostic test (DO warm up test) and read 88.8% saturation, which is outside the acceptable range.

07/26/07 08:15 – 08/09/2007 09:00- The turbidity failed at post-deployment calibration; the turbidity probe read 72.8 NTU in the 123 NTU standard, which is outside the acceptable range.

11/14/07 09:45  - 12/05/07 11:15 – The DO data was considered suspect because the DO failed the post-deployment calibration diagnostic test (DO warm up test) and the o-ring was missing from probe upon retrieval.

12/20/07 12:00 – 12/31/07 23:45 – The turbidity data was rejected because turbidity failed at post-deployment calibration. The turbidity probe read -9.1 in the 0 NTU standard and 113.3 in the 123 NTU standard, which is outside the acceptable range. It is possible that the 0 NTU standard was contaminated at pre-deployment calibration causing inaccurate readings.

Fakahatchee Bay (rkbfbwq):

For the following deployments, the data flagged as missing are due to the sample interval being set at 30 minutes rather than 15 minutes:

01/01/07 00:00 – 2/08/07 10:15

11/01/07 09:15 – 11/14/07 09:45

12/05/07 11:45 – 12/20/07 09:30

02/22/07 10:00 – 03/08/07 10:45-  The DO data was considered suspect because the DO failed the post-deployment calibration diagnostic test (DO warm up test) and it read 116.6%, which is outside the acceptable range. The DO readings during this deployment didn’t match up with the deployments before and after it. Therefore, that data may be suspect.

On the below dates and times, the DO data was considered suspect because the DO failed the post-deployment calibration diagnostic test (DO warm up test): 04/05/07 10:00 -04/19/07 11:15, 06/07/07 09:45 – 06/21/07 09:30, 07/12/07 09:30 – 07/26/07 08:30

07/26/07 – 08/09/07 09:15-  The DO data was considered suspect because the DO failed the post-deployment calibration diagnostic test (DO warm up test) and the wiper was off on the bottom of the guard.

10/18/07 09:45  - Added ROX optical probe

Weather Conditions:

January: Due to a lack of cold fronts, significantly dry conditions prevailed across the region.

February: Drier than normal conditions prevailed across the region.

March: Continued high atmospheric pressure across the area.

April: Continued drought conditions prevailed, with scatter rain fall across the region the second and fourth week.

May: Some sea breeze activity generated some scattered showers, not enough to be significant.

June: Tropical storm Barry kicked off the rainy season, but didn’t deliver much rainfall and June continued to be one of the driest on record.

July: Although within “wet season”, the typical summer rain fall patterns still haven’t shown up, it continues to be dry for the month of July.

August: As hurricane Dean moved across the Caribbean drier air was forced across the state, therefore continuing the dry conditions. The summer time rainfall has yet to appear in August.

September: Although some summer time weather patterns started to appear across SW Florida, there continued to be lower than normal rain fall amounts.

October: Same.

November: Wide spread dry air from Tropical Storm Noel, and lack of cold fronts prevented rain fall activity.

December: A low atmospheric trough extended from the remnants of Tropical Storm Olga, generating some moderate rain fall amounts.

Acknowledgement: The data included with this document were collected by the staff of the Florida Department of Environmental Protection at the Rookery Bay National Estuarine Research Reserve with funding through NOAA’s Estuarine Research Division.  Any products derived from these data should clearly acknowledge this source (please use the attached logos).  This recognition is important for ensuring that this long-term monitoring program continues to receive the necessary political and financial support.