San Francisco Bay (SFB) NERR Meteorological Metadata

January 1, 2010 – December 31, 2010

Latest Update: 01/18/2013

I.  Data Set and Research Descriptors

1)  Principal investigator(s) and contact persons

Matt Ferner, Ph.D., Research Coordinator (November 2008 – Present)

San Francisco Bay NERR

3150 Paradise Drive

Tiburon, CA	94920

Ph: 415-338-3724

Email: mferner@sfsu.edu

Lara Martin, Monitoring Technician (2006-Present)

San Francisco Bay NERR

3150 Paradise Drive

Tiburon, CA	94920

Ph: 415-338-3719

Email: laram@sfsu.edu

Caitlin Coleman-Hulbert, Monitoring Technician (January 2010- Present)

San Francisco Bay NERR

3150 Paradise Drive

Tiburon, CA 94920

Ph: 415-338-3719

Email: cech@sfsu.edu

2)  Entry verification

Data (.DAT files) are uploaded from the CR1000 data logger to a Personal Computer (Windows compatible).  Files are exported from LoggerNet in a comma-delimited format and uploaded to the Centralized Data Management Office (CDMO) where they undergo automated primary QAQC and become part of the CDMO’s online provisional database.  During primary QAQC, data are flagged if they are missing or out of sensor range.  The edited file is then returned to the Reserve where it is opened in Microsoft Excel and processed using the CDMO’s NERRQAQC macro.  The macro inserts station codes, creates metadata worksheets for flagged data, and graphs the data for review.  It allows the user to apply QAQC flags to outliers, suspect and erroneous data identified during secondary QAQC based on review and comparisons with field readings, code the data, append files and export the resulting data file to the CDMO for tertiary QAQC and assimilation into the CDMO’s authoritative online database.  For more information on QAQC flags and QAQC codes, see Sections 11 and 12.

Persons responsible for data verification are the Monitoring Technician (Rachel Endicott from November 2008 to January 2010 and Caitlin Coleman-Hulbert from January 2010 to present) and the Research Coordinator (Matt Ferner from November 2008 to present).

3)  Research objectives

As part of the System Wide Monitoring Program (SWMP), a meteorological station is used to supplement the collection of water quality and biological data.  Overall goals of the San Francisco Bay NERR/SWMP meteorological monitoring effort are to collect weather data to: 1.) Document short-term variability and long-term changes in atmospheric conditions at the Reserve sites (e.g., localized impacts of seasonal storm events, interannual differences in rainfall, magnitude and influence of El Niño – La Niña events, etc.). 2.) Provide insight into patterns or occurrences in meteorological conditions that have direct relevance to estuarine communities and resource-management practices. 3.) Integrate these data into the broader Central and Northern California Ocean Observing System (CeNCOOS) that links to a wide variety of real-time data on water quality, currents and weather around the region.  Together these data provide an ecosystem-based network for understanding the environmental conditions and ecological processes of the San Francisco Bay.

4)  Research methods

The Campbell Scientific weather station samples a variety of parameters every five seconds continuously throughout the year.  High-resolution data are averaged over 15-minute intervals and recorded for air temperature (°C), relative humidity (%), barometric pressure (mb), wind speed (m/s), wind direction (degrees), and battery voltage (volts).  Maximum wind speed (m/s) from the five-second data and standard deviation of wind direction (degrees) are also recorded.  Maximum and minimum readings and their times from 5-second data for air temperature are available directly from the Reserve through the contacts provided above.  Fifteen minute totals are recorded for precipitation (mm), photosynthetically available radiation or PAR (millimoles/m2), and cumulative precipitation (mm).

The meteorological station is visited monthly to download the data to a laptop computer through direct connection via a nine-pin serial cable.  At the time of download, the current readings of the CR1000 are displayed and compared to readings from an external device (Kestrel 4000) to cross check sampling accuracy. The station is also maintained during these visits; the wind sensors are checked to make sure there are no obstructions and can move freely, and PAR, temperature/relative humidity sensors and the tipping rain bucket are cleaned as needed.  Sensors are calibrated on a one to two year schedule by swapping the existing sensors with ones that have been freshly calibrated, and by manually calibrating the precipitation gauge.  The following is a list of which sensors were swapped in 2010 and when they were calibrated:

Par from Jan - April     Calibrated 5/7/09

Par from April - Oct     Calibrated 3/11/10

Par from Oct - Dec      Calibrated 9/8/10 (consecutively with the 3/11/10 sensor)

BP             Calibrated 8/26/09

Wind        Sent in at same time as BP so probably Aug 2009

Temp/RH  Calibrated Aug 28 or 31, 2010

During 2010, the following programs were in use:

Jan 1- April 22       SANFRAN_V6_20080922

April 22-Oct 18      SANFRAN_V6.1_04022010

Oct 18 - Dec 31     SANFRAN_V6.1_NewPar101810

Campbell Scientific data telemetry equipment was installed at the Rush Ranch Station on 8/16/2006 and transmits data to the NOAA GOES satellite, NESDIS ID # 3B016066 (where 3B016066 is the GOES ID for that particular station).  The transmissions are scheduled hourly and contain four (4) data sets reflecting a series of 15-minute data-sampling intervals.  Upon receipt by the CDMO, data undergo the same automated primary QAQC process detailed in Section 2 above.  The “real-time” telemetry data become part of the provisional dataset until undergoing secondary and tertiary QAQC and assimilation in the CDMO’s authoritative online database.  Provisional and authoritative data are available at http://nerrsdata.org.

5)  Site location and character

The San Francisco Estuary has the largest watershed on the west coast.  There are four main bays in the San Francisco Bay portion of this system; Central Bay, South Bay, San Pablo Bay and Suisun Bay.  The San Francisco Bay National Estuarine Research Reserve (SFBNERR) is a multi-component site totaling approximately 3,700 acres.  Reserve sites include tidal salt marsh at China Camp State Park (38°0'4.49"N, 122°27'37.28"W) in San Pablo Bay and brackish tidal marsh at Rush Ranch Open Space Preserve (38°13'21.00"N, 122°01'33.60"W) in Suisun Bay. The two component sites include open water, tidal wetland, and adjacent upland buffer habitats and are well-protected remnants of the diverse plant and animal communities that once occupied the salinity gradient within the estuary.  Development within the watersheds of the sites ranges from predominantly urban/suburban to protected land/agricultural.

The meteorological station is located within the Rush Ranch Open Space Preserve in the transition zone from the brackish tidal marsh to the upland grassland habitat.  The station is placed just above mean higher high water in a flat grass area approximately 10 meters from marsh vegetation and is free from obstructions overhead or nearby. The latitude and longitude of the meteorological station are 38° 12’ 01.6 N, 122° 1' 35.3 W and its elevation is approximately 3.96 meters. The upland habitat is undeveloped aside from a nature center, historical barn, and horse corrals, all of which are almost a kilometer away from the meteorological station.  The surrounding grassland is used intermittently for agricultural grazing of cattle.

The tower itself is a Campbell Scientific 20 foot (6.096 m) tower. The sensors are located at 7.5 m (wind, temperature and humidity, and PAR), 2.5 m (barometric pressure), and 1 m (rain gauge) above the ground.  The rain gauge is a separate unit located about 3 m from the tower.

There are two water quality monitoring stations at Rush Ranch Open Space Preserve. The First Mallard station is located at the intersection of Cutoff Slough and First Mallard Creek (38° 11' 41.70 N, 122° 1' 58.02 W) and is approximately 3.13 km southwest of the weather station. The Second Mallard station is located at the intersection of Cutoff Slough and Second Mallard Creek (38° 11' 02.8 N, 122° 0' 48.9 W) and is approximately 4.52 km southeast of the weather station.

The weather station’s location is representative of the atmospheric conditions at the water quality stations at Rush Ranch.   Rush Ranch is located about 40 kilometers inland of the other Reserve site, China Camp State Park.  At China Camp there are two water quality monitoring stations and due to the distance from and distinct differences in climate at Rush Ranch, this meteorological station is not representative of the conditions at China Camp.

6)  Data collection period

Meteorological data has been collected routinely at Rush Ranch since August 2006.  Data was collected for all of 2010, January 1 00:00 through December 31 23:45.

7)  Distribution

NOAA/ERD retains the right to analyze, synthesize and publish summaries of the NERR SWMP data.  The PI retains the right to be fully credited for having collected and processed the data.  Following academic courtesy standards, the PI and NERR site where the data were collected will be contacted and fully acknowledged in any subsequent publications in which any part of the data are used.  Manuscripts resulting from this NOAA/OCRM supported research that are produced for publication in open literature, including refereed scientific journals, will acknowledge that the research was conducted under an award from the Estuarine Reserves Division, Office of Ocean and Coastal Resource Management, National Ocean Service, National Oceanic and Atmospheric Administration.  The data set enclosed within this package/transmission is only as good as the quality assurance and quality control procedures outlined by the enclosed metadata reporting statement.  The user bears all responsibility for its subsequent use/misuse in any further analyses or comparisons.  The Federal government does not assume liability to the Recipient or third persons, nor will the Federal government reimburse or indemnify the Recipient for its liability due to any losses resulting in any way from the use of this data.

SFBNERR meteorological data and metadata can be obtained from the Research Coordinator at the individual NERR site (please see Principal investigators and contact persons), from the Data Manager at the Centralized Data Management Office (please see personnel directory under the general information link on the CDMO home page) and online at the CDMO home page http://nerrsdata.org.  Data are available in comma-separated format.

8)  Associated researchers and projects

These meteorological data are part of the NERR SWMP. SFBNERR monitors a total of four water quality stations where data is collected continuously with YSI data sondes. There are two stations located at the China Camp State Park location and two at the Rush Ranch Open Space Preserve location. Nutrient sampling is also conducted at the four sites at water depths adjacent to the sondes. The Reserve is currently monitoring benthic invertebrates in both subtidal habitats (via trapping) and intertidal habitats (via quadrat surveys) with plans to expand these efforts over the coming year.

II. Physical Structure Descriptors

9)  Sensor specifications:

The CR1000 has 2 MB of Flash EEPROM that is used to store the Operating System.  Another 128 K Flash is used to store configuration settings.  A minimum of 2 MB SRAM is (4 MB optional upgrade) available for program storage (16K), operating system use, and data storage.  Additional storage is available by using a compact flash card in the optional CFM100 Compact Flash Module.

Date CR1000 installed:  August 16, 2006

Parameter: Temperature

Units: Celsius

Sensor type: Platinum resistance temperature detector (PRT)

Model #:  HMP45C Temperature and Relative Humidity Probe

Operating Temperature:  -40°C to +60°C

Range: -40°C to +60°C

Accuracy: ± 0.2 °C @ 20°C

Date of Last calibration:

2006: s/n A5250046 was installed (calibration date uknown)

12/05/2007: s/n C3610004 was installed (calibration date 09/03/2007)

09/22/2009: s/n A5250046 was installed (calibration date 08/09/2009)

09/23/2010: s/n F4610016 was installed (calibration date 08/28/2010)

Parameter: Relative Humidity

Units: Percent

Sensor type: Vaisala HUMICAP© 180 capacitive relative humidity sensor

Model #:  HMP45C Temperature and Relative Humidity Probe

Range:  0-100% non-condensing

Accuracy at 20°C:  +/- 2% RH (0-90%) and +/- 3% (90-100%)

Temperature dependence of RH measurement: +/- 0.05% RH/°C

Date of Last calibration:

2006: s/n A5250046 was installed (calibration date uknown)

12/05/2007: s/n C3610004 was installed (calibration date 09/03/2007)

09/22/2009: s/n A5250046 was installed (calibration date 08/09/2009)

09/23/2010: s/n F4610016 was installed (calibration date 08/28/2010)

Parameter: Barometric Pressure

Units: millibars (mb)

Sensor type: Vaisala Barocap © silicon capacitive pressure sensor

Model #: CS-105

Operating Range:	 Pressure: 600 to 1060 mb; Temperature:  -40°C to +60°C;

Humidity:  non-condensing

Accuracy: ± 0.5 mb @ 20°C; +/- 2 mb @ 0°C to 40°C; +/- 4 mb @ -20°C to 45°C; +/- 6 mb @ -40°C to 60°C

Stability: ± 0.1 mb per year

Date of Last calibration:

12/05/2007: s/n C2430021 was installed (calibration date 06/15/2007, multiplier 0.240/500)

09/22/2009: s/n A0240014 was installed (calibration date 08/26/2009, multiplier 0.184/600)

Parameter: Wind speed

Units: meter per second (m/s)

Sensor type: 18 cm diameter 4-blade helicoids propeller molded of polypropylene

Model #: R.M. Young 05103 Wind Monitor

Range:  0-60 m/s (134 mph); gust survival 100 m/s (220 mph)

Accuracy: +/- 0.3 m/s

Date of last calibration:

12/05/2007: s/n 81058 was installed (Campbell check date unknown)

09/22/2009: s/n 65477 was installed (Campbell check date 08/26/2009)

Parameter: Wind direction

Units: degrees

Sensor type: balanced vane, 38 cm turning radius

Model #: R.M. Young 05103 Wind Monitor

Range:  360° mechanical, 355° electrical (5° open)

Accuracy: +/- 3 degrees

Date of last calibration:

12/05/2007: s/n 81058 was installed (Campbell check date unknown)

09/22/2009: s/n 65477 was installed (Campbell check date 08/26/2009)

Parameter: Photosynthetically Active Radiation (PAR)

Units: mmoles m-2 (total flux)

Sensor type: High stability silicon photovoltaic detector (blue enhanced)

Model #: LI190SB

Light spectrum waveband:  400 to 700 nm

Temperature dependence: 0.15% per °C maximum

Stability:  <±2% change over 1 yr

Operating Temperature:  -40°C to 65°C; Humidity: 0 to 100%

Sensitivity:  typically 5 µA per 1000 µmoles s-1 m-2

Date of last calibration:

12/14/2007: s/n Q38103 was installed (calibration 06/12/2007, multiplier 1.27)

09/22/2009: s/n Q41616 was installed (calibration 09/12/2009, multiplier 1.2997)

04/02/2010: s/n Q38103 was installed (calibration 03/10/2010, multiplier 1.3245)

10/18/2010: s/n Q39861 was installed (calibration 09/08/2010, multiplier 1.4809)

Parameter: Precipitation

Units: millimeters (mm)

Sensor type: Tipping Bucket Rain Gauge

Model #:  TE525

Rainfall per tip: 0.01 inch

Operating range: Temperature: 0° to 50°C; Humidity: 0 to 100%

Accuracy:  +/- 1.0% up to 1 in./hr; +0, -3% from 1 to 2 in./hr; +0, -5% from 2 to 3 in./hr

Date of Last calibration: 12/05/2007, 10/30/2009, 10/18/2010,

10)  Coded variable definitions

Sampling Site Code 	Station Code		Sampling Station

RR			sfbrrmet		Rush Ranch Meteorological Station

11)  QAQC flag definitions

QAQC flags provide documentation of the data and are applied to individual data points by insertion into the parameter’s associated flag column (header preceded by an F_).   During primary automated QAQC (performed by the CDMO), -5, -4, and -2 flags are applied automatically to indicate data that is above or below sensor range, or missing.  All remaining data are then flagged 0, as passing initial QAQC checks.   During secondary and tertiary QAQC 1, -3, and 5 flags may be used to note data as suspect, rejected due to QAQC, or corrected.

-5	Outside High Sensor Range

-4	Outside Low Sensor Range

-3	Data Rejected due to QAQC

-2	Missing Data

-1	Optional SWMP supported parameter

0		Passed Initial QAQC Checks

1		Suspect Data

2		Open - reserved for later flag

3		Open - reserved for later flag

4		Historical Data:  Pre-Auto QAQC

5		Corrected Data

12)  QAQC code definitions

QAQC codes are used in conjunction with QAQC flags to provide further documentation of the data and are also applied by insertion into the associated flag column.  There are three (3) different code categories, general, sensor, and comment.  General errors document general problems with the CR1000, sensor errors are sensor specific, and comment codes are used to further document conditions or a problem with the data.  Only one general or sensor error and one comment code can be applied to a particular data point.

General Errors

GIM	Instrument Malfunction

GIT	Instrument Recording Error, Recovered Telemetry Data

GMC	No Instrument Deployed due to Maintenance/Calibration

GMT 	Instrument Maintenance

GPD	Power Down

GPF	Power Failure / Low Battery

GPR	Program Reload

GQR	Data Rejected Due to QA/QC Checks

GSM	See Metadata

Sensor Errors

SIC	Incorrect Calibration Constant, Multiplier or Offset

SIW	Incorrect Wiring

SMT	Sensor Maintenance

SNV	Negative Value

SOC	Out of Calibration

SSN	Not a Number / Unknown Value

SSM	Sensor Malfunction

SSR	Sensor Removed

Comments

CAF	Acceptable Calibration/Accuracy Error of Sensor

CCU	Cause Unknown

CDF	Data Appear to Fit Conditions

CRE	Significant Rain Event

CSM	See Metadata

CVT	Possible Vandalism/Tampering

13)  Other remarks/notes

MISSING DATA

Data are missing due to equipment or associated specific sensors not being deployed, equipment failure, time of maintenance or calibration of equipment, or repair/replacement of a sampling station platform.  Any NANs in the dataset stand for “not a number” and are the result of low power, disconnected wires, or out of range readings.  If additional information on missing data is needed, contact the Research Coordinator at the reserve submitting the data.

04/02/2010 10:30-11:15 <-2> [GMT] Data missing due to the operating system 18 upload.

04/30/2010 16:45 <-2> [GPR] Data missing due to program reload.

09/23/2010 12:15 – 13:00 <-2> [GMT] Data lost due to maintenance of sensors.

10/18/2010 11:00 – 12:00 <-2> [GMT] Data lost due to maintenance of sensors.

OPPERATING SYSTEM RELOAD

04/02/2010 10:30 - 13:15 Campbell Scientific operating system 18 and met station operating program SANFRAN_V6.1_04022010 were loaded along with a freshly calibrated PAR sensor.  Data was not collected between 10:30 and 11:15, 11:30-13:00 is flagged as rejected as the mast was lowered and the wind instrument adjusted.  13:15 is rejected as it is the first data point after the power up and does not accurately reflect the 15 minute average.

04/30/2010 16:45 – 17:00 Campbell Scientific operating system 19 was loaded.  The 16:45 data point was missed and the 17:00 data point rejected as it was the first data point after a program upload, and does not accurately reflect the 15 minute average.

10/18/2010 12:15 We uploaded a new program.

PAR

Small negative PAR values are within range of the sensor and are due to normal errors in the sensor and the CR1000 datalogger and are flagged <1> (CAF). The Maximum signal noise error for the LiCor sensor is +/- 2.214 mmoles/m2 over a 15 minute interval.

From 09/01/2008 to 07/26/2011 SFBNERR has experienced periods where the nighttime PAR readings have been significantly above 0.  After substantial troubleshooting, LI-COR definitively declares this problem was due to water intrusion.  This diagnosis is considered a sensor malfunction.  In accordance with CDMO/DMC policy, all SFBNERR PAR data that show signs of water intrusion (i.e. after the dry freshly calibrated PAR sensor begins recording elevated nighttime readings until the sensor is replaced) have been rejected, this includes day and night time readings.

04/02/2010 10:30 - 13:15 Freshly calibrated PAR sensor was installed. Multiplier was changed to 1.3245.

10/18/2010 11:00 – 13:15 We installed a freshly calibrated PAR sensor (multiplier 1.4809)

RELATIVE HUMIDITY

Relative Humidity data greater than 100 are within range of the sensor accuracy of +/-3%.

08/31/2010 07:00 – 07:15 <-3> [GQR] Relative humidity was over the acceptable 103 sensor range so the data was rejected.

09/17/2010 Multiple points <-3> [GQR] Relative humidity was over the acceptable 103 sensor range so the data was rejected.

09/20/2010 Multiple points<-3> [GQR] Relative humidity was over the acceptable 103 sensor range so the data was rejected.

09/23/2010 12:58 Swapped the temperature/relative humidity sensor for a freshly calibrated sensor.  By 13:00 it looked like the temperature had stabilized.  All data are missing 12:15 to 13:00.

RAIN GAUGE

Cumulative precipitation data are recorded from 00:00 to 23:59 with the daily total recorded at the midnight mark (00:00).  The midnight CumPrcp value is actually the total from the previous day.

12/03/2009 – 01/13/2010 17:00 <0> (CSM) During the January 13th weather station check the screen from the rain bucket was found on the ground.  This likely did not affect data collected during this time.

02/22/2010 11:00 – 11:45 TotPrcp = <-3> [GMT] & CumPrcp = <1>(CSM) 11:00 – 00:00 The rain gauge was disconnected to fix the wiring.  During this time we lost a ½ bucket of water.

04/30/2010 16:00 TotPrcp and 16:00 – 16:30 CumPrcp = <5> (CSM) During this weather station check the rain bucket was tipped at 15:30 so the precipitation logged was recalculated to “0” along with the daily average.

09/23/2010 13:15 – 09/24/2010 00:00 <5> (CSM) TotPrcp(09/23/2010 13:30-14:00) and CumPrcp recalculated to “0” as it has tipped many times at calibration.

9/23/2010 14:00 – 10/18/2010 12:00 [SOC] (CDF) The rain gauge has been off due to incomplete calibration.  Timestamps that had values during this time period were rejected because we cannot be certain how many tips occurred therefore do not know how much precipitation occurred.  All zeros were QAQC checked to make sure there was no precipitation that day then left unflagged.

10/18/2010 12:30 – 13:00 TotPrcp and CumPrc from  10/18/2010 12:30 – 10/19/00:00 <5> (CSM) Recalculated to “0” after calibration was complete.

11/16/2010 10:55 Rain gauge tipped and value recalculated to 0 for both TotPrcp (at 11:00)  and CumPrcp (11/16/2010 11:00-11/17/2010 00:00).

11/16/2010 – 12/17/2010 During the December 17th weather station check the screen from the rain bucket was found on the ground.  This likely did not affect data collected during this time.

12/17/2010 12:00 Rain gauge tipped with an almost full bucket.  Because it was raining that day, the data were retained and marked as suspect.

GENERAL

12/17/2010 12:00 – 02/01/2011 16:15  Data are flagged as {CSM} in the F_Record and <1> (CDF) in the data for  ATemp and  RH.   Sometime between the two weather station visits the mast dropped about a foot and a half and rotated  and the cross arm that hold the wind and PAR sensors was loosely swinging in the wind.  The exact time couldn’t be determined as all data looks normal.

All wind parameters have been flagged <-3> [SQR](CSM).   PAR data were already rejected.