South Slough (SOS) NERR Meteorological Metadata

January – December 2007

Latest Update: 7/23/2012

I.  Data Set and Research Descriptors

- Principal investigators and contact persons:

### Address: 	South Slough National Estuarine Research Reserve

### PO Box 5417

### Charleston, OR 97420

### Phone:  (541) 888-5558

### Fax:      (541) 888-2733

Contact Persons:

###### Dr. Steven Rumrill, Research Program Coordinator

### E-mail:  steve.rumrill@state.or.us; (541) 888-2581 ext. 302

Alicia Helms, Estuarine Monitoring Coordinator

E-mail: alicia.r.helms@state.or.us; (541) 888-2581 ext 304

Adam DeMarzo, Estuarine Monitoring Assistant

E-mail: adam.demarzo@state.or.us; (541) 888-2581 ext. 307

- Entry verification

CR10X

From January 1 – February 5, 2007, the meteorological information from each sensor was sampled every 5 seconds and stored on the Campbell Scientific, Inc. (CSI) CR10X datalogger.  Data were output to a file in three arrays: array 15 stores 15 minute averages and max and min data; array 60 stores hourly averages and max and min data; and array 144 stores daily averages and max and min data. The CDMO datalogger program (NERR_4.csi) was uploaded into the CR10X to control the sensors. The CR10X interfaces with the PC208W software supplied by Campbell Scientific. The station was hardwired into the Estuarine and Coastal Science laboratory IBM-compatible PC with Campbell Scientific NL100 network Link Interface device in 2003 (see 2003 metadata, no date reported), which allows the CR10x to communicate with a computer using TCP/IP over a local area network (Oregon Institute of Marine Biology).  The weather data was managed from the PC with CSI’s Real-Time Data Management Program; however, on Oct 6, 2006, the hardwire connection was removed.

CR1000

On February 5, 2007, a CR1000 datalogger was installed to replace the CR10x.   The local area network link was reestablished with the help of Jay Poucher, NERRS Telemetry Project Coordinator, on June 7, 2007 using LoggerNet software until 1/22/08 15:15.  From June-Dec 2007, data was retrieved from a real-time file that was set up on a lab computer, which was automatically uploading data from the CR1000 via the NL100 link.  After Jan 2008, several attempts to troubleshoot and connect with the NL100 device were made with no success and it was removed.

Data are uploaded from the CR1000 data logger to a Personal Computer (IBM compatible).  Files are exported from LoggerNet in a comma-delimited format and uploaded to the CDMO where they undergo automated primary QAQC and become part of the CDMO’s online provisional database.  During primary QAQC, data are flagged if they are missing or out of sensor range.  The edited file is then returned to the Reserve where it is opened in Microsoft Excel and processed using the CDMO’s NERRQAQC Excel macro.  The macro inserts station codes, creates metadata worksheets for flagged data and summary statistics, and graphs the data for review.  It allows the user to apply QAQC flags and codes to the data, append files, and export the resulting data file to the CDMO for tertiary QAQC and assimilation into the CDMO’s authoritative online database.  For more information on QAQC flags and QAQC codes, see Sections 11 and 12.

Steve Rumrill has managed all aspects of the weather component of SWMP since August 2001.  Sue Powell was responsible for the weather station from installation in August 2001 through December 16, 2005, including QA/QC processing of the data and compiling the metadata.  Ali Helms has been responsible for the weather station and data management since July 2006 with the help of Adam DeMarzo since October 2006.

- Research objectives:

The South Slough NERR operates the SWMP Meteorological Station to provide continuous measurements of local weather conditions affecting the South Slough estuary.  The nearest National Weather Service certified station is located at the North Bend Airport (43.34N, 124.32 W), 11.5 km NE from the mouth of South Slough.  At that location, the weather observers’ primary function is to provide current meteorological information to airplane pilots.  Data is available real-time every 15 minutes; however, 24-hour weather data is reported, including humidity, wind speed and wind direction (prevailing compass), pressure, and temperature.

Because the SWMP weather station is located near the mouth of South Slough, recorded data provides information more pertinent to the microclimates affecting the estuary.  The principal objective is to record long-term meteorological data that is particularly relevant to the South Slough estuary in order to gain a better understanding of climatic forcing and to observe any environmental changes or trends over time.  In addition, the station provides real-time data to assess the short-term effects of local weather on measurements of water parameters within the estuary.  The station is part of a framework to evaluate the relative contributions of oceanic forcing and watershed inputs to hydrodynamics, nutrient dynamics, and other ecological processes within the South Slough estuary.  Data can be used as 1) a vital reference of atmospheric data for various research projects at the Reserve – an integral part of the NERR mission to provide a platform for estuarine research, 2) for meteorological context or atmospheric forcing for 30 minute water quality data, and other long-term environmental monitoring at the Reserve

- Research methods:

Data collection

CR10X

The Campbell Scientific CR10X datalogger (in use Jan-Feb 5, 2007) records measurements every 5 seconds to produce 15 minute, hourly and daily averages of Air Temperature, Relative Humidity, Barometric Pressure, Wind Speed and Wind Direction. Precipitation and Photosynthetic Active Radiation (PAR) are summed and recorded as 15 minute, hourly and daily totals.  Array 15 stores fifteen minute averages and max and min data, array 60 stores hourly averages, and max and min data; and array 144 stores daily averages, and max and min data.

CR1000

The CR1000 datalogger (in use Feb 5, 2007 – present) records measurements of Air Temperature, Relative Humidity, Barometric Pressure, Wind Speed, Wind Direction, PAR, and Precipitation.  The data are automatically downloaded to a PC in the laboratory via a local network link.  A 4MB storage module was installed on 10/20/03 to increase the capacity of the CR10X data storage.  When the CR10X was replaced with the CR1000 datalogger, the storage module was removed since it was not compatible with the CR1000 datalogger.   When the direct network line fails, a laptop computer is used to download data directly from the CR1000.

The 15 minute Data are collected in the following formats for the CR1000:

Averages from 5-second data:

Air Temperature (°C), Relative Humidity (%), Barometric Pressure (mb), Wind Speed (m/s), Wind Direction (degrees),  Battery Voltage (volts)

Maximum and Minimum Air Temperature (°C) and their times from 5-second data (these data are available from the Reserve)

Maximum Wind Speed (m/s) and time from 5-second data (not a standard parameter until 2008)

Wind Direction Standard Deviation (degrees) (not a standard parameter until 2008)

Totals:

Precipitation (mm), PAR (millimoles/m2), and Cumulative Precipitation (mm) (not a standard parameter until 2008)

Datalogger

The CR10X and CR1000 are powered by a 12 V sealed lead acid battery.  Two solar power panels (40W and 20W) generate power during the daylight hours to maintain battery charge.  The panels also generate power to an Optimus 12V deep-cycle sealed gel battery (installed on 11/17/05) during the day to provide backup power source.  We removed the 40W solar panel and the backup 12 V deep-cycle battery in June 2007 to clean up the station and eliminate excess equipment.   During November 2007, the station began to have low power/battery voltage and in December (12/19/07 20:15) data loss began occurring intermittently due to the low battery voltage. The problem was solved by recharging the 12 V battery.  The 40 W panel will be reinstalled by Winter 2008.

Telemetry

Campbell Scientific data telemetry equipment was installed at the South Slough NERR SWMP Charleston Meteorological station on 11/17/05 and transmits data to the NOAA GOES satellite, NESDIS ID # 3B0032E0.  The transmissions are scheduled hourly and contain four (4) data sets reflecting fifteen minute data sampling intervals.  Upon receipt by the CDMO, the data undergoes the same automated primary QAQC process detailed in Section 2 above.  The real-time telemetry data become part of the provisional dataset until undergoing secondary and tertiary QAQC and assimilation in the CDMO’s authoritative online database.  Provisional and authoritative data are available at http://cdmo.baruch.sc.edu.

Sensor QAQC

The meteorological sensors, solar panel, and associated components are inspected monthly for damage or fouling and the solar panels, PAR sensor, and the rain gauge are inspected and cleaned if necessary.  The Spike Needle Strips (Bird-X, Inc.) that were installed on a frame around the RG-2000-C rain gauge to prevent bird roosting have been very effective.  A Kestrel 4000 handheld and a sling psychrometer are used as a general check of the sensors. In the laboratory, data recorded as close as possible to the field check are downloaded from the nearby North Bend Weather Station (relative humidity) or the NOAA NOS/CO-OPS Charleston Station (9432780) and compared to station data.  The NOAA/NOS tide station has all weather checks available except relative humidity.

Safety apparatus for ascending the tower is in accordance with OSHA regulations.  A sliding trolley runs in a track during ascension of the ladder.  The trolley is connected by a carabiner to a chest harness and belt.  While working on the tower platform, a 4-meter-long tether secures the person to a tower strut.

Sensor Calibration

Annually, all sensors were calibrated on site by Robb Hibbs of MeasureTek, Inc.  Beginning in Sept 2007, South Slough sensors that are calibrated biannually were exchanged with spare sensors from the CDMO and South Slough sensors were mailed off for calibration.  The barometric pressure, PAR, and wind sensor are calibrated every 2 years while the temperature/relative humidity (mailed off for calibration) and rain gauge (calibrated with static method on site by staff) are calibrated annually. Data recorded during sensor calibrations/sensor removal are rejected.

Details for 2007 Sensor Calibration:

The following sensors were removed for calibration and replaced with spare CDMO sensors-Barometric Pressure, Relative Humidity & Temperature on 8/8/07 and Wind Speed, Wind Direction, and PAR/Licor Quantum Sensor on 8/9/07.  The CDMO temp/rel humidity sensor was removed and the calibrated SOS Temp/Rel humidity sensor was replaced on 11/1/07.  The remaining CDMO sensors were removed and the calibrated South Slough sensors were replaced on 11/5/07. The temp/relative humidity and the wind direction/speed sensors were incorrectly wired after calibration and were fixed on 11/7/07.

Weather programs

The following records tracks the weather program updates and includes the date that it was uploaded to the CR1000, program version name, and a brief description of the reason for the program change.

2/07/07   SOUTHSLOUGH_v3.0_020707.CR1-wrong precip multiplier 0.254

3/05/07  15:30 SOUTHSLOUGH_v3.0_022207.CR1-corrected precip multiplier from 0.254 to 0.1 for RG 2000C

5/04/07 13:45   SOUTHSLOUGH_v3.0_050207.CR1-Added TE525 WS rain gauge on P2; changed multiplier for new precip. gauge TE525 WS (0.254), put old rain gauge on standby.

6/06/07  11:30 SOUTHSLOUGH_v4.0_060107.CR1-Fixed an error with Wind data caused by time changes near noon

8/08/07 14:30   SOUTHSLOUGH_v4.1_061807.CR1-added pressure in inches of mercury to OIMBMET table; Added “DailyRain” to the GOES table so a new parameter should “appear” at HADS and the CDMO as “Cumulative Rainfall”; remarked measurements for the CDMO sensor inventory to accommodate local sensor calibration; added OIMBMET table including Cumulative Rainfall as well as all the other MET measurements in both Metric and English units.

8/09/07 10:15 SOUTHSLOUGH_v4.1_070907.CR1-CDMO sensors online/south slough sensors standby (removed for calibration)

11/7/2007 10:45 SOUTHSLOUGH_v4.1_110507.CR1- south slough sensors online (calibrated)/CDMO sensor standby.

- Site location and character:

South Slough is characterized as a drowned river-mouth estuary.  Located 8.5 km southwest of Coos Bay, Oregon, South Slough is one of seven tidal inlets that collectively form the Coos estuary.  The South Slough National Estuarine Research Reserve is representative of the mid-Pacific portion of the Lower Columbian Bioregion, extending from Cape Mendocino, CA north to the Columbia River, OR.  The Reserve includes 4,700 acres with 600 acres of tidal marshes, mudflats and open water channels.

Within the South Slough, there are three primary sub-systems that provide much of the freshwater input to the area.  Only two sub-systems, the Sengstacken and Winchester arms, are within Reserve boundaries.  The Sengstacken arm drains a semi-pristine area with no development and serves as a reference site for the riverine portion of the estuary.  The Winchester arm is similarly located in an area with development limited to one small farm (<10 cattle) high in the watershed.  The area functions as a management-treatment site for salt marsh restoration activities in the Winchester Tidelands Restoration Project.  Valino Island marks the northern boundary of the Reserve and is a marine-dominated area.

Charleston Meteorological Station:  [43 20.703’ N, 124 19.724’ W]

The station is located 1.2 km northwest of the Charleston water quality monitoring site.  A 10-meter tall navigation tower that formerly belonged to the Corps of Engineers is situated on the northeast corner of the Oregon Institute of Marine Biology campus.  The tower is about 20 meters from the inner boat marina at Charleston harbor.  Three 12-13 meter tall shore pines lie directly south of the tower and may slightly shelter the station from southerly winds, but the sensor array is never shaded.  The weather station is above the nearest building, which is about 8 meters away.  A 40-meter-tall forested ridge that has a northeast to southwest alignment is located about 60 meters from the weather tower and may interfere with winds from the southwest.  On the tower platform, a 3-meter tall mast holds the Wind Monitor and PAR sensors array.  The barometric pressure sensor is mounted inside the enclosure.  The Air Temperature/Relative Humidity Sensor is mounted on one of the triangular tower rods, just above the enclosure on the north side.  Because there is no open field available, the Rain Gauge is set on top of a platform on the southwest corner of the tower.  Sensor heights from the ground are Air Temperature/Relative Humidity 9.68 m, Rain 8.73 m, Barometric Pressure 9.18 m, Wind 11.08 m, and PAR 11.08 m.

- Data collection period:

The Charleston station was installed on 08/21/01.  The station was hardwired to the SWMP laboratory in 2003, and the data were downloaded automatically until October 2006.  On Oct 6, 2006, the hardwire connection was removed to provide for personnel space/equipment use.  Jay Poucher reestablished the real-time lan link on June 7, 2007 using LoggerNet software; however, it failed on December 19, 2007 and intermittently transmitted data to the real-time file until 1/15/08 15:00.  After several attempts to troubleshoot and connect with the NL100 device failed, it was removed.  Data was retrieved directly from the CR1000 instrument with a laptop and RS232 cable (Feb 2007-June 2007).

Equipment was installed on 11/17/05 to telemeter the meteorological data through the NOAA GOES satellite system.

Data were collected for most of the year in 2007, from January 1, 2007 at 00:00 to December 31, 2007 at 23:45.  All raw data from the CR1000 is missing from 2/5/07 13:45 to 3/5/07 15:15, 11/6/07 11:15 – 11/7/07 10:30, 12/19/07 20:15-20:30, 12/19/07 21:15-12/20/07 10:00, 12/20/07 23:30 – 12/21/07 10:00,  12/21/07 19:00-12/24/07 10:00, 12/24/07 10:30, 12/24/07 18:30-12/26/07 10:00, 12/26/07 10:30, 12/26 12:15, 12/26/07 18:00-12/30/07 9:45, 12/30/07 10:45 –11:45, 12:15 – 14:30, 12/30/07 16:00-12/31/07 9:15.

- Distribution

According to the Ocean and Coastal Resource Management Data Dissemination Policy for the NERRS System-wide Monitoring Program:  NOAA/ERD retains the right to analyze, synthesize and publish summaries of the NERRS System-wide Monitoring Program data.  The PI retains the right to be fully credited for having collected and processed the data.  Following academic courtesy standards, the PI and NERR site where the data were collected will be contacted and fully acknowledged in any subsequent publications in which any part of the data are used.  Manuscripts resulting from the NOAA/OCRM-supported research that are produced for publication in open literature, including refereed scientific journals, will acknowledge that the research was conducted under an award from the Estuarine Reserves

Division, Office of Ocean and Coastal Resource Management, National Ocean Service, National Oceanic and Atmospheric Administration.  The data set enclosed within this package/transmission is only as good as the quality assurance/quality control procedures outlined by the enclosed metadata reporting statement.  The user bears all responsibility for its subsequent use/misuse in any further analyses or comparisons.  The Federal government does not assume liability to the Recipient or third persons, nor will the Federal government reimburse or indemnify the Recipient for its liability due to any losses resulting in any way from the use of this data.

NERR weather data and metadata can be obtained from the Research Coordinator at the individual NERR site (please see Section 1 - Principal investigators and contact persons), from the Data Manager at the Centralized Data Management Office (please see personnel directory under the general information link on the CDMO home page) and online at the CDMO home page http://cdmo.baruch.sc.edu/.  Data are available in comma separated format.

- Associated researchers and projects:

The South Slough NERR maintains four water quality monitoring stations along the estuarine gradient within the South Slough estuary.  These stations collect continuous information to address three primary topics:  1) short-term variability and long-term changes in estuarine water parameters within different regions of the estuary (i.e., localized impacts of seasonal storm events, variability in water quality parameters in association with re-establishment of tidal circulation, interannual differences in rainfall, magnitude and influence of El Niño – La Niña events, spatial extent of oceanic and tidal forcing, and long-term changes associated with the Pacific Decadal Oscillation, etc.); 2) verification, calibration, and future refinement of an existing two-dimensional water quality model (CE-QUAL-W2) developed to predict levels of coliform bacteria and dispersion of environmental toxins; and 3) collection of fundamental baseline data for future development of a three-dimensional tidal hydrodynamic model for the greater South Slough/Coos Bay estuaries.

A water chemistry component is also associated with the core monitoring program.  Variables collected include: ammonium, nitrite, nitrate, nitrate + nitrate, orthophosphate, dissolved inorganic nitrogen, silicate, chlorophyll a, and phaeophytin.  Grab samples are collected for both high and low tides at the four water quality sampling sites, and automated samplers collect diel samples to compare marine and riverine sites during wet and dry seasons.

The SSNERR water quality monitoring project has been incorporated into the State of Oregon Coastal Non-Point Pollution Control Program (S. 6217) developed by the Oregon Department of Land Conservation and Development and the Oregon Department of Environmental Quality in July 1995.

#### II. Physical Structure Descriptors

- Sensor specifications, operating range, accuracy, date of last calibration

Parameter:Temperature

Units:  Celsius

Sensor Type:  Platinum resistance temperature detector (PRT)

Model #: HMP45C Temperature and Relative Humidity probe, changed to HMP35C on11/17/05 and returned to manufacturer for repair

Operating Temperature:  -40°C to +60°C

Range: -40°C to +60°C

Accuracy:  ± 0.2 °C @ 20°C

Date of Last calibration: 08/28/07

Parameter:Relative Humidity

Units:Percent

Sensor Type:Vaisala HUMICAP© 180 capacitive relative humidity sensor

Model #:HMP45C Temperature and Relative Humidity Probe, changed to HMP35C on 11/17/05 and returned to manufacturer for repair

Range:	0 - 100% non-condensing

Accuracy at 20oC: ± 2% RH (0-90%) and ± 3% (90-100%)

Temperature dependence of RH measurement:± 0.05% RH/°C

Date of Last calibration: 08/28/07

CDMO sensor:

Parameter:  Air temperature and relative humidity

HMP45C (SN Z4840089) installed from 8/8/2007-11/1/2007

Parameter:Barometric Pressure

Units:millibars (mb)

Sensor Type:Vaisala Barocap © silicon capacitive pressure sensor

Model #:CS-105

Operating Range:Pressure: 600-1060 mb; Temperature: -40 °C to +60°C;

Humidity:non-condensing

Accuracy: ±0.5 mb at 20°C; ± 2 mb 0°C to 40°C; ± 4 mb @ -20°C to 45°C; ± 6 mb -40°C to 60°C

Stability: ± 0.1 mb per year

Date of Last calibration: 09/04/07

CDMO sensor:

Parameter:  Barometric Pressure

PTB101B  (SN Z4950018) installed from 8/8/2007 – 11/5/2007

Parameter:Wind Speed

Units:meters per second (m/s)

Sensor Type:18 cm diameter 4-blade helicoid propeller molded of polypropylene

Model:RM Young 05106 Wind Monitor-MA

Range:	0 to 60 m/s (130 mph); gust survival 100 m/s (220 mph)

Accuracy: ±0.3 m/s (±0.6 mph)

Date of last calibration: 09/20/07

Parameter:Wind Direction

Units:degrees

Sensor Type: balanced vane, 38 cm turning radius

Model:	RM Young 05106 Wind Monitor-MA

Range:	360° mechanical, 355° electrical (5° open)

Accuracy:± 3%

Date of last calibration: 09/20/07

CDMO sensor:

Parameter: Wind Speed

RM Young Wind Anemometer (SN 66120) installed from 8/9/2007 – 11/5/2007

Parameter: Photosynthetically Active Radiation (PAR)

Units:mmoles m-2 (total flux)

Sensor Type:High stability silicon photovoltaic detector (blue enhanced)

Model #:LI190SB (LiCor Quantum Sensor)

Light spectrum wavelength:400 to 700 nm

Temperature Dependence:0.15% per °C maximum

Stability:<±2% change over 1 yr

Operating Temperature:-40°C to 65°C; Humidity: 0 to 100%

Sensitivity:typically 5 μA per 1000 μmoles s-1 m-2

Date of Last calibration: s/n Q22455 5/3/2001; s/n Q31375 installed new 12/12/2002 calibrated 8/7/2002, 10/20/03, 9/15/05, 10/01/07

s/n Q22455: 7/96, 5/3/01

s/n Q31375: Calibration Constant: 6.57; Multiplier: 1.26 (1.2599915)

s/n Q34301 CDMO sensor calibrated 12/14/2004, installed 8/9/07 – 11/05/07Cal constant 6.34, Multiplier: 1.305701

CDMO sensor:

Parameter: Photosynthetically Active Radiation

LICOR 190SB (SN Q34301) installed from 8/9/2007-11/5/2007

Parameter: Precipitation

Units:  millimeters

Sensor Type:  Tipping Bucket Rain Gauge

Model#:  RG-2000-C

Rainfall per tip:  0.01 inch

Operating Range:  0°C to ± 50°C; Humidity: 0 to 100%

Accuracy:  ±1.0% at 14”/hr

Date of Last calibration:  09/15/05

Parameter: Precipitation

Units: millimeters

Sensor Type: Tipping Bucket Rain Gage

Model # :  TE525WS (8 in. collector)

Rainfall per tip:  0.01 inch

Environmental Limits: Temp: 0ºC to +50ºC; Humidity: 0 to 100 %

Accuracy: + 1 % up to 1 in. hr-1; +0, -3% from 1 to 2 in./hr; +0, -5% from 2 to 3 in./hr

Date of Last calibration: Installed 4/30/07, 3/5/08

Storage Module

Model:				SM4M

Storage Capacity:		2 million low-resolution data values

Program Storage:		up to 8 programs with a total capacity of 128 KB

Processor:			Hitachi H8S

Operating System:		64 KB, flash memory based, user downloadable

Operating Range:		Temperature -35° to +65°C

Baud Rates:			9600, 76800

Memory Type:		user selectable for either ring style (default) or fill and drop

Power Requirements:		5 ± 0.3 VDC at 100 mA

Campbell Scientific CR10X Wiring Panel has 128K of flash memory (EEPROM), in which it stores the operating system and its program (that it uses to run the weather station).  Additionally, there is 128K of SRAM, which it uses to run the program and store its measurements and for final data storage.

The Campbell Scientific CR1000 Wiring Panel has two MB Flash EEPROM that is used to store the Operating System. Another 128 K Flash is used to store configuration settings. A minimum of 2 MB SRAM is (4 MB optional) is available for program storage (16K), operating system use, and data storage. Additional storage is available by using a compact flash card in the optional CFM100 Compact Flash Module.

CR1000 installed:  February 5, 2007

- Coded variable definitions:

Sampling Station		Sampling Site Code		Station Code

Charleston Meteorological		CM			soscmmet

11) QAQC flag definitions – This section details the automated primary and secondary QAQC flag definitions.

QAQC flags provide documentation of the data and are applied to individual data points by insertion into the parameter’s associated flag column (header preceded by an F_).   During primary automated QAQC (performed by the CDMO), -5, -4, -2 flags are applied automatically to indicate data that is above or below sensor range or missing.  All remaining data are then flagged 0, as passing initial QAQC checks.   During secondary and tertiary QAQC 1, -3, and 5 flags may be used to note data as suspect, rejected due to QAQC, or corrected.

-5	Outside High Sensor Range

-4	Outside Low Sensor Range

-3	Data Rejected due to QAQC

-2	Missing Data

-1	Optional SWMP supported parameter

0		Good Data

1		Suspect Data

2		Open – reserved for later flag

3		Open – reserved for later flag

4		Historical Data:  Pre-Auto QAQC

5		Corrected Data

12)  QAQC code definitions – This section details the secondary QAQC Code definitions used in combination with the QAQC flags above.

QAQC codes are used in conjunction with QAQC flags to provide further documentation of the data and are also applied by insertion into the associated flag column.  There are three (3) different code categories, general, sensor, and comment.  General errors document general problems with the deployment or the CR1000, sensor errors are sensor specific, and comment codes are used to further document conditions or a problem with the data.  Only one general or sensor error and one comment code can be applied to a particular data point.

General Errors

GIM	Instrument Malfunction

GIT	Instrument Recording Error, Recovered Telemetry Data

GMC	No Instrument Deployed due to Maintenance/Calibration

GMT 	Instrument Maintenance

GPF	Power Failure / Low Battery

GQR	Data Rejected Due to QA/QC Checks

GPR	Program Reload

GPD	Power Down

GSM	See Metadata

Sensor Errors

SIC		Incorrect Calibration Constant, Multiplier or Offset

SSN		Not a Number / Unknown Value

SNV		Negative Value

SOC		Out of Calibration

SSM		Sensor Malfunction

SSR		Sensor Removed

Comments

CAF	Acceptable Calibration/Accuracy Error of Sensor

CCU 	Cause Unknown

CDF	Data Appear to Fit Conditions

CRE	Significant Rain Event

CSM       See Metadata

CVT	Possible Vandalism/Tampering

13) Other remarks/notes

Data are missing due to equipment or associated specific sensors not being deployed, equipment failure, time of maintenance or calibration of equipment, or repair/replacement of a sampling station platform.  Any NANs in the dataset stand for “not a number” and are the result of low power, disconnected wires, or out of range readings.  If additional information on missing data is needed, contact the Research Coordinator at the reserve submitting the data.

Small negative PAR values are within range of the sensor and are due to normal errors in the sensor and the CR1000 Datalogger. The Maximum signal noise error for the Licor sensor is +/- 2.214 mmoles/m2 over a 15 minute interval.

Elevated nighttime PAR values for 2007 ranging from 2.214 to 168.90 mmoles/m2 were rejected and flagged <-3> (CSM); these elevated readings/offsets occurring at night are most likely due to slow moisture intrusion into the body of the sensor over time.  Night time values were determined using sunrise/set table for one year from US naval observatory (http://aa.usno.navy.mil/data/docs/RS_OneYear.php), and any data that fell between the set and rise, excluding the time of set and rise, or one minute after set to one minute before rise, were flagged -3 CSM.  The sensor was calibrated 10/01/07 with no problems detected.   There were 4 values of 2.20 in January that were flagged (1/9/07 17:00, 1/10/07 17:15, 1/12 7:45, 1/22 17:30) since the CR10X only recorded values of 2.2 or 2.3.

Relative humidity values 101-103 are within the range of the sensor and are flagged as <1> [CAF].  Relative humidity values > 103 are rejected.

Relative humidity data from 1/1/2007 00:00 - 3/23/2007 10:00 and 11/12/2007 04:45-06:00 were rejected <-3> [SSM] (CCU); there were highly variable fluctuations due to unknown reasons.  RH data for the remainder of the year were flagged as suspect <-1> CSM, unless the data were rejected or missing.  The following are calibration notes from the 8/28/2007 calibration:

8/28/07 Met conformance/calibration standards by Campbell Scientific.

RH Test Points (90%: initial 87.1 %/calibrated 89.4%; 50%: initial 49.0%/calibrated 51.1; 20%: initial 18.2%, calibrated 19.1%); Tempe was within +/ 0.3 degree C of a reference.

11/18/2009 Tested & inspected. RH was non-functional.  Tried replacing chip, would not respond. Hydrid circuit card is damaged, can no longer replace this circuit.  Sensor is non-repairable.

The sensor was calibrated 8/28/07 with no problems detected.

On 2/2/2007 15:45 data were missing due to an error during data retrieval from the CR10X, and this line of data was flagged <-2> [GIM] (CSM).

On 2/5/2007 13:30 Air temperature was rejected (NAN value) <-3> [GMC] (CSM) due to rewiring of this instrument after calibration.  As a result, RH data was also rejected with the same code.

Most data (exception see comment just below) are missing for 2/5/07 13:45 – 2/8/07 13:00 and flagged <-2> [GMT] (CSM) due to a datalogger replacement of the CR10x with the CR1000 on 2/5/07 13:30-15:30. All parameters for 2/8/2007 13:15 were rejected since data weren’t a full 15 minutes.

On 2/8/2007 00:00 data are rejected (or missing) and flagged as <-3> or <-2> [GQR] (CSM because this was the only line of data in a string of missing data (see above comment).

2/8/2007 13:15 – 2/9/2007 8:30 ATemp, RH, BP, Tot Prcp, and TotPAR data for these dates and times are flagged <-3> [GIT] (CSM) or <-2> [GMT] (CSM) due to raw data being lost because of failure to retrieve data before a program upload; data were recovered from telemetry but data from this period were rejected due to inconsistent values.  Raw data (lost to the same program upload) from 2/9/207 8:45-3/5/2007 10:30 were recovered from telemetry and accepted.  Most of the data for February 2007 and beginning of March were reconstructed from the CDMO telemetered file; these data were from 2/9/2007 8:45 – 3/5/2007 10:30 and are flagged <0> [GIT] (CSM) with the exception of RH <-3> [GIT] (CSM), some missing records <-2> [GIM] (CSM), Wind direction <-3> [SSM] (CSM), and Tot Precipitation (<-3> [SIC] (CSM).   See below for details on RH and missing records.  Wind direction and TotPrcp are explained in the next paragraph.

Relative Humidity: RH data were recovered from telemetry; however, data were rejected because there were highly variable fluctuations due to unknown reasons (continued until 3/23/2007 at 10:00 and 11/12/2007 04:45-06:00).   Unless rejected or missing, all other RH data were flagged as suspect for the year <-1> CSM.   See above for RH sensor calibration notes.  The sensor was calibrated 8/28/07 with no problems detected.  Air temperature data recovered from telemetry for this same period were accepted since they were compared to North Bend Airport Temp data and data were similar.

Missing records: Data from 2/17/07 4:45-6:30 were missing due to unknown reasons and are coded <-2> [GIM] (CSM).  Data from 2/17/07 6:45 were rejected as they may not be a full 15 minutes.  Data were missing 3/5/2007 00:00 due to a gap in the telemetered data <-2> [GIM] (CSM) and data were rejected on 3/5/2007 00:15 as they may not be a full 15 minutes.

From 2/8/07 13:15 Wind speed (until 2/9/2007 08:30) and direction data (until 8/8/07 13:15) were rejected <-3> [SSM] (CSM) or missing <-2> [SSM] (CSM) due to incorrect wiring; the 5106 was wired as a 5103.  Wind direction was incorrectly wired from 2/8/2007 13:15-8/8/07 13:15 <-3> [SSM] (CSM).

From 2/9/07 8:15 until 3/5/2007 10:30, precipitation data were rejected and flagged <-3> [SIC] (CSM) after attempting to correct them for the incorrect multiplier.  An incorrect multiplier for the rain gauge was used.  The old rain sensor RG-2000-C was still in use (multiplier 0.1) at this time; however, the program set up for the new rain gauge TE525WS with a multiplier of 0.254 was uploaded before the new rain gauge was installed.  Data from this period were corrected by multiplying by (0.1/0.254) or 0.394.  However, comparison of daily sums of rain for the corrected data, uncorrected data, and NB airport data showed discrepancies.  With the combination of the loss of raw data, and thus data retrieval from telemetry, and the correction due to the wrong multiplier, the data ultimately were rejected.

From 2/10/2007 until 3/4/2007, midnight readings (00:00) were rejected.  These data do not appear to be correct.  The error may have occurred with the old CR10X and the arrays that were being recorded at the time.

Data from 3/5/07 10:45-15:15 are missing due to failure to retrieve data before a program upload; these data were not recovered through telemetry.  Data on 3/5/2007 15:30 were rejected as they may not be a full 15 minutes.

Beginning 3/5/07 20:00 – 12/24/07 18:15, there are repeating low values of PAR (0.1, 0.6- 0.9, 1.5, 1.6) that were not considered elevated nighttime PAR (below 2.214); these data were flagged <1> (CSM). These repeated low PAR values occurred with both the SOS sensor s/n Q31375 as well as the CDMO sensor s/n Q34301.

Total Precipitation data from 3/29/2007 10:00 to 4/30/07 15:15 were rejected <-3> [SOC] due to a corroded/broken adjustable calibration stop preventing the static calibration from being completed.  The rain gauge was replaced on 4/30/2007 ~ 15:30.

PAR data from 5/4/2007 14:00 – 6/5/2007 9:15 were rejected <-3> [SIC] (CSM). The PAR multiplier wasn’t inserted after program upload on 5/4/07.

6/5/2007 9:15, data were collected prior to station maintenance with CDMO. All data were missing 6/5/2007 9:30-10:30 and 6/5/2007 13:45 – 6/6/07 11:15 due to station maintenance.  Data from 6/5/2007 10:45-13:30 were retrieved through telemetry file. PAR data began recording and were recovered from telemetry 6/5/07 10:45 – 13:30 but the values were low therefore data were rejected.  The PAR multiplier was inserted after the program upload 6/6/07 11:30.

New program uploads occurred on the following dates and time and therefore, data are rejected <-3> [GPR]:  (3/5/2007 15:30 SOUTHSLOUGH_v3.0_022207.CR1), 5/4/2007 13:45 (SOUTHSLOUGH_V3.0_050207.CR1), 6/6/07 11:30 (SOUTHSLOUGH_V4.0_060107.CR1), 8/9/07 10:15 (SOUTHSLOUGH_V4.1_070907.CR1), 11/7/07 10:45 (SOUTHSLOUGH_V4.1_110507.CR1).

All records for 8/8/07 14:15 are missing due to the removal of the barometric pressure and Air/Temperature Relative Humidity sensors for recalibration and data are flagged <-2> [GMT] (CSM). South Slough BP and Air Temp/RH sensors were replaced with CDMO sensors; however, the met program changes for the CDMO sensors were not uploaded until 8/9/07 10:15 therefore, all data records from 8/8/07 13:30- 8/9/07 10:00 are rejected (or missing 8/8/07 14:15) & flagged with instrument maintenance <-3> [GMT] (CSM) (or <-2> [GMT] (CSM)). The rejected data values are from raw values that were set up to collect automatically on a lab workstation, instead of climbing the tower and retrieving the files with a laptop computer.

SOS wind and par sensors were replaced with CDMO sensors and the program (SOUTHSLOUGH_V4.1_070907.CR1) was uploaded 8/9/07 10:15; data records were rejected 8/9/2007 9:45-10:00 <-3> [GMT] (CSM).   PAR Data were rejected 8/9/07 10:30-8/10/07 14:00 due to low PAR values (0.00) during the day; the incorrect multiplier was used 1.26 instead of 1.31; the correct multiplier for the CDMO PAR sensor was changed 8/10/07 ~ 13:50.

There was a noticeable shift in the PAR values following the 8/09/2007 PAR sensor switch.  This is probably due to using a newly installed CDMO PAR sensor followed by the freshly calibrated SOS PAR sensor.

All data collected with the CDMO sensors were flagged <1> [SOC] ( CSM) due to no accurate calibration history available for these sensors during the following period (exceptions for data during this period that were already flagged rejected or suspect for another reason):

Air Temp and Relative Humidity: 8/9/2007 10:30 – 11/1/07 14:00

Barometric Pressure: 8/9/07 10:30 – 11/5/07 12:00

Wind Speed/Direction: 8/9/10:30 - 11/5/07 12:15

PAR: 8/10/07 14:15 - 11/5/07 12:15

10/8/2007 22:15 – 10/9/2007 00:30 data were missing probably due to the real time file not getting transmitted to the lab computer via the lan link and were flagged <-2> [GIM] (CSM).

11/1/2007 14:00-14:15 BP data were rejected <-3> [GQR] (CSM) due to low numbers; this problem occurred just before other maintenance and rewiring issues for other sensors and the cause was unknown.

Temperature and RH data records from 11/1/07 14:15-11/6/07 11:00 were rejected <-3> (CSM) due to incorrect wiring (values -40 C).  This occurred when the CDMO sensor was replaced with the SOS calibrated sensor; the sensor was rewired and the problem was fixed 11/6/07 11:00.

11/5/2007 12:15-13:00 The BP CDMO sensor was replaced with the calibrated SOS sensor.

Wind speed and direction data records were rejected or missing from 11/5/07 12:30 – 11/7/07 11:45; on 11/5/07 the CDMO wind sensor was removed and the calibrated SOS wind sensor was installed; however, the wiring color code for the RM Young 5106 instrument was different than the RM Young 5106 manual (Oct 2001) and the CDMO manual; this difference was discovered during annual recalibration of sensors with Ed Chemosky from RM Young. The white wire (wind direction excitation voltage) should be attached to an excitation terminal of the datalogger and the black remains on ground.  The brown wire is connected to ground and the red is connected to the Pulse 1 channel (P1).  The SOS calibrated wind sensor was properly wired on 11//7/07 11:00-11:45 pm PST.

Data were being collected and sent automatically to a real-time file in the lab (in place of climbing the tower and downloading the data with a laptop computer), but the computer was not always left on, which resulted in gaps; therefore, data from 11/6/07 11:30-13:30, 11/6/07 14:30-11/7/07 10:30 were present in the file; however, some data were missing from the file for 11/6/07 11:15, 13:45-14:15.  ATemp and RH were rejected 11/06/07 at 12:30 due to the low ATemp reading (0.0).  PAR data were rejected from 12:30-13:00.  All data were rejected 11/6/07 11:30, 14:30 because data weren’t a full 15 minutes due to the preceding data value missing.

11/7/2007 10:45 data were rejected <-3> [GPR] (CSM) because the wind sensor was rewired during this time and this does not reflect a full 15 minutes and there was a program reload (SOUTHSLOUGH_V4.1_110507.CR1).

There are intermittent missing data records (-2 GPF CSM) due to low battery/ power failure

on 12/19/07 20:15 – 20:30, 12/19/07 21:15-12/20/07 10:00, 12/20/07 23:30-12/21/07 10:00, 12/21/07 19:00 – 12/24/07 10:00, 10:30, 12/24/18:30 -12/26/07 10:00, 10:30, 12:15, 12/26/07 18:00-12/30/07 9:45, 12/30/07 10:45 –11:45, 12:15 – 14:30, 12/30/07 16:00-12/31/07 9:15.  The data records collected following the missing records were rejected because they were not a full 15 minutes of data.

Precipitation data recorded at the North Bend Airport compared to soscmmet weather station:

2007 Monthly Precipitation in mm

| Site | Jan | Feb | Mar | Apr | May | Jun | Jul | Aug | Sep | Oct | Nov | Dec |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| soscmmet | 132.00 | 232.00 | 143.37 | 78.23 | 39.21 | 22.72 | 31.50 | 74.72 | 69.60 | 143.51 | 177.66 | 89.9 |
| NBWS | 155.70 | 237.49 | 125.22 | 81.03 | 35.56 | 31.50 | 24.64 | 44.45 | 63.25 | 127.25 | 179.83 | 277.37 |
| Histor | 258.32 | 202.69 | 194.56 | 114.30 | 76.71 | 41.40 | 11.43 | 18.03 | 42.93 | 131.32 | 241.05 | 270.26 |

soscmmet		2007 data from SWMP Charleston Meteorological Station – located 11.5 km

SW of NBWS and 8 km SW of local weather spotter

NBWS		North Bend Weather Station at FAA Airport 2005 data

Histor		Mean data recorded at NBWS from 1931-2001