South Slough (SOS) NERR Meteorological Metadata

January - December 2014

Latest Update: 09/07/2016

I.  Data Set and Research Descriptors

- Principal investigators and contact persons:

### Address: 	South Slough National Estuarine Research Reserve

### PO Box 5417

### Charleston, OR 97420

### Phone:  (541) 888-5558

### Fax:      (541) 888-2733

Contact Persons:

Bree Yednock, Lead Scientist/Research Coordinator

E-mail:  bree.yednock@state.or.us; (541) 888-8270 ext. 302

Alicia Helms, Estuarine Monitoring Coordinator

E-mail: alicia.r.helms@state.or.us; (541) 888-8270 ext. 304

Adam DeMarzo, Estuarine Monitoring Assistant

E-mail: adam.demarzo@state.or.us; (541) 888-8270 ext. 305

- Entry verification

Data are uploaded from the CR1000 data logger to a Personal Computer (IBM compatible).  Files are exported from LoggerNet in a comma-delimited format and uploaded to the CDMO where they undergo automated primary QAQC and become part of the CDMO’s online provisional database.  During primary QAQC, data are flagged if they are missing or out of sensor range.  The edited file is then returned to the Reserve where it is opened in Microsoft Excel and processed using the CDMO’s NERRQAQC Excel macro.  The macro inserts station codes, creates metadata worksheets for flagged data and summary statistics, and graphs the data for review.  It allows the user to apply QAQC flags and codes to the data, append files, and export the resulting data file to the CDMO for tertiary QAQC and assimilation into the CDMO’s authoritative online database.  For more information on QAQC flags and QAQC codes, see Sections 11 and 12.  Ali Helms was responsible for data management 2006-2011, 2016.  Adam DeMarzo was responsible for data management 2012-2015.

- Research objectives:

The South Slough NERR operates the SWMP Meteorological Station to provide continuous measurements of local weather conditions affecting the South Slough estuary.  The nearest National Weather Service certified station is located at the North Bend Airport, 11.5 km NE from the mouth of South Slough.  At that location, the weather observers’ primary function is to provide current meteorological information to airplane pilots.  Only 24-hour weather data is recorded, including maximum and minimum temperature, precipitation, and prevailing wind direction.

Because the SWMP weather station is located at the mouth of South Slough, recorded data provides information more pertinent to the microclimates affecting the estuary.  The principal objective is to record long-term meteorological data that is particularly relevant to the South Slough estuary in order to gain a better understanding of climatic forcing and to observe any environmental changes or trends over time.  In addition, the station provides real-time data to assess the short-term effects of local weather on measurements of water parameters within the estuary.  The station is part of a framework to evaluate the relative contributions of oceanic forcing and watershed inputs to hydrodynamics, nutrient dynamics, and other ecological processes within the South Slough estuary.  Data can be used as 1) a vital reference of atmospheric data for various research projects at the Reserve – an integral part of the NERR mission to provide a platform for estuarine research, 2) for meteorological context or atmospheric forcing for 15 minute water quality data, and other long-term environmental monitoring paradigms at the Reserve.

- Research methods:

Campbell Scientific data telemetry equipment was installed at the Charleston Meteorological (soscmmet) station on 11/17/05 and transmits data to the NOAA GOES satellite, NESDIS ID ##3B0032E0.  The equipment includes a GOES satellite antenna, a Global Positioning System (GPS) receiver, a 20 watt solar panel and charge controller, a 12 V DC power supply, and a weatherproof enclosure. The transmissions are scheduled hourly and contain four data sets reflecting fifteen minute data sampling intervals.  Upon receipt by the CDMO, the data undergoes the same automated primary QAQC process detailed in Section 2 above.  The “real-time” telemetry data become part of the provisional dataset until undergoing secondary and tertiary QAQC and assimilation in the CDMO’s authoritative online database.  Provisional and authoritative data are available at http://cdmo.baruch.sc.edu.

The 15 minute Data are collected in the following formats for the CR1000:

Averages from 5-second data:

Air Temperature (°C), Relative Humidity (%), Barometric Pressure (mb), Wind Speed (m/s), Wind Direction (degrees), Battery Voltage (volts)

Maximum and Minimum Air Temperature (°C) and their times from 5-second data (these data are available from the Reserve):

Maximum Wind Speed, (m/s) from 5-second data

Wind Direction Standard Deviation (degrees)

Totals: Precipitation (mm), PAR (millimoles/m2), and Cumulative Precipitation (mm)

The meteorological sensors are inspected monthly for damage or fouling and the solar panels, PAR sensor, and the rain gauge are cleaned if necessary.  A Kestrel 4000 handheld is used as a general check of the sensors. In the laboratory, data recorded as close as possible to the field check are downloaded from the nearby North Bend Weather Station and compared to station data.  Beginning 9/30/2008, the data check for the SWMP met station incorporated the nearby Charleston NWLON (National Water Level Observation Network) station ID 9432780 as it is physically closer to the SWMP Meteorological Station (approximately 500 m ); this NOAA station collects Meteorological Observations of Wind Speed (m/s), Wind Direction (Degrees), , and Barometric Pressure (mbars) at 6 minute intervals; therefore, the Relative Humidity sensor is compared with the North Bend Airport or the Kestrel.  Most of the sensors are mailed for recalibration yearly or every two years with the exception of the rain gauge which is calibrated by a static method annually in house.

Recommended calibration frequency for the MET station sensors:

- Temperature/Humidity- yearly recalibration

- Rain Gauge- yearly recalibration

- Wind Speed/Direction- yearly or every 2 years (depending on the sensor)

- Barometric Pressure- every 2 years recalibration

- PAR- every 2 years recalibration

- CR1000-every 5 years (required beginning 2014, one year initial grace period)

- Site location and character:

South Slough is characterized as a drowned river-mouth estuary.  Located 8.5 km southwest of Coos Bay, Oregon, South Slough is one of seven tidal inlets that collectively form the Coos estuary.  The South Slough National Estuarine Research Reserve is representative of the mid-Pacific portion of the Lower Columbian Bioregion, extending from Cape Mendocino, CA north to the Columbia River, OR.  The Reserve includes 4,700 acres with 600 acres of tidal marshes, mudflats and open water channels.

Within the South Slough, there are three primary sub-systems that provide much of the freshwater input to the area.  Only two sub-systems, the Sengstacken and Winchester arms, are within Reserve boundaries.  The Sengstacken arm drains a semi-pristine area with no development and serves as a reference site for the riverine portion of the estuary.  The Winchester arm is similarly located in an area with development limited to one small farm (<10 cattle) high in the watershed.  The area functions as a management-treatment site for salt marsh restoration activities in the Winchester Tidelands Restoration Project.  Valino Island marks the northern boundary of the Reserve and is a marine-dominated area.

Charleston Meteorological Station:  43.34505 N, -124.328733 W The station is located 1.2 km northwest of the Charleston water quality monitoring site.  A 10-meter tall navigation tower that formerly belonged to the Corps of Engineers is situated on the northeast corner of the Oregon Institute of Marine Biology campus.  The tower is about 20 meters from the inner boat marina at Charleston harbor.  Three 12-13 meter tall shore pines lie directly south of the tower and may slightly shelter the station from southerly winds, but the sensor array is never shaded.  The weather station is above the nearest building, which is about 8 meters away.  A 40-meter-tall forested ridge that has a northeast to southwest alignment is located about 60 meters from the weather tower and may interfere with winds from the southwest.  On the tower platform, a 3-meter tall mast holds the Wind Sentry, Temperature and Humidity, Barometric, and PAR sensors array.  Because there is no open field available, the Rain Gauge is set on top of a platform on the southwest corner of the tower.  Sensor heights from the ground are Temperature/Relative Humidity 10.1 m, Barometric Pressure 9.4 m, Wind 11.4 m, PAR 11.2 m, and Rain 9.0 m.   Other sensor placement is in accordance with the CDMO Manual (Version 4.0, Part D, Section 2.0).

The CR1000 is powered by an internal battery.  Two solar power panels (40W and 20W) generate power during the daylight hours to maintain battery charge. The panels also generate power to the Optimus 12V deep-cycle sealed gel battery (installed on 11/17/05) during the day. The Optimus provides power during the night.

Safety apparatus for ascending the tower is in accordance with OSHA regulations.  A sliding trolley runs in a track during ascension of the ladder.  The trolley is connected by a carabiner to a chest harness and belt.  While working on the tower platform, a 4-meter-long tether secures the person to a tower strut.

- Data collection period:

The Charleston station was installed on 08/21/01.  The station was hardwired into the Estuarine and Coastal Science laboratory IBM-compatible PC originally using CSI’s Real-Time Data Management Program on December 13, 2002, and the data were downloaded automatically to a PC in the laboratory via a local network link (Campbell Science NL100).  A 4MB storage module was installed on 10/20/03 to increase the capacity of the CR10X data storage.  Instruments were installed on 11/17/05 to telemeter the meteorological data to NOAA GOES satellite.  On Oct 6, 2006, the hardwire connection was removed due to conflict with server usage for a new employee.  On February 5, 2007, a CR1000 datalogger was installed to replace the CR10x.  The CR1000 interfaces with the LoggerNet 3.2.2 software supplied by Campbell Scientific.  In June 2007, the NL100 Lan Link network connection was re-established using LoggerNet 3.2; this allowed the weather data to be automatically downloaded to a real-time file on a personal computer; however, there were problems with the hardware and IP address and the NL100 connection was not maintained.  Beginning Jan 2008, a laptop computer is connected directly to the CR1000 datalogger via an RS232 connection to download data approximately on a monthly basis.

Data were collected in 2014 from 01/01/2014 00:00 to 12/31/2014 23:45.  Raw data were downloaded from the station during the following periods:

12/30/2013 12:00 –  03/10/2014 09:15

03/10/2014 09:30 – 04/07/2014 12:15

04/07/2014 12:30 – 05/08/2014 11:15

05/08/2014 11:30 – 06/20/2014 14:00

06/20/2014 14:15 – 07/31/2014 15:15

07/31/2014 15:30 – 09/03/2014 15:00

09/03/2014 15:15 – 10/13/2014 14:45

10/13/2014 15:00 – 12/18/2014 13:00

12/18/2014 13:15 – 02/02/2015 13:45

- Distribution

NOAA retains the right to analyze, synthesize and publish summaries of the NERRS System-wide Monitoring Program data.  The NERRS retains the right to be fully credited for having collected and process the data.  Following academic courtesy standards, the NERR site where the data were collected should be contacted and fully acknowledged in any subsequent publications in which any part of the data are used.  The data set enclosed within this package/transmission is only as good as the quality assurance and quality control procedures outlined by the enclosed metadata reporting statement.  The user bears all responsibility for its subsequent use/misuse in any further analyses or comparisons.  The Federal government does not assume liability to the Recipient or third persons, nor will the Federal government reimburse or indemnify the Recipient for its liability due to any losses resulting in any way from the use of this data.

Requested citation format:

NOAA National Estuarine Research Reserve System (NERRS). System-wide Monitoring Program. Data accessed from the NOAA NERRS Centralized Data Management Office website: http://www.nerrsdata.org/; accessed 12 October 2012.

NERR meteorological data and metadata can be obtained from the Research Coordinator at the individual NERR site (please see Principal Investigators and Contact Persons), from the Data Manager at the Centralized Data Management Office (please see personnel directory under the general information link on the CDMO home page) and online at the CDMO home page www.nerrsdata.org.  Data are available in comma delimited format.

- Associated researchers and projects:

The South Slough NERR maintains four water quality monitoring stations along the estuarine gradient within the South Slough estuary.  These stations collect continuous information to address three primary topics:  1) short-term variability and long-term changes in estuarine water parameters within different regions of the estuary (i.e., localized impacts of seasonal storm events, variability in water quality parameters in association with re-establishment of tidal circulation, interannual differences in rainfall, magnitude and influence of El Niño – La Niña events, spatial extent of oceanic and tidal forcing, and long-term changes associated with the Pacific Decadal Oscillation, etc.); 2) verification, calibration, and future refinement of an existing two-dimensional water quality model (CE-QUAL-W2) developed to predict levels of coliform bacteria and dispersion of environmental toxins; and 3) collection of fundamental baseline data for future development of a three-dimensional tidal hydrodynamic model for the greater South Slough/Coos Bay estuaries.

A water chemistry component is also associated with the core monitoring program.  Variables measured include: nitrogen (nitrate, nitrite, and ammonium), orthophosphate, chlorophyll a, phaeophytin and silica.  Grab samples are collected for both high and low tides at four water quality sampling sites, and automated samplers collect diel samples to compare marine and riverine sites during wet and dry seasons.

The SSNERR water quality monitoring project has been incorporated into the State of Oregon Coastal Non-Point Pollution Control Program (S. 6217) developed by the Oregon Department of Land Conservation and Development and the Oregon Department of Environmental Quality in July 1995.

#### II. Physical Structure Descriptors

- Sensor specifications, operating range, accuracy, date of last calibration, dates of sensor use

Parameter:  Temperature

Units:	Celsius

Sensor Type:  Platinum resistance temperature detector (PRT)

Model#:  HMP45C and HMP35C Temperature and Relative Humidity Probes

The specifications for the HMP45C and the HMP35C sensors are as follows (https://s.campbellsci.com/documents/af/manuals/hmp45c.pdf; https://s.campbellsci.com/documents/br/manuals/hmp35c.pdf ):

Operating Temperature:	-40°C to +60°C (HMP45C); -35°C to +60°C (HMP35C)

Range: -40°C to +60°C (HMP45C); -35°C to +55°C (HMP35C)

Accuracy: ± 0.2 °C @ 20°C (HMP45C); ± 0.4 °C over the range - 24°C to 48°C (HMP35C)

Serial numbers: s/n D3520106 (HMP45C), s/n R2720011 (HMP35C), s/n 467113 (HMP35C)

Dates of Last calibration:  10/01/2015, 08/26/2008 (s/n D3520105); 07/09/2012, 04/26/2007 (s/n R2720011); 11/18/2009, 08/28/2007 (s/n 467113)

Dates of Sensor Use: D3520105 06/22/2012-05/04/2015; R2720011: 01/11/2010-05/11/2012; 467113: 11/01/2007-11/12/2009

Parameter:  Relative Humidity

Units:	Percent

Sensor Type:  Vaisala HUMICAP© 180 capacitive relative humidity sensor

The specifications for the HMP45C and the HMP35C sensors are as follows (https://s.campbellsci.com/documents/af/manuals/hmp45c.pdf; https://s.campbellsci.com/documents/br/manuals/hmp35c.pdf ):

Model#:  HMP45C and HMP 35C Temperature and Relative Humidity Probe

Range:	 0 - 100% non-condensing

Accuracy at 20 oC:  ± 2% RH (0-90%) and ± 3% (90-100%) (same for HMP45C and HMP35C)

Temperature dependence of RH measurement: ± 0.05% RH/°C (HMP45C); ± 0.04% RH/°C (HMP35C)

Serial Numbers: s/n D3520106 (HMP45C), s/n R2720011 (HMP35C), s/n 467113 (HMP35C)

Date of Last calibration:  10/01/2015, 08/26/2008 (s/n D35201015); 07/09/2012, 04/26/2007 (s/n R2720011); 11/18/2009, 08/28/2007 (s/n 467113)

Dates of Sensor Use: D3520105 06/22/2012-05/04/2015; R2720011: 01/11/2010-05/11/2012; 467113: 11/01/2007-11/12/2009

Parameter:  Barometric Pressure

Units:  millibars (mb)

Sensor Type:  Vaisala Barocap © silicon capacitive pressure sensor

Model #:  CS105, CS106

Operating Range:  Pressure: 600-1060 mb (CS105), 500 mb-1100mb (CS106); Temperature: -40°C to +60°C (CS105, CS106)

Humidity:  non-condensing

Accuracy: ±0.5 mb at 20°C,  ± 2 mb at  0°C-40°C, ± 4 mb at 20°C to 45°C, ± 6 mb -40 –60°C (CS105); ±0.3 mb at 20°C, ± 0.6 mb at  0°C-40°C, ± 1 mb at -20°C to 45°C, ± 1.5 mb -40 –60°C (CS106)

Stability: ± 0.1 mb per year

Serial Numbers: s/n X3540012 (CS105); s/n H0160056 (CS106)

Date of Last calibration: 10/15/2015, 12/02/2009 (X3540012); 01/10/2012, 07/09/2012 (H0160056)

Dates of Sensor Use: 11/05/2007-11/12/2009, 01/11/2010 - 5/11/2012, 06/15/2012-05/04/2015 (X3540012); 05/11/2012-06/15/2012 (H0160056)

Parameter:  Wind Speed

Units:	meters per second (m/s)

Sensor Type:  18 cm diameter 4-bladed helicoids propeller molded of polypropylene

Model:  RM Young 05106-5A Wind Monitor

Operating Temperature: -50°C to 50°C

Range:	 0-100 m/s (224 mph)

Accuracy:  ± 0.3 m/s or 1 % of reading

Serial Number: WM115774

Date of last calibration: 10/15/2015, 11/22/2011

Dates of Sensor Use: 06/11/2012-05/04/2015

Parameter: Wind Direction

Units: degrees

Sensor Type:  balanced vane, 38 cm turning radius

Model:	 RM Young 05106 Wind Monitor

Range:	  360° mechanical, 355° electrical (5° open)

Accuracy:  ± 3 degrees

Serial Number: WM115774

Date of last calibration:  10/15/2015, 11/22/2011

Dates of Sensor Use:  06/11/2012-05/04/2015

Parameter:  Photosynthetically Active Radiation

LI-COR Quantum Sensor

Units:  mmoles m-2 (total flux) conversion, sensor measures micromoles per second per square meter (umol s-1 m-2)

Sensor Type:  High stability silicon photovoltaic detector (blue enhanced)

Model:	  LI 190SZ

Light spectrum wavelength:	400 to 700 nm

Temperature Dependence:  0.15% per °C maximum

Stability:  < ± 2% change over 1 yr

Operating Temperature:  -40°C to 65°C; Humidity: 0 to 100%

Sensitivity:  typically 5 μA per 1000 μmoles s-1 m-2

Multiplier: s/n Q46942 1.196264

Serial Number: s/n Q46942

Date of Last calibration: 01/20/2012

Dates of Sensor Use: 06/11/2012-05/04/2015

Parameter:  Precipitation

Units:  milliliters (mm)

Sensor Type:  Tipping Bucket Rain Gauge

Model:	  TE525USW

Range:	  0.01 in or 0.2 mm per bucket tip

Operating Range:  Temperature: 0° to ± 70°C; Humidity: 0 to 100%

Accuracy:  ±1.0% at 2”/hr or less

Serial Number: s/n 41930-307

Date of Last calibration:  10/15/2015, 4/21/2010

Dates of Sensor Use: 04/30/2007-05/04/2015

The CR1000 has 2 MB of Flash EEPROM that is used to store the Operating System.  Another 128 K Flash is used to store configuration settings.  A minimum of 2 MB SRAM is (4 MB optional upgrade) available for program storage (16K), operating system use, and data storage.  Additional storage is available by using a compact flash card in the optional CFM100 Compact Flash Module.

Date CR1000 Installed: 2/5/2007

Date CR1000 Calibrated: 05/01/2006; 10/30/2015

CR1000 Firmware Version (s):  CR1000.Std.21

CR1000 Program Version(s): SOSCMMET_5.7_062212

- Coded variable definitions:

Sampling Station		Sampling Site Code		Station Code

Charleston Meteorological		CM			soscmmet

11) QAQC flag definitions – This section details the automated primary and secondary QAQC flag definitions.

QAQC flags provide documentation of the data and are applied to individual data points by insertion into the parameter’s associated flag column (header preceded by an F_).   During primary automated QAQC (performed by the CDMO), -5, -4, and -2 flags are applied automatically to indicate data that is above or below sensor range, or missing.  All remaining data are then flagged 0, as passing initial QAQC checks.   During secondary and tertiary QAQC 1, -3, and 5 flags may be used to note data as suspect, rejected due to QAQC, or corrected.

-5	Outside High Sensor Range

-4	Outside Low Sensor Range

-3	Data Rejected due to QAQC

-2	Missing Data

-1	Optional SWMP supported Parameter

0		Passed Initial QAQC Checks

1		Suspect Data

2		Open - reserved for later flag

3		Open - reserved for later flag

4		Historical Data:  Pre-Auto QAQC

5		Corrected Data

12)  QAQC code definitions – This section details the secondary QAQC Code definitions used in combination with the QAQC flags above.

QAQC codes are used in conjunction with QAQC flags to provide further documentation of the data and are also applied by insertion into the associated flag column.  There are three (3) different code categories, general, sensor, and comment.  General errors document general problems with the CR1000, sensor errors are sensor specific, and comment codes are used to further document conditions or a problem with the data.  Only one general or sensor error and one comment code can be applied to a particular data point, but some comment codes (marked with an * below)  can be applied to the entire record in the F_Record column.

General Errors

GIM	Instrument Malfunction

GIT	Instrument Recording Error, Recovered Telemetry Data

GMC	No Instrument Deployed due to Maintenance/Calibration

GMT 	Instrument Maintenance

GPD	Power Down

GPF	Power Failure / Low Battery

GPR	Program Reload

GQR	Data Rejected Due to QA/QC Checks

GSM	See Metadata

Sensor Errors

SDG	Suspect due to sensor diagnostics

SIC	Incorrect Calibration Constant, Multiplier or Offset

SIW	Incorrect Wiring

SMT	Sensor Maintenance

SNV	Negative Value

SOC	Out of Calibration

SQR	Data rejected due to QAQC checks

SSD	Sensor Drift

SSN	Not a Number / Unknown Value

SSM	Sensor Malfunction

SSR	Sensor Removed

Comments

CAF	Acceptable Calibration/Accuracy Error of Sensor

CCU	Cause Unknown

CDF	Data Appear to Fit Conditions

CML	Snow melt from previous snowfall event

CRE*	Significant Rain Event

CSM*	See Metadata

CVT*	Possible Vandalism/Tampering

CWE*	Significant weather even

13) Other remarks/notes

Data are missing due to equipment or associated specific sensors not being deployed, equipment failure, time of maintenance or calibration of equipment, or repair / replacement of a sampling station platform.  Any NANs in the dataset stand for “not a number” and are the result of low power, disconnected wires, or out of range readings. If additional information on missing data is needed, contact the Research Coordinator at the reserve submitting the data.

Small negative PAR values are within range of the sensor and are due to normal errors in the sensor and the CR1000 Datalogger. The Maximum signal noise error for the Licor sensor is +/- 2.214 mmoles/m2 over a 15 minute interval.

Relative Humidity data greater than 100 are within range of the sensor accuracy of +/-3%.

Data recorded for all parameters (with the exception of cumulative precipitation) at the midnight timestamp (00:00) are the 15 minute averages and totals for the 23:45-23:59 time period of the previous day.  Cumulative precipitation data at the midnight timestamp (00:00) are the sum of raw (unrounded) precipitation data from 00:00 to 23:59 of the previous day.  Summing each individual 15-minute total precipitation value from the same period will result in small differences from cumulative precipitation due to rounding.  It is especially important to note how data at the midnight timestamp are recorded when using January 1st and December 31st data.  Note:  Cumulative precipitation is no longer available via export from the CDMO.  Please contact the Reserve or the CDMO for more information or to obtain these data.

The Air Temperature/Relative Humidity sensor (s/n D35201016) was considered out of calibration for part of 2014 (calibrated 08/26/2008, 10/02/2015) and air temperature and relative humidity data records from 06/22/2014 00:15 - 12/31/2014 23:45 were flagged 1 SOC CSM.  The sensor was calibrated 10/02/15 and was documented as out of tolerance.   This sensor was from the telemetry repository and was meant to be installed temporarily during 2012 sensor calibrations until the South Slough sensor s/n R2720011 was recalibrated (calibrated 07/09/2012); however, the repository sensor was never replaced.

The Rain gauge (s/n 41930-307) was out of calibration (calibrated 04/21/2010, 10/15/2015) and Total Precipitation and Cumulative Precipitation records are flagged 1 SOC CSM for all of 2014.  Calibration results from 10/15/2015 indicated the sensor did not pass calibration tests.  The sensor averaged 70 tips per 653 mL reservoir (+/ 1 %), which was outside the acceptable range of 78.4-80.0 tips per 653 mL reservoir.

Slightly elevated nighttime PAR values were observed throughout the year and were flagged and coded as suspect, 1 CSM using http://aa.usno.navy.mil/data/docs/RS_OneYear.php to determine sunrise and sunset times.  Elevated nighttime PAR values ranged from 0.1 - 4.0 mmoles/m2 and occurred beginning 01/04/2014 18:15 through 12/31/2014 20:45.