South Slough (sos) NERR Water Quality Metadata

January-June 2017

Latest Update: 08/01/2017

Note: This is a provisional metadata document; it has not been authenticated as of its download date.  Contents of this document are subject to change throughout the QAQC process and it should not be considered a final record of data documentation until that process is complete.  Contact the CDMO (cdmosupport@belle.baruch.sc.edu) or Reserve with any additional questions.

I.  Data Set and Research Descriptors

1)  Principal investigator(s) and contact persons –

Bree Yednock, Acting Manager

bree.yednock@state.or.us (541) 888-8270 ext. 302

Alicia Helms, Estuarine Monitoring Coordinator

alicia.r.helms@state.or.us (541) 888-8270 ext. 304

Adam DeMarzo, Monitoring Technician

adam.demarzo@state.or.us (541) 888-8270 ext. 310

Mailing Address: South Slough NERR

Reserve Science Center

P.O. Box 5417 / 63466 Boat Basin Rd.

Charleston, OR 97420

Phone: (541) 888-8270

2)  Entry verification –

Deployment data are uploaded from the YSI data logger to a Personal Computer (IBM compatible).  Files are exported from EcoWatch in a comma-delimited format (.CDF) , EcoWatch Lite in a comma separated file (CSV) or KOR Software in an Excel File (.XLS)  and uploaded to the CDMO where they undergo automated primary QAQC; automated depth/level corrections for changes in barometric pressure (cDepth or cLevel parameters); and become part of the CDMO’s online provisional database.  All pre- and post-deployment data are removed from the file prior to upload.  During primary QAQC, data are flagged if they are missing or out of sensor range.  The edited file is then returned to the Reserve for secondary QAQC where it is opened in Microsoft Excel and processed using the CDMO’s NERRQAQC Excel macro.  The macro inserts station codes, creates metadata worksheets for flagged data and summary statistics, and graphs the data for review.  It allows the user to apply QAQC flags and codes to the data, remove any overlapping deployment data, append files, and export the resulting data file for upload to the CDMO.  Upload after secondary QAQC results in ingestion into the database as provisional plus data, recalculation of cDepth or cLevel parameters, and finally tertiary QAQC by the CDMO and assimilation into the CDMO’s authoritative online database.  Where deployment overlap occurs between files, the data produced by the newly calibrated sonde is generally accepted as being the most accurate.  For more information on QAQC flags and codes, see Sections 11 and 12.

Outliers in the data are evaluated to determine whether they should be flagged as rejected (-3), suspect (1), or corrected (5), following the CDMO water quality QAQC Flag and Code suggestions.  Data are flagged if the values are outside the range expected for the site (see Site Location and Character, Section 5).  Data outside the acceptable range of water quality parameters for a monitoring site are investigated for validity based on weather data, field observations, post calibration notes and instrument diagnostics.  Data outliers are flagged if they are attributed to sensor malfunction or exposure of sensors during low tide if it can be determined that freshwater inflow was not responsible — identified by unusual Depth, Temperature, and Specific Conductivity data.  Data are flagged as suspect or rejected if 1) there is evidence that fouling of sensors by aquatic organisms, debris or sediment influence data values, 2) there is significant sensor drift detected from field calibration checks or post calibration procedures and/or, 3) comparisons of previous deployment data with new deployment data indicate a significant, variable, or unnatural change at the time of sonde retrieval/deployment.   Field calibration checks are made by comparing field check measurements of the retrieved sonde at all sites to either newly deployed sonde at Valino Island, Elliot Creek, Winchester Creek, or to handheld sonde at Charleston Bridge.  Large positive Turbidity spikes (>500 FNU) are flagged as suspect, and spikes larger than 1400 FNU are rejected.

Ali Helms has been responsible for data management since July 2006.  Adam DeMarzo has provided assistance with data management since Oct 2006.  Steve Rumrill oversaw SWMP program operations from 1995-Oct 2011.  Sue Powell was responsible for data management from 1999-2006.

3)  Research objectives –

Four long-term monitoring stations have been established along the salinity gradient of the South Slough estuary to collect essential baseline information and to improve our understanding of tidal dynamics and watershed inputs.  In particular, the South Slough Reserve monitoring stations collect continuous information to address two primary topics:  1) short-term variability and long-term changes in estuarine water parameters within different hydrodynamic regions of the estuary (i.e., localized impacts of seasonal storm events, variability in water quality parameters in association with re-establishment of tidal circulation, interannual differences in rainfall, magnitude and influence of El Niño – La Niña events, spatial extent of oceanic and tidal forcing, localized impacts of regional upwelling and long-term changes associated with the Pacific Decadal Oscillation) and 2) collection of baseline data for research and monitoring projects, including the development of a three-dimensional tidal hydrodynamic model for the Coos Bay estuary.

The South Slough estuary contains three primary sub-systems (Joe Ney Arm, Sengstacken Arm, Winchester Arm) that provide much of the freshwater input to the area.  Only two sub-systems, the Sengstacken and Winchester arms, are located within the administrative boundaries of the Reserve.  One long-term monitoring station was sited in the riverine portion of each arm in 1995.  The Sengstacken Arm station (Talbot Creek) was located in a semi-pristine area surrounded by undisturbed marshes and upland forests with no residential or rural development and serves as a reference site for the riverine portion of the estuary; this site was discontinued on June 16, 2012.   The Elliot Creek station is also located in the Sengstacken Arm and now serves as a primary SWMP station (replacing Talbot Creek); data collection began on July 23, 2012.  The Winchester Creek station is similarly located in an area characterized by marshes and upland forests, and where rural development is limited to one small farm (<15 cattle) high in the watershed.  This station functions as a management-treatment site for saltmarsh restoration activities encompassed by the Winchester Tidelands Restoration Project.  The Valino Island station was established in June 1999 in the middle (southern extent of the marine-dominated region) of the South Slough estuary, downstream from the two riverine stations.  Charleston Bridge, the fourth monitoring station, was installed in April 2002 close to the marine/estuarine interface.  Both the Valino and Charleston stations are located outside the South Slough NERR boundaries.

4)  Research methods –

Deployment & Retrieval Methods

YSI EXO2 water quality sondes are deployed at each permanent monitoring station in the South Slough estuary.  The sensor arrays are held 15 cm (Sengstacken) to 50 cm (Winchester, Valino, Charleston, Elliot) off the bottom of the tidal channels that are approximately 3 meters deep.  The data sondes are programmed to record water quality parameter measurements in Pacific Standard Time every 15 minutes over deployment periods of approximately one month.  Parameters measured include Water Temperature, Salinity, Specific Conductivity, pH, Dissolved Oxygen (DO), Turbidity, and Depth.

Sites are accessed at high tide, preferably above 5.0’, in a small skiff or the SSNERR KatanaCraft R/V Joy Tally, both equipped with outboard motors.  Due to limitations in the availability of data loggers during the period of January-April 07, the sondes were not interchanged among sites but were returned to the same site approximately 24 hours later after uploading data, cleaning, and calibrations were completed.  South Slough acquired enough datasondes and sensors by the April 2007 deployment to begin interchanging the sondes.  Occasionally, we revert to the old deployment methods if sondes are being repaired or under maintenance.

During transport, each sonde is wrapped with calibration cup and placed horizontally in a padded and vented cooler for insulation against jarring.  To deploy, the data loggers are attached to stainless steel chain and lowered into protective cases (constructed from 4"-diameter ABS plumbing pipe).  The ABS cases are equipped with rectangular windows (2” x 6”) to ensure adequate tidal flushing and exposure of the probes to ambient water conditions.  ABS cases are painted with two, thin coats of Petit Unepoxy Plus anti-fouling paint.  When dry, the cases are affixed vertically to existing log pilings driven into the mud or stainless steel towers.  A stainless steel bolt, located at the bottom of the pipe, prevents the data logger from descending beyond the ABS pipe and ensures that the logger monitors at the same depth on every deployment.  YSI sensors are wrapped in copper tape and sensor guards are painted with anti-fouling paint according to procedures outlined in the YSI manual.  In the past, the sensor guard was wrapped with a piece of nylon cloth mesh (20 x 30 cm, 6 mm opening) secured with duct tape, and the bottom of the sonde guard was covered with plastic mesh (diameter 7.6 cm, 5 mm opening) secured with cable ties.  Beginning Oct 2008, guards were deployed with no mesh or plastic mesh because young fish, shrimp, and crabs get into the guard and then cannot get out and fouling on the cloth mesh itself was creating problems.  The plastic mesh had problems with fouling as well; therefore, currently guards are deployed with no mesh.  Occasionally, algae gets wrapped around the wiper at the VA station and we tried copper mesh for a few deployments as an alternative to cloth or plastic as a solution but decided to use no mesh for all deployments.

As a quality assurance check, field calibrations are recorded during sonde deployment and retrieval  using a YSI Professional Plus handheld or EXO2 sonde to collect measurements of Water Temperature, Specific Conductivity, Salinity, pH, Turb, DO, and depth.   We began using the YSI Professional Plus handheld model July 22, 2013,: Pro 101020030; 10 m cable length which is deployed near the bottom at the level of the sonde sensor array.  The near bottom measurements of water parameters are reported on the digital field logs.  The field calibration check method performed with the EXO2 sonde involves collecting surface water in a bucket and placing the retrieved old EXO2 sonde and the new deployment sonde in the bucket at the same time and recording measurements from each instrument. Beginning 4/6/2012, a biweekly field check was implemented during periods of the heavy fouling season, where technicians make a visit (approximately 2 weeks after deployment) to the water quality sites to check and document fouling, clean probes, verify the file is collecting data, and check telemetry equipment.  This field check has improved the quality assurance/quality control procedures and helps with flagging problem datasets.

Calibration procedures

Sonde cleaning and calibration of the ROX DO, Conductivity, Depth, pH, and Turbidity probes are performed as outlined in the YSI manual.  For Conductivity and Salinity, YSI solution 3168 (10,000 μS/cm) is used and for pH calibrations YSI 3822 pH 7 and 3823 pH 10 solutions are used with temperature compensation.  Turbidity calibrations are completed with a two-point calibration using 0 NTU (RO water) and 124 FNU (YSI 6073G).  The depth sensor is calibrated in air at sea level.    For DO post-calibration, recording Dissolved Oxygen in 100% saturated water method is used: two data points are recorded while the sonde is still sampling in unattended mode and has been placed in a bucket of saturated tap water.  The bucket is kept saturated with a small aquarium pump/air stones.  The other post-deployment calibrations are performed in the lab after some sonde/sensor cleaning to prevent contamination of the standards and still determine if instrument drift has occurred and evaluate the validity of the data.

Chlorophyll

A one-point (0 NTU) Chlorophyll calibration was performed until 04/02/04. Beginning April 2004, a two-point calibration is completed for chlorophyll using 0 ug/L (RO water) and 118 ug/L (rhodamine B dye) with 118 ug/L adjusted for temperature (range 100-144 µg/L corresponding to temperature ranges from 30ºC - 8ºC).  We began uploading chlorophyll fluorescence data in 2009 (Jan 2009 for Charleston and Valino stations and March 2009 for Winchester and Sengstacken stations).  Current QA/QC of the chlorophyll fluorescence parameter includes flagging erroneous spike values (> 50 ug/L) as -3, CSM, small negative values (0 to -1) with 1, CAF and negative values (<-1) -3, 1 SIC.  Currently, we do not calibrate the chlorophyll sensor using a known concentration of phytoplankton from fluorometric analysis or adjust the data with this correction, as suggested by YSI 6-Series Manual; however, we plan to complete this using grab samples collected during nutrient monitoring.  The threshold for the maximum value (50 ug/L) was chosen based on extractive fluorometric measurements from our grab samples and other investigations, which include field and lab measurements of chlorophyll within South Slough (Roegner and Curtis 2001; Cowlishaw 2002).  For example, maximum summer concentrations measured by a Seabird 10 CTD equipped with a Wetstar fluorometer were 15 mg m-3, and total chlorophyll maximum concentrations measured by extractive analysis in August 2001 by a different study were 6.4 ug l-1.  Maximum concentrations measured by our nutrient grab samples were 40.60 ug l-1 (July 2008 at the Boathouse station) and maximum concentrations measured by the ISCO diel sampling (CH station) were 45.58 ug l-1 in Sept 2008.  Despite the high seasonal maximum chlorophyll concentrations measured from our nutrient grab and isco samples, the mean concentrations are generally low (< 5 mg m-3) at all sites (O’Higgins and Rumrill 2007).

Chlorophyll Fluorescence qualifier:  YSI chlorophyll sensors (6025) are designed to estimate chlorophyll concentrations in the field for monitoring applications and complement the traditional lab extractive methods ; therefore, there are accuracy limitations associated with the data that are detailed in the YSI manual including interference from other fluorescent species, lack of calibration reagents, and effects of cell structure, particle size, organism type, temperature, and light on sensor measurements.

Additional post calibration procedures

South Slough added an additional post-calibration procedure.  From April 2011-December 2014, when sondes returned from the field, they were immediately put into a flowing seawater bath in the lab along with a clean and freshly calibrated sonde designated the “Cal Check Sonde”.  The sondes were left to run for at least three 15 minute cycles.  Then the files were opened and the data from those three 15 minute cycles were recorded and compared.  This provided additional information about bio-fouled conditions of the sondes retrieved from the field.  The routine SWMP-mandated post-deployment calibration check on cleaned sondes using calibration solutions was used to determine sensor drift independent of fouling effects.  This helped provide additional information about our data accuracy.

From April 2011-December 2014, we also started a procedure for checking the accuracy of used solutions by comparing them against new solutions using the clean (freshly cleaned) cal check sonde.  Once a used calibration solution gives a relative reading that exceeds the specified variation when compared to a new batch of solution, it is retired to the “rinse” bottle to be used for rinsing sondes prior to immersing them in new solutions.  This should improve the accuracy of our used calibration solutions most of the time.  The dates of periodic checks on the calibration solutions, the hand-held (against the cal check sonde), and the calibration of the cal check sonde are recorded on a white board located in the lab.

In January 2015, once we began deploying EXO2 sondes, we returned to standard post calibration procedures without the use of a Cal Check Sonde.

Real-time data

Sutron Sat-Link2 transmitters were installed at the Charleston Bridge and Valino Island stations on 11/15/06, at the Winchester Creek station on 12/19/2009, and at the Elliot Creek station on 8/9/2012  and transmits data to the NOAA GOES satellite, NESDIS ID #3B002196, #3B03F0F0, #3B04414A, and #3B0437DA.  (Where 3B002196, 3B03F0F0, 3B04414A, and 3B0437DA are the GOES IDs for these particular stations.)  The transmissions are scheduled hourly (soschwq 34:30; sosvawq 59:30; soswiwq 35:50; sosecwq 35:40) and contain four (4) data sets reflecting fifteen minute data sampling intervals.  Upon receipt by the CDMO, the data undergoes the same automated primary QAQC process detailed in Section 2 above.  The “real-time” telemetry data become part of the provisional dataset until undergoing secondary and tertiary QAQC and assimilation in the CDMO’s authoritative online database.  Provisional and authoritative data are available at http://cdmo.baruch.sc.edu.

5)  Site location and character –

South Slough is characterized as a drowned river-mouth estuary in the Lower Columbia Biogeographic Region.  South Slough National Estuarine Research Reserve is representative of the mid-Pacific portion of the Columbian Bioregion, extending from Cape Mendocino, CA north to the Columbia River, OR estuary.  The South Slough, located five miles southwest of the city of Coos Bay, is one of seven tidal inlets that collectively form the Coos estuary (Rumrill, 2006).  The four sites within the South Slough that are monitored with YSI data loggers are:

- Sengstacken Arm (EC) (Elliot Creek): tower [43.296501212 N,  -124.310729269 W]

This station serves as a reference site as it is located within the riverine (southern) region of the more pristine Sengstacken Arm and is located adjacent to the Elliot Creek confluence.  The site is approximately 5 km from the mouth of the South Slough estuary and is located across from the Elliot Creek drainage, which receives freshwater input on order of 799 ha.  The maximum tidal ranges can reach 2.3 m, and monthly mean Salinity ranges between 3-28 ppt.  Mean temperatures range from 7-18 oC.  Freshwater inputs have not been quantified.  Tidal waters in the riverine channels are typically well mixed in summer when the Salinity ranges from 5-30 ppt, and they are well flushed and partially stratified during the rainy season when the Salinities fluctuate from 0-21 ppt.  The data sonde is deployed in a tidal channel approximately 3 m deep and the sensor arrays are located 50 cm off the bottom.   Data is collected from this location beginning 7/23/12 14:45 – present.

Discontinued Sengstacken Arm (SE) (Talbot Creek):  piling [43.280683 N, -124.302617 W]

This station served as a reference site and was located within the riverine (southern) region of the Sengstacken Arm (Talbot Creek) immediately adjacent to 4.3 ha of natural salt marsh.  The site was approximately 6.6 km from the mouth of the South Slough estuary.  The maximum tidal ranges can reach 2.3 m, and monthly mean Salinity ranges between 3-28 ppt.  Mean temperatures range from 7-18 oC.  Freshwater inputs have not been quantified.  Tidal waters in the riverine channels are typically well mixed in summer when the Salinity ranges from 5-30 ppt, and they are well flushed and partially stratified during the rainy season when the Salinities fluctuate from 0-21 ppt.  The data sonde was deployed in a tidal channel approximately 3 m deep and the sensor arrays were located 15 cm off the bottom.   Data was collected from this location from 4/4/95 11:45- 6/16/12 07:15.

2) Winchester Arm (WI) (Winchester Tidelands Restoration Project site): tower [43.282464258 N,

-124.320424086 W]; piling [43.282472526 N, -124.320316882 W]

This station serves as a management-treatment site and is located within the riverine (southern) portion of the Winchester Arm (Kunz Marsh) and is approximately 7.1 km from the mouth of the South Slough estuary.  The site is adjacent to an experimental marsh restoration area, where dikes were removed from a 5.1 ha area in 1996 and 1998.  About 4.2 hectares of natural marsh surround the project site. Maximum tidal amplitudes can reach 2.6 m, monthly mean Salinities are between 4-20 ppt, and mean Temperatures range from 8-18 oC.  Freshwater inputs vary seasonally between 1-23 cfs, depending on precipitation events.  Similar to the Sengstacken site, the YSI data sonde is deployed in a 3 m deep tidal channel with the sensor arrays located 50 cm above the bottom.

The Winchester station was moved about 8.75 m northwest of the previous piling location on 12/18/2009; this new location was selected for telemetry reasons because it allows a clear line of sight to the GOES west satellite.  The old location will be referred to as piling and the new location will be referred to as tower since equipment is installed on a stainless steel tripod tower.  Data was collected from the piling location 4/4/95 11:45- 12/18/09 15:30 and from the new tower location beginning 12/18/09 15:45 to present.

Both sites 1 and 2 are located in tidal channels immediately adjacent to intertidal marshes in habitats characterized by soft mud and eelgrass (Zostera marina).  The primary contrasts between the two sites include differences in channel width, maximum channel depth, and orientation of the channel with regard to the predominant north-south winds.  The hillsides above both sites are composed primarily of second-growth coniferous forest, and there are no industrial sites in the watershed.  One small rural/residential dwelling is located in the Upper Winchester watershed outside the South Slough Reserve’s boundary.

- Valino Island (VA):  [43.317237439 N, -124.321647335 W]

This station serves as a secondary reference area and is located immediately northeast of Valino Island, in the southern portion of marine-dominated region of the estuary.  Freshwater input is mainly from rainfall and via three primary creeks that drain adjacent watershed sub-basins.  Additional freshwater enters this region indirectly during ebb tides from the Winchester and Sengstacken arms of the estuary.  The Valino site is approximately 3.6 km from the mouth of South Slough estuary.  Situated in the primary tidal channel, the station is adjacent to productive eelgrass beds, tidal mudflats, sandflats, and commercial oyster cultivation areas.  Maximum tidal ranges can reach 3.8 m and monthly mean Salinities fluctuate seasonally from 5-31 ppt.  The tidal channel is about 4 m deep and the YSI sensors are located 50 cm off the bottom.  There are no known point source pollutants although high bacterial levels from unknown sources cause frequent regulatory problems for commercial oyster growers.

- Charleston Bridge (CH):  [43.337669435 N, -124.320532584 W]

This station serves as another reference area and is located immediately south of the Charleston Bridge in the northern end of the marine-dominated portion of the estuary.  The site is approximately 1.2 km from the mouth of South Slough estuary and 2.4 km north of the Valino site.  The monitoring station is affixed to a timber and piling dock with controlled access by land for security purposes.  Situated in the primary tidal channel, the station is adjacent to extensive eelgrass beds, tidal mudflats, sandflats, and commercial oyster cultivation areas.  The maximum tidal amplitudes can reach 3.9 m, monthly mean Temperatures range from 8-15 oC, and the monthly mean Salinities fluctuate seasonally from 20-33 ppt.  The tidal channel is 4 meters deep and the YSI sensors are located 50 cm off the bottom.  The adjacent Joe Ney Slough has a decade-long history of tributyltin pollution.  The area was voluntarily cleaned up in the mid-1990s and passed DEQ site review.  Occasional high bacterial levels are noted.  Freshwater input is limited to seasonal rainwater and a small transient salinity/density signal from Joe Ney Slough.

6)  Data collection period –

The data loggers were first deployed at the Winchester Creek site on April 04, 1995 and at the discontinued Sengstacken (Talbot Creek) site on April 19, 1995.  The Valino Island data logger was deployed on June 03, 1999 and the Charleston Bridge station sonde was deployed on April 19, 2002.  Data loggers were deployed at Sengstacken Arm, Elliot Creek on July 23, 2012.  All monitoring is long term with a data collection period of approximately one month.  The collection dates and times are included in the annual metadata document.

| Site | Deployment Date/Time | Retrieval Date/Time |
| --- | --- | --- |
| Elliot Creek | 12/31/2016 15:30 | 01/26/2017 10:30 |
|  | 01/27/2017 12:15 | 02/28/2017 14:00 |
|  | 02/28/2017 14:15 | 03/27/2017 12:15 |
|  | 03/27/2027 12:30 | 04/27/2017 14:00 |
|  | 04/27/2017 14:15 | 05/26/2017 13:00 |
|  | 05/26/2017 13:15 | 06/23/2017 12:30 |
| Winchester Creek | 12/31/2016 15:15 | 01/26/2017 10:45 |
|  | 01/26/2017 11:15 | 02/28/2017 14:15 |
|  | 02/28/2017 14:45 | 03/27/2017 12:30 |
|  | 03/27/2017 13:00 | 04/27/2017 13:30 |
|  | 04/27/2017 14:00 | 05/26/2017 12:15 |
|  | 05/26/2017 12:45 | 06/23/2017 12:00 |
| Valino Island | 12/31/2016 16:00 | 01/26/2017 11:15 |
|  | 01/26/2017 11:30 | 02/28/2017 15:00 |
|  | 02/28/2017 15:15 | 03/27/2017 13:00 |
|  | 03/27/2017 13:30 | 05/01/2017 16:30 |
|  | 05/01/2017 16:45 | 05/26/2017 12:00 |
|  | 05/26/2017 12:15 | 06/23/2017 12:45 |
| Charleston Bridge | 12/09/2016 15:00 | 01/06/2017 15:00 |
|  | 01/06/2017 15:30 | 02/08/2017 15:15 |
|  | 02/08/2017 15:30 | 03/07/2017 15:15 |
|  | 03/07/2017 15:30 | 04/07/2017 15:00 |
|  | 04/07/2017 15:15 | 05/05/2017 15:45 |
|  | 05/05/2017 16:00 | 06/06/2017 15:00 |
|  | 06/06/2017 15:30 | 07/07/2017 12:30 |

7)  Distribution –

NOAA retains the right to analyze, synthesize and publish summaries of the NERRS System-wide Monitoring Program data.  The NERRS retains the right to be fully credited for having collected and process the data.  Following academic courtesy standards, the NERR site where the data were collected should be contacted and fully acknowledged in any subsequent publications in which any part of the data are used.  The data set enclosed within this package/transmission is only as good as the quality assurance and quality control procedures outlined by the enclosed metadata reporting statement.  The user bears all responsibility for its subsequent use/misuse in any further analyses or comparisons.  The Federal government does not assume liability to the Recipient or third persons, nor will the Federal government reimburse or indemnify the Recipient for its liability due to any losses resulting in any way from the use of this data.

Requested citation format:

NOAA National Estuarine Research Reserve System (NERRS). System-wide Monitoring Program. Data accessed from the NOAA NERRS Centralized Data Management Office website: http://www.nerrsdata.org/; accessed 12 October 2012..

NERR water quality data and metadata can be obtained from the Research Coordinator at the individual NERR site (please see Principal Investigators and Contact Persons), from the Data Manager at the Centralized Data Management Office (please see personnel directory under the general information link on the CDMO home page) and online at the CDMO home page www.nerrsdata.org.  Data are available in comma delimited format.

8)  Associated researchers and projects (link to other products or programs) –

As part of the SWMP long-term monitoring program, South Slough NERR also monitors 15-minute meteorological along with monthly grab samples and diel sampling for nutrient data which may be correlated with this water quality dataset.  These data are available at www.nerrsdata.org.

Several local researchers are currently using water quality and nutrient data recorded at the South Slough NERR SWMP stations.  Three master’s students, funded on the NERRS Science Collaborative project “A Collaborative Approach to Address Reproduction, Larval Supplies, and Settlement During Recovery of Native Olympia Oysters” (2011-2014), have accessed water quality data for their projects.  Mark Oates defended his thesis: Observations of Gonad Structure and Gametogenic Timing in a Recovering Population of Ostrea lurida (Carpenter 1864) in June 2013.  Cate Pritchard defended her thesis on native oyster larval supplies and larval distribution patterns on December 5, 2013.  Rose Rimler studied larval settlement and post-settlement mortality as determinants of the spatial distribution of Olympia oysters in Coos Bay and defended her thesis on March 7, 2013.    Matt Gray, Oregon State University, is a PhD student and Graduate Research Fellow studying interspecific competition between Olympia oysters (Ostrea lurida) and the introduced pacific oyster (Crassostrea  gigas) focusing on feeding ecology.  Molly O’Neill studied seasonal hydrography and hypoxia in Coos Bay, Oregon and utilized water quality data from the Charleston Bridge station; defended 7/17/2014.

The Partnership for Coastal Watersheds, a NERRS Science Collaborative project, utilizes SWMP water, nutrient, and weather data for the State of the Watersheds Assessment and Environmental Indicators of the project, which includes development of environmental status and trends/condition summaries and report for South Slough and Coos Bay.  The goal of the project is to provide tools for the community to understand long-term environmental and socioeconomic change in the watershed including the expansion of South Slough NERR’s estuarine water quality network to the Coos estuary in partnership with local tribes, the facilitation of a Coos estuary hydrodynamic model with University of Oregon professor Dave Sutherland, and the development of a suite of environmental and socioeconomic indicators for the community.

The SSNERR water quality monitoring project has been incorporated into the State of Oregon Coastal Non-Point Pollution Control Program (S. 6217) developed by the Oregon Department of Land Conservation and Development and the Oregon Department of Environmental Quality in July 1995.

South Slough NERR is a partner in the Northwest Association of Networked Ocean Observing Systems (NANOOS).  SWMP data collected by the South Slough NERR is incorporated into the NANOOS data management system to address critical environmental issues for sustainable development of the Pacific Northwest, including hydrodynamic modeling, estuarine water quality ie real-time water quality data products for oyster growers, estuarine ecosystem management and restoration, coastal storms and erosion, and search and rescue operations.

II.  Physical Structure Descriptors

9)  Sensor specifications –

[Instructions/Remove: Include the parameter description, units, sensor type, model #, range of measurement, accuracy and resolution for each sensor for all measuring devices (6600, 6600 EDS, 6600 EDS V2, 6600 V2, or EXO).  Specify if all of your sondes are the same model and have the same configuration.  If not, detail how many of each model you have, what different sensor configurations you use, and where the different models/configurations are deployed.  See the following example, update for your sondes/sensors (do not include sondes/sensors that were not in use), and include the disclaimers below. ]

South Slough deployed EXO2 sondes in 2017. EXO2 sondes were first deployed at Charleston Bridge on 1/30/15 14:15, at Winchester Creek on 3/4/15 at 11:45, at Elliot Creek on 3/6/15 13:00, and at Valino Island on 3/6/15 14:45.  ROX DO sensors were deployed at all sites.

YSI EXO Sonde:

Parameter: Temperature

Units: Celsius (C)

Sensor Type: CT2 probe, Thermistor

Model#: 599870

Range: -5 to 50 C

Accuracy: -5 to 35: +/- 0.01, 35 to 50: +/- .005

Resolution: 0.01 C

Parameter: Conductivity

Units: milli-Siemens per cm (mS/cm)

Sensor Type: CT2 probe, 4-electrode cell with autoranging

Model#: 599870

Range: 0 to 200 mS/cm

Accuracy: 0 to 100: +/- 0.5% of reading or 0.001 mS/cm; 100 to 200: +/- 1% of reading

Resolution: 0.001 mS/cm to 0.1 mS/cm (range dependant)

Parameter: Salinity

Units: practical salinity units (psu)/parts per thousand (ppt)

Sensor Type: CT2 probe, Calculated from conductivity and temperature

Range: 0 to 70 psu

Accuracy: +/- 1.0% of reading pr 0.1 ppt, whichever is greater

Resolution: 0.01 psu

OR

Parameter: Temperature

Units: Celsius (C)

Sensor Type: Wiped probe; Thermistor

Model#: 599827

Range: -5 to 50 C

Accuracy: ±0.2 C

Resolution: 0.001 C

Parameter: Conductivity

Units: milli-Siemens per cm (mS/cm)

Sensor Type: Wiped probe; 4-electrode cell with autoranging

Model#: 599827

Range: 0 to 100 mS/cm

Accuracy: ±1% of the reading or 0.002 mS/cm, whichever is greater

Resolution: 0.0001 to 0.01 mS/cm (range dependent)

Parameter: Salinity

Units: practical salinity units (psu)/parts per thousand (ppt)

Model#: 599827

Sensor Type: Wiped probe; Calculated from conductivity and temperature

Range: 0 to 70 ppt

Accuracy: ±2% of the reading or 0.2 ppt, whichever is greater

Resolution: 0.01 psu

Parameter: Dissolved Oxygen % saturation

Sensor Type: Optical probe w/ mechanical cleaning

Model#: 599100-01

Range: 0 to 500% air saturation

Accuracy: 0-200% air saturation: +/- 1% of the reading or 1% air saturation, whichever is greater 200-500% air saturation: +/- 5% or reading

Resolution: 0.1% air saturation

Parameter: Dissolved Oxygen mg/L (Calculated from % air saturation, temperature, and salinity)

Units: milligrams/Liter (mg/L)

Sensor Type: Optical probe w/ mechanical cleaning

Model#: 599100-01

Range: 0 to 50 mg/L

Accuracy: 0-20 mg/L: +/-0.1 mg/l or 1% of the reading, whichever is greater

20 to 50 mg/L: +/- 5% of the reading

Resolution: 0.01 mg/L

Parameter: Non-vented Level - Shallow (Depth)

Units: feet or meters (ft or m)

Sensor Type: Stainless steel strain gauge

Range: 0 to 33 ft (10 m)

Accuracy: +/- 0.013 ft (0.004 m)

Resolution: 0.001 ft (0.001 m)

Parameter: pH

Units: pH units

Sensor Type: Glass combination electrode

Model#: 599701(guarded) or 599702(wiped)

Range: 0 to 14 units

Accuracy: +/- 0.01 units within +/- 10° of calibration temperature, +/- 0.02 units for entire temperature range

Resolution: 0.01 units

Parameter: Turbidity

Units: formazin nephelometric units (FNU)

Sensor Type: Optical, 90 degree scatter

Model#: 599101-01

Range: 0 to 4000 FNU

Accuracy: 0 to 999 FNU: 0.3 FNU or +/-2% of reading (whichever is greater); 1000 to 4000 FNU +/-5% of reading

Resolution: 0 to 999 FNU: 0.01 FNU, 1000 to 4000 FNU: 0.1 FNU

Parameter: Chlorophyll

Units: micrograms/Liter

Sensor Type: Optical probe

Model#: 599102-01

Range: 0 to 400 ug/Liter

Accuracy: Dependent on methodology

Resolution: 0.1 ug/L chl a, 0.1% FS

Depth Qualifier:

The NERR System-Wide Monitoring Program utilizes YSI data sondes that can be equipped with either vented or non-vented depth/level sensors.  Readings for both vented and non-vented sensors are automatically compensated for water density change due to variations in temperature and salinity; but for all non-vented depth measurements, changes in atmospheric pressure between calibrations appear as changes in water depth.  The error is equal to approximately 1.02 cm for every 1 millibar change in atmospheric pressure, and is eliminated for vented sensors because they are vented to the atmosphere throughout the deployment time interval.

Beginning in 2006, NERR SWMP standard calibration protocol calls for all non-vented depth sensors to read 0 meters at a (local) barometric pressure of 1013.25 mb (760 mm/hg).  To achieve this, each site calibrates their depth sensor with a depth offset number, which is calculated using the actual atmospheric pressure at the time of calibration and the equation provided in the SWMP calibration sheet or digital calibration log.  This offset procedure standardizes each depth calibration for the entire NERR System.  If accurate atmospheric pressure data are available, non-vented sensor depth measurements at any NERR can be corrected.

In 2010, the CDMO began automatically correcting depth/level data for changes in barometric pressure as measured by the Reserve’s associated meteorological station during data ingestion.  These corrected depth/level data are reported as cDepth and cLevel, and are assigned QAQC flags and codes based on QAQC protocols.  Please see sections 11 and 12 for QAQC flag and code definitions.

NOTE:  older depth data cannot be corrected without verifying that the depth offset was in place and whether a vented or non-vented depth sensor was in use.  No SWMP data prior to 2006 can be corrected using this method.  The following equation is used for corrected depth/level data provided by the CDMO beginning in 2010:

((1013-BP)*0.0102)+Depth/Level = cDepth/cLevel.

Salinity Units Qualifier:

In 2013, EXO sondes were approved for SWMP use and began to be utilized by Reserves. While the 6600 series sondes report salinity in parts per thousand (ppt) units, the EXO sondes report practical salinity units (psu). These units are essentially the same and for SWMP purposes are understood to be equivalent, however psu is considered the more appropriate designation. Moving forward the NERR System will assign psu salinity units for all data regardless of sonde type.

Turbidity Qualifier:

In 2013, EXO sondes were approved for SWMP use and began to be utilized by Reserves. While the 6600 series sondes report turbidity in nephelometric turbidity units (NTU), the EXO sondes use formazin nephelometric units (FNU). These units are essentially the same but indicate a difference in sensor methodology, for SWMP purposes they will be considered equivalent.  Moving forward, the NERR System will use FNU/NTU as the designated units for all turbidity data regardless of sonde type. If turbidity units and sensor methodology are of concern, please see the Sensor Specifications portion of the metadata.

Chlorophyll Fluorescence Disclaimer:

YSI chlorophyll sensors (6025 or 599102-01) are designed to serve as a proxy for chlorophyll concentrations in the field for monitoring applications and complement traditional lab extraction methods; therefore, there are accuracy limitations associated with the data that are detailed in the YSI manual including interference from other fluorescent species, differences in calibration method, and effects of cell structure, particle size, organism type, temperature, and light on sensor measurements.

10)  Coded variable definitions –

Sampling station:		Sampling site code:	Station code:

Sengstacken Arm		EC			sosecwq

Winchester Arm		WI			soswiwq

Valino Island		VA			sosvawq

Charleston Bridge		CH			soschwq

11)  QAQC flag definitions –

QAQC flags provide documentation of the data and are applied to individual data points by insertion into the parameter’s associated flag column (header preceded by an F_).   During primary automated QAQC (performed by the CDMO), -5, -4, and -2 flags are applied automatically to indicate data that is missing and above or below sensor range.  All remaining data are then flagged 0, passing initial QAQC checks.   During secondary and tertiary QAQC 1, -3, and 5 flags may be used to note data as suspect, rejected due to QAQC, or corrected.

-5	Outside High Sensor Range

-4	Outside Low Sensor Range

-3		Data Rejected due to QAQC

-2		Missing Data

-1		Optional SWMP Supported Parameter

0		Data Passed Initial QAQC Checks

1		Suspect Data

2		Open - reserved for later flag

3		Calculated data: non-vented depth/level sensor correction for changes in barometric pressure

4		Historical Data:  Pre-Auto QAQC

5		Corrected Data

12)  QAQC code definitions –

QAQC codes are used in conjunction with QAQC flags to provide further documentation of the data and are also applied by insertion into the associated flag column.  There are three (3) different code categories, general, sensor, and comment.  General errors document general problems with the deployment or YSI datasonde, sensor errors are sensor specific, and comment codes are used to further document conditions or a problem with the data.  Only one general or sensor error and one comment code can be applied to a particular data point, but some comment codes (marked with an * below) can be applied to the entire record in the F_Record column.

General Errors

GIC	No instrument deployed due to ice

GIM	Instrument malfunction

GIT	Instrument recording error; recovered telemetry data

GMC 	No instrument deployed due to maintenance/calibration

GNF	Deployment tube clogged / no flow

GOW	Out of water event

GPF	Power failure / low battery

GQR	Data rejected due to QA/QC checks

GSM	See metadata

Corrected Depth/Level Data Codes

GCC	Calculated with data that were corrected during QA/QC

GCM	Calculated value could not be determined due to missing data

GCR	Calculated value could not be determined due to rejected data

GCS	Calculated value suspect due to questionable data

GCU 	Calculated value could not be determined due to unavailable data

Sensor Errors

SBO	Blocked optic

SCF	Conductivity sensor failure

SCS		Chlorophyll spike

SDF	Depth port frozen

SDG	Suspect due to sensor diagnostics

SDO	DO suspect

SDP	DO membrane puncture

SIC		Incorrect calibration / contaminated standard

SNV	Negative value

SOW	Sensor out of water

SPC	Post calibration out of range

SQR	Data rejected due to QAQC checks

SSD	Sensor drift

SSM	Sensor malfunction

SSR		Sensor removed / not deployed

STF	Catastrophic temperature sensor failure

STS		Turbidity spike

SWM	Wiper malfunction / loss

Comments

CAB*	Algal bloom

CAF	Acceptable calibration/accuracy error of sensor

CAP	Depth sensor in water, affected by atmospheric pressure

CBF	Biofouling

CCU	Cause unknown

CDA*	DO hypoxia (<3 mg/L)

CDB*	Disturbed bottom

CDF	Data appear to fit conditions

CFK*	Fish kill

CIP	*	Surface ice present at sample station

CLT*	Low tide

CMC*	In field maintenance/cleaning

CMD*	Mud in probe guard

CND	New deployment begins

CRE*	Significant rain event

CSM*	See metadata

CTS	Turbidity spike

CVT*	Possible vandalism/tampering

CWD*	Data collected at wrong depth

CWE*	Significant weather event

13)  Post deployment information –

Included in annual metadata document.

14)  Other remarks/notes –

Data are missing due to equipment or associated specific probes not being deployed, equipment failure, time of maintenance or calibration of equipment, or repair/replacement of a sampling station platform.  Any NANs in the dataset stand for “not a number” and are the result of low power, disconnected wires, or out of range readings.  If additional information on missing data is needed, contact the Research Coordinator at the reserve submitting the data.

Elliot Creek

Data are missing 01/26/17 10:45-01/27/17 12:00 and are flagged -2 GMC CSM.

pH post calibration for the February deployment is high and data between old and new deployments is slightly disjunct.  Data from 02/26/17 12:30 – 02/28/17 14:00 are flagged 1 SPC CSM.

Specific conductivity, salinity, DO mg/l, and depth data are flagged 1 SWM CSM from 01/25/17 11:30 – 01/26/17 10:30.  Retrieval notes indicated a wiper malfunction and field checks for specific conductivity/salinity were much lower for the old sonde compared to the new sonde readings.

Turbidity values for the April, May, and June deploymenst spike during low tides and then recover to low values, values > 100 FNU are flagged 1 STS CSM and values > 4000 FNU are flagged -3 STS CSM.

Specific conductivity, salinity, and pH are slightly disjunct between the old and new deployment on 04/27/17 14:15, but post calibrations for all sensors were good.  Data for the last week of the deployment 4/21/17 2:30 – 04/27/17 14:00 were flagged 1 SSD CSM.

Depth values on 5/25/17 6:15-6:45 were negative and flagged 1 SNV CAP.

Specific conductivity for the May deployment post calibration was high (10.18), but data from old to new deployment match up well and the field check matches well with the old deployment values for 5/26/17 13:00; therefore, data were not flagged.

Water temperature drops on 06/06/17 4:45 and data is flagged 1 SSM CSM.

Dissolved oxygen values drop (3.9-5.9 mg/L)  during the following periods at low spring tides and following tide exchanges (flooding or ebbing) and data are flagged 1 GSM CLT: 6/01/17 6:30, 6/01/17 6:30, 06/02/17 00:45, 06/03/17 1:30, 06/04/17 2:45-3:00, 6/9/17 1:00, 6/10/17 1:30, 6/11/17 2:00-2:15, 6/12/17 2:30, 6/15/17 9:30, 19:15, 6/16/17 2:30, 5:30, 11:30, 12:15,  6/17/17 4:15-30, 5:45-6:30, 6/19/17 2:30, 6/20/17 9:45.

Winchester Creek

pH data on 01/08/17 20:15, 01/10/17 6:00, 01/11/17 7:30, 01/22/17 18:30, and 01/25/18 05:00, 2/10/17 18:15-18:45, 3/7/17 4:00, 3/8 5:30, 3/19/17 16:15, 3/21/17 16:45, 3/22/17 17:15, 3.23/17 17:45, 3/24/17 20:45, 3/26/17 7:30 are flagged 1 SSM CSM; data values drop during flooding tides and/or low salinities.

Turbidity data from 01/21/17 – 01/26/17 10:45 are flagged 1 SWM CSM or -3 SWM CSM (if > 4000 FNU). Field notes indicate the wiper was missing and turbidity data are elevated at the end of the deployment.

Dissolved oxygen values drop during the following periods and data are flagged 1 SSM CSM: 01/01/17 7:45-8:00, 15:30, 01/04/17 05:00, 7:00, 8:00-15, 01/05/17 9:00, 9:30, 10:15, 01/06/17 9:45, 10:30, 01/08/17 8:15, 8:45, 01/15/17 10:00, 01/16/17 5:15, 14:30, 01/17/17 10:15, 10:45, 01/18/17 9:00, 01/19/17 7:30, 01/21/17 11:00-12:15, 01/22/17 12:30-13:00, 01/23/17 13:00, 01/24/17 11:00-15, 01/26/17 6:15.

Specific conductivity, salinity, DO %, DO mg/l, and depth are flagged -3 SSM CSM from 2/14/17 4:15 – 2/28/17 14:15.  Data are very disjunct from old to new deployment, post calibration was low for sp cond but ok for Dissolved oxygen, and data also appear to be impacted significantly by rain events.

pH data from 2/14/17 4:15 – 2/28/17 14:15 are flagged 1 SPC CSM, data are disjunct between old and new deployments and post calibration was high (7.14, 10.14).

Turbidity data during the March deployment was impacted by rain events and data from 2/5/17 12:15 – 2/28/17 14:15 are flagged 1 GSM CRE or -3 GSM CRE if values were > 4000 FNU.

Turbidity data from 3/8/17 20:45 – 3/27/17 13:00 are flagged 1 SWM CSM or -3 SWM CSM if values were > 4000 NTU; data values are noisy and the wiper was missing during retrieval.  There was also a lot of rain during the month of March.

Dissolved oxygen values drop unusually low during the March deployment and low values are flagged 1 SSM CSM:    3/2/17 15:30, 3/4/17 17:30-45, 3/5/17 8:00, 16:15, 17:15, 3/7/17 10:30, 3/8/17 23:00, 3/9/17 10:30, 3/12/17 8:30, 3/13/17 9:00, 3/16/17 17:00, 3/17/17 15:45,  3/19/17 16:15-17:00, 3/20/17 15:45, 3/23/17 8:30-9:00, 3/24/17 9:00-15, 11:45, 3/25/17 6:45, 3/27/17 12:15.

Data are missing 4/27/17 13:45 and 5/26/17 12:30 during sonde field deployment exchanges.

Turbidity data during the first half of the May deployment peaks during low tides from the beginning of the deployment 04/27/17 14:00 – 05/14/1711:00 and values > 100 FNU are flagged 1 STS CSM.  Then for the second half of the deployment turbidity gets noisy and remains elevated through the end of deployment; the wiper was missing during retrieval and data records from 5/16/17 6:45 – 5/26/17 12:15 were flagged -3 SWM CSM.

Dissolved oxygen values drop below 6 mg/L during the May deployment during low tides or following tide exchanges (flooding, or ebbing) and values are flagged 1 GSM CLT: 5/14/17 3:00, 17:45, 20:45-21:15, 5/16/17 20:45, 5/17/17 21:15-30, 5/18/17 22:30-23:00, 5/19/17 7:45-8:15, 22:00, 5/20/17 3:15.

Turbidity data (June deployment) spikes on 6/17/17 11:00 and 6/23/ 17 1:30 and values were flagged 1 STS CSM.

Valino Island

Turbidity data for the January deployment is noisier than usual with large spikes > 4000.  There were significant rain events as well as a missing wiper and crab in the guard upon retrieval.  Data from 1/7/17 17:30 – end of deployment are flagged 1 SWM CSM or -3 SWM CSM (if >4000 FNU).

pH data are disjunct between the old and new deployment on 01/26/17 11:30.  Post calibrations were slightly high and retrieval notes indicated the wiper was missing.  Data from 01/21/17 9:00 – 01/26 17 11:!5 are flagged 1 SPC CSM.

pH data are disjunct between the old and new deployment on 02/28/17 15:15.  Post calibrations were high.  Data from 02/17/17 19:00 – 02/28/17 15:00 are flagged 1 SPC CSM.

Specific conductivity post calibration for the April deployment is low (9.68) and sp cond/salinity data values are disjunct between the old and new deployment.  Specificu conductivity and salinity data are flagged 1 SPC CSM from 04/26/17 00:00 – 05/01/17 16:30.

Water temperature spikes on 04/3/17 14:00 and 06/01/17 13:15 during flooding tide following a spring low tide, values are flagged 1 SSM CSM.

Turbidity values spike intermittently during the April deployment and increase in frequency toward the end of the deployment.  There were 4 crabs noted in the guard upon retrieval.  Values > 97 FNU are flagged 1 STS CSM and values > 4000 FNU are flagged -3 STS CSM.

pH data is slightly disjunct at the old and new deployment switch, and the first record for the new deployment 05/01/17 16:45 is flagged 1 GSM CSM as data recovers and matches well after this record.

Data records on 5/26/17 6:15 – 6:45, 5/27/17 7:0-7:45, 5/28/17 7:45- 8:30, 6/23/17 5:00-6:00 are flagged -3 GOW CSM, and specific conductivity/salinity on 5/26/17 6:00, 7:00,  5/27/17 8:00, 5/28/17 8:30, 6/16/17 4:45 are flagged -3 SOW CSM;  temp, sp conductivity/salinity values drop suddenly during a spring low tide.

pH drops on 5/12/17 6:45 and this occurs during a low tide and although other parameters were not data was flagged -3 SOW CSM.

Turbidity values during the May deployment spike intermittently and from 5/1/17 6:45 – 5/21/17 5:15,  values > 100 FNU are flagged 1 STS CSM, values > 4000 FNU are flagged -3 STS CSM .  Toward the end of the deployment, spikes start to increase and data from 05/21/17 5:30 – 5/26/17 11:30 values are flagged 1 GSM CBF.

There are occasional turbidity spikes during the June deployment, values > 200 FNU are flagged 1 STS CSM, those > 4000 FNU are flagged -3 STS CSM.

DO % and mg/L values drop on 5/12/17 7:00-7:15, 5/21/17 4:30, 5/22/17 5:15, 6/4/17 4:15-5:00, 6/5/17 3:15-4:00, 6/6/17 4:00, 6/7/17 4:45-5:15, 6/9/17 6:15, 6/10/17 6:30, 6/11/17 7:15, 6/12/17 7:45, 6/13/17 9:15, 6/14/17 8:45, 6/16/17 10:15-10:45, 6/20/17 2:30, 6/21/17 3:15, 6/22/17 5:15 during spring low tides and values are flagged 1 GSM CLT.

Charleston Bridge

DO post calibration for the December deployment was slightly high (104.1 %), and data values at the end of deployment were a little higher compared to the following new deployment, data from 01/01/17 10:30 – 01/06/17 15:00 are flagged 1 SPC CSM.

The first pH reading at 4/7/17 15:15 for a new deployment was low and flagged 1 GSM CND.

pH and/or Dissolved oxygen readings spike during flooding tides and data are flagged 1 SSM CSM: 4/20/17 13:45, 4/21/17 14:00-15:15, 4/22/17 15:30-15:45.

Turbidity values for the April deployment have some spikes at the beginning and mid deployment, values > 87 NTU are flagged 1 STS CSM and one value > 4000 NTU was flagged -3 STS CSM.

The wiper fails during the May deployment and turbidity, DO %, and DO mg/L data from 05/23/17 01:00 – 06/06/2017 15:00 are flagged -3 SWM CSM.

Most data values on 6/05/17 12:00 are NAN, missing, or outside of sensor range and depth drops to 0.0, pH was unaffected.  There are no notes for maintenance at the site during this period.  Data are flagged -3 GOW CSM.

pH values are slightly disjunct between the old and new deployment on 5/5/17 16:00 and there was a new pH module installed.  Data from the new deployment 5/5/17 16:15-5/8/17 2:45 are flagged 1 SSM CSM as it appears the sensor is reading slightly high and then recovers.

Specific conductivity post calibration is slightly high (10.15) for the April deployment, but data from old to new deployment match up well and the field check matches well with the old deployment values for 5/5/17 15:45; therefore, data were not flagged.

Data are missing on 6/06/17 15:15 during the field exchange and data are flagged -2 GMC CSM.

Specific conductivity post calibration is low (9.75) for the June deployment, data from 6/27/17 14:45 – 7/07/17 12:30 are flagged 1 SPC CSM. Field checks between the old and new sonde also showed lower sp cond and salinity readings for the old sonde.

Chlorophyll data for all sites

As detailed in section 4 Research Methods under the Chlorophyll subheading, chlorophyll values that are > 50 ug/l are considered erroneous spike values and flagged suspect <1> SCS.  This maximum threshold number was chosen based on field studies as well as extractive chlorophyll measurements from the nutrient grab samples made within South Slough and will be used until the recommended YSI procedure is completed, using a known concentration of phytoplankton derived from extractive fluorometric analysis of our grab samples and applying the correction ratio to the field fluorescence data.  Chlorophyll probes were removed from all stations in October 2014: at Valino Island on 10/03/14 10:45, 10/08/14 17:00 Charleston Bridge, 10/03/14 10:15 Winchester Creek, 10/08/14 12:45 Elliot Creek.

Turbidity data for all sites

Turbidity events up to 500 NTU are common for this estuary, especially during the wet season.  Data over 4000 FNU are rejected according to probe specifications.  Small negative turbidity down to -2 FNU are within probe specifications, but less than that level (-3 to -10 FNU) are probably due to calibration errors resulting from water left on the probes during calibration.  These negative turbidity values (-1 to -10) are kept in the data files and documented as anomalous; -1 to -2 are coded with <1> (CAF); -5 and smaller are coded with <-3> [SIC] (CSM).  Occasional Turbidity spikes may be due to short optical interferences from biological debris, detritus, and/or organisms (fish, crab) stuck in the sonde guard or significant rain events; data records and field notes are used to determine the best possible code for these events.

References

Cornu, C.E., and J. Souder, J. Hamilton, A. Helms, R. Rimler, B. Joyce, F. Reasor, T. Pedersen, E. Wright, R. Namitz, J. Bragg, B. Tanner. 2012.  Partnership for Coastal Watersheds State of the South Slough and Coastal Frontal Watersheds.  Report prepared for the Partnership for Coastal Watersheds Steering Committee.  South Slough National Estuarine Research Reserve and the Coos Watershed Assocation.  225 pp.

O’Higgins T. and Rumrill, SS.  2007.  Tidal and watershed forcing of nutrients and dissolved oxygen stress within four pacific coast estuaries.  CICEET Final Report, 58 pp.

Rumrill, S.S.  2006.  The Ecology of the South Slough Estuary: Site Profile of the South Slough National Estuarine Research Reserve.  NOAA / Oregon Department of State Lands. 126 pp.

Cowlishaw, R.J. 2002.  Microzooplankton trophic interactions and their impact on phytoplankton production and community structure in the South Slough arm of Coos Bay, Oregon.  PhD thesis University of Oregon, 175 pp.

Roegner, GC and Shanks AL.  April 2001.  Import of coastally-derived chlorophyll a to South Slough, Oregon.  Estuaries 24 (2): 244-256.

Powell, S.L. and Rumrill, S.S.  March 2001.  South Slough National Estuarine Research Reserve:  Monitoring short-term variability and change in the estuary’s water quality.  Earth System Monitor, Vol. 11, No. 3, pp. 5-11.

Rumrill. S.S. and C.E. Cornu.  1995.  South Slough Coastal Watershed Restoration:  A case study in integrated ecosystem restoration.  Restoration and Management Notes 13:53-57.

Precipitation

Significant precipitation events greater than 6.35 mm (0.25 inch) often affect SpCond, DO, pH, and Temp.  The data are recorded at the North Bend Airport Weather Station (KOTH).  Precipitation recorded at the Tom’s Creek Meteorological SWMP Station can be downloaded from http://cdmo.baruch.sc.edu/.

Significant precipitation events above 0.25 inches of daily rain for 2017 will be reported for annual submission below as the month/day & daily precipitation in inches from North Bend Airport station.

January

1/1 0.43, 1/7 0.83, 1/8 1.97, 1/9 0.77, 1/10 1.00, 1/17 0.55, 1/18 1.57, 1/20 0.32, 1/22 0.73

February

2/3 1.23, 2/4 1.26, 2/5 1.63, 2/7 0.30, 2/8 1.22, 2/9 0.81, 2/10 0.24, 2/15 1.41, 2/16 1.85, 2/20 0.95, 2/21 1.36, 2/24 0.48, 2/26 0.62, 2/27 0.33

March

3/3 0.39, 3/4 0.59, 3/5 0.84, 3.6 0.56, 3/7 0.93, 3/8 0.44, 3/9 0.30, 3/15 0.59, 3/18 0.99, 3/22 0.25, 3/23 0.30, 3/24 2.09, 3/26 0.36, 3/29 0.63

April

4/6 0 .93, 4/9 0.38, 4/11 0.44, 4/12 0.31, 4/14 0.29, 4/16 0.39, 4/17 0.48, 4/19 0.33, 4/20 0.25, 4/23 0.57, 4/24 0.55, 4/26 0.77

May

5/11 1.01, 5/12 0.85, 5/13 1.11, 5/16 1.00