Weeks Bay (WKB) NERR Water Quality Metadata

January – June 2017

Latest Update: July 2017

Note: This is a provisional metadata document; it has not been authenticated as of its download date.  Contents of this document are subject to change throughout the QAQC process and it should not be considered a final record of data documentation until that process is complete.  Contact the CDMO (cdmosupport@belle.baruch.sc.edu) or Reserve with any additional questions.

- Data set and Research Descriptors

1. Principal investigators and contact persons

Scott Phipps, Research Coordinator, Scott.Phipps@dcnr.alabama.gov

Weeks Bay NERR

11300 U.S. Hwy 98

Fairhope, Al 36532

Ph: (251)928-9792

2. Entry verification

Deployment data are uploaded from the YSI data logger to a Personal Computer (IBM compatible).  Files are exported from EcoWatch in a comma-delimited format (.CDF) and uploaded to the CDMO where they undergo automated primary QAQC and become part of the CDMO’s online provisional database.  Pre- and post-deployment data are removed from the file prior to upload. During primary QAQC, data are flagged if they are missing, out of sensor range, or outside 2 or 3 standard deviations from the historical seasonal mean.  The edited file is then returned to the Reserve where it is opened in Microsoft Excel and processed using the CDMO’s NERRQAQC Excel macro.  The macro inserts station codes, creates metadata worksheets for flagged data, and graphs the data for review.  It allows the user to apply QAQC flags and codes to the data, remove remaining pre- and post-deployment data, append files, and export the resulting data file to the CDMO for tertiary QAQC and assimilation into the CDMO’s authoritative online database.  Where deployment overlap occurs between files, the data produced by the newly calibrated sonde is generally accepted as being the most accurate.  For more information on QAQC flags and codes, see Sections 11 and 12.

Scott Phipps was responsible for data management (see contact information above).

3. Research objectives

The Data loggers are programmed to record water quality parameters every 15 minutes.  Presently, four data logger -sites are located in the Weeks Bay Reserve.  Data was collected from Fish River, Weeks Bay, Mid Bay, and Magnolia River. (See section 5 for site description)

Beyond the mouth of Weeks Bay is a larger embayment referred to as Bon Secour Bay. Hydrographic profiles made in Bon Secour Bay in past years show that bottom waters (which flow into Weeks Bay on a rising tide) are often anoxic in summer and fall months.  Fish River water quality is also in question, as largely non-point source nutrient loading to this river may be quite high during periods of high flow in the river.  Therefore, it is of interest to determine if poor water quality periodically observed in Weeks Bay is generated within the bay proper, or is being introduced from Fish River or Bon Secour Bay with changing tides.  By using salinity and water level as trace, data from these two sites may be used to infer the source and fate of low oxygen water.

There is a great deal of value in having a baseline data set for any management site.   Changes or trends in any monitored parameters are readily apparent.  Weeks Bay is useful as a model for estuary function and this data is necessary for the profile of the model.

4. Research Methods

Data loggers are moored on private docks and channel marker pilings. At each location, a PVC tube is attached to the pilings using ratchets equipped with zinc disks (to help reduce corrosion) and nylon straps.  A chain, attached to the cap of the tube, is shackled to the bail of the Data logger.  Chains and tubes are cut to a length that places the Data logger probes just below the bottom to the tube and 0.5 m above the sediment when in place.

The Data loggers are programmed to record temperature, specific conductivity, salinity, dissolved oxygen (DO) percent saturation and concentration, pH, water level and turbidity every 15 minutes central standard time (CST).  When available, a newly calibrated Data logger is also used to record the above parameters at the time of Data logger exchange.  The instruments are usually serviced (instruments post-calibrated, data uploaded, instruments cleaned and re-calibrated) every 7 to 14 days; however, they are programmed to log for 365 days in the event they cannot be serviced on schedule.

Before the Data loggers are retrieved, a newly calibrated Data logger (when available) is lowered to the depth of the deployed instrument, and readings are logged.  When time and weather conditions allow, this is done at the same time that the deployed instruments are programmed to record.  Post-calibration values are obtained in the lab by placing unclean probes in standard solutions and recording the values given by the Data loggers.

Uploading, cleaning, maintenance, and calibration are conducted as described in the YSI Operating Manual.  Calibration standards, required for pH are purchased from a scientific supply house.  Salinity standards are created in house per Standard Methods (20th Ed.).  Dissolved oxygen is calibrated using a calibration cup filled with about ½ cm tap water, which creates a 100% water-saturated air environment for the sensors when the sonde is placed in the cup.  The sensors are allowed to equilibrated for at least 1 hour in the cup before DO (% Saturation) is calibrated. Turbidity standard is purchased from YSI.

A Sutron Sat-Link2 transmitter was installed at the Fish River (FR) station on 10/11/2006 and transmits data to the NOAA GOES satellite, NESDIS ID #3B02B100.  The transmissions are scheduled hourly and contain four (4) data sets reflecting fifteen minute data sampling intervals.  Upon receipt by the CDMO, the data undergoes the same automated primary QAQC process detailed in Section 2 above.  The “real-time” telemetry data become part of the provisional dataset until undergoing secondary and tertiary QAQC and assimilation in the CDMO’s authoritative online database.  Provisional and authoritative data are available at http://cdmo.baruch.sc.edu.

5. Site location and character

Weeks Bay (30 23 ‘ N, 87 50’ W) is a small, shallow, microtidal sub-estuary,  located on the eastern shore of Mobile Bay in the northern Gulf of Mexico.  It is nearly diamond shaped, and its longitudinal axis (3.4 km long) runs nearly north –south from the head, where the Fish River flows in, to the mouth, where water is exchanged with Mobile Bay.  Its widest point (3.1 km) is located near the center of the estuary, where the Magnolia River discharges into eastern side of Weeks Bay.  Average depth is 1.4 m, although there are two areas where depths are significantly greater.  The first is in the mouth of the bay, where the average depth is 6m; the second is about 100m upstream of the mouth of the Fish River, where the average depth is 3.5m.  Tides are principally diurnal, and have a mean range of 0.4m.

The Fish River drainage basin encompasses 14300 hectares and contributes approximately 73% to the total incoming freshwater flow with the Magnolia River supplying the rest.  Mean combined discharge is 9 cubic meters per second; although freshets up to 4 times larger occur throughout the year.  These characteristics result in a freshwater residence time of 13 days under average discharge conditions, with a range from 0.5 to 100 days.  Salinity in Weeks Bay varies substantially both temporally and spatially.  During periods of high flow in the river, salinity in the bay may be fresh from the head to the mouth, except in the deeper holes of the estuary that are not as easily flushed.  However, during periods of low flow in the river, wind velocity and tidal stage are strong factors influencing salinity structure.  Salinity greater than 25 ppt is infrequently observed in Weeks Bay and is usually restricted to the southern portion of the estuary near the mouth.

Site FR (Fish River; 30 24.97’N, 87 49.37’W) is located near the mouth of Fish River at a mean depth of about 2m.  Sediment type is sandy-silt, and there are small patches of Vallisneria sp. growing near (but not directly under) the Data logger.  Land use in the water shed is agricultural, forested and residential with the residential portion rapidly increasing.  Directly surrounding the site, land use is residential and forested,  Nutrient concentrations at this site are variable and may be quite high, with nitrate and phosphate concentrations ranging from approximately 4-100 µM and <0.5 to 7.5 µM, respectively, over the previous 10 years (Alabama Department of Environmental Management, Geological Survey of Alabama).  Over the past year, salinity ranged from 0 to 20 ppt at this site. This site is transmitted to the NOAA GOES satellite (see section 4 above).

Site WB (Weeks Bay; 30 22.85’N, 87 49.92’W) is located near the southeast shore of Weeks Bay, about 0.5 kilometers from the mouth of the estuary.  Mean water depth at this site is about 0.9m; sediment type is sandy-silt.  Land use around this site is almost exclusively residential, with agriculture occurring inland.  Nutrient concentrations are much lower at this site than at Fish River.  Salinity concentrations from data collected indicates a range from approximately 5 to 25 ppt over the past year.

Site MB (Middle Bay; 30 23.52’ N, 87 49.37’ W) is located near the middle of the north/south axis of Weeks Bay.  Mean water depth at this site is about 1.5m; sediment type is sandy-silt.  The datalogger at this site is secured to a navigational channel marker.  There is no indication of pollutants and no sub-aquatic vegetation at this site.  Salinity concentrations from data collected indicates a range from approximately 0 to25 ppt.

Site MR (Magnolia River; 30 23.31’ N, 87 49.03’W) is located near the mouth of Magnolia River at a mean depth of about 1.1m.  Land use around this site is bottomland forest and shrub, with a fringing marsh.  Upriver land use is substantially residential.  The Magnolia river watershed is predominantly agricultural including turf, cotton and peanut crops as well as cattle grazing.  There is no indication of pollutants and the bottom sediment is silty to clayey with no sub-aquatic vegetation.  Salinity concentrations from data collected indicates a range from approximately 0 to 20ppt.  The datalogger at this site is secured to a navigational channel marker.

6. Data collection period

The Data loggers were first deployed at sites FR and WB on 27 October 1995, and have been continuously in service since then.  Data loggers were first deployed for sites MB and MR on 05 May 2003.

7. Distribution

NOAA retains the right to analyze, synthesize and publish summaries of the NERRS System/wide Monitoring Program data.  The NERRS retains the right to be fully credited for having collected and process the data.  Following academic courtesy standards, the NERR site where the data were collected should be contacted and fully acknowledged in any subsequent publications in which any part of the data are used.  The data set enclosed within this package/transmission is only as good as the quality assurance and quality control procedures outlined by the enclosed metadata reporting statement.  The user bears all responsibility for its subsequent use/misuse in any further analyses or comparisons.  The Federal government does not assume liability to the Recipient or third persons, nor will the Federal government reimburse or indemnify the Recipient for its liability due to any losses resulting in any way from the use of this data.

Requested citation format:

NOAA National Estuarine Research Reserve System (NERRS). System/wide Monitoring Program. Data accessed from the NOAA NERRS Centralized Data Management Office website: http://www.nerrsdata.org/; accessed 12 October 2012.

NERR water quality data and metadata can be obtained from the Research Coordinator at the individual NERR site (please see Principal Investigators and Contact Persons), from the Data Manager at the Centralized Data Management Office (please see personnel directory under the general information link on the CDMO home page) and online at the CDMO home page http://cdmo.baruch.sc.edu/.  Data are available in comma delimited format.

8. Associated researchers and projects

SWMP Nutrient monitoring. Nutrient data (ammonia, nitrate, nitrite, dissolved reactive phosphate, total dissolved phosphate and chlorophyll) are collected from all water quality sites in triplicate, monthly at low tide and from FR site over a tidal cycle monthly. These data are available at the CDMO home page http://cdmo.baruch.sc.edu/ or by contacting the Research Coordinator at Weeks Bay Reserve (see contact information above).

SWMP Meteorological monitoring. Meteorological data (air temperature, relative humidity, barometric pressure, wind speed, wind direction, rainfall, and photosynthetically active radiation) are collected every 15 minutes and are available at the CDMO home page http://cdmo.baruch.sc.edu/ or by contacting the Research Coordinator at Weeks Bay Reserve (see contact information above).

Alabama Water Watch, Auburn University.  Establishing a citizen-monitoring group in the Weeks Bay Watershed, Weeks Bay Reserve established twenty-four sites for monitoring water quality in the watershed by volunteers in May 1995.  These sites are monitored once every two weeks for fecal coliform counts, secchi depth, DO concentration, temperature, pH, alkalinity, hardness, and at some sites salinity, nitrate, phosphate, and fecal coliform levels.

United States Geological Survey.  Two USGS stage and one in stream gauging stations are maintained in the Weeks Bay watershed.

Baldwin County Soil and Water Conservation District.  Weeks Bay pollution prevention project: works with farmers, developers, and homebuilders to promote BMP’s within the watershed.

Weeks Bay Management Plan.  This effort completes the development of a management plan for the Weeks Bay NERR.

II. Physical Structure Descriptors

9. Sensor specifications

FR site for all dates incorporate only YSI 6600EDS dataloggers. WB,MR and MB sites utilized only EXO2 dataloggers:

YSI 6600EDS sensors:

Parameter: Temperature

Units: Celsius (C)

Sensor Type: Thermistor

Model #: 6560

Range: -5 to 45 ºC

Accuracy: +/-0.15 ºC

Resolution: 0.01 ºC

Parameter: Conductivity

Units: milli-Siemens per cm (mS/cm)

Sensor Type: 4-electrode cell with autoranging

Model #: 6560

Range: 0 to 100mS/cm

Accuracy: +/- 0.5% of reading + 0.1 mS/cm

Resolution: 0.001 mS/cm to 0.1 mS/cm (range dependent)

Parameter: Salinity

Units: parts per thousand (ppt)

Sensor Type: Calculated from conductivity and temperature

Range: 0 to 70 ppt

Accuracy: +/- 1.0% of reading or 0.1 ppt, whichever is greater

Resolution: 0.01 ppt

Parameter: Dissolved Oxygen % Saturation

Units: percent air saturation (%)

Sensor Type: ROX Optical Dissolved Oxygen Sensor

Model #: 6150

Range: 0 to 500 % air saturation

Accuracy: 0-200 % air saturation, +/- 1% of the reading or 1% air saturation, whichever is greater; 200-500 % air saturation, +/- 15% of the reading

Resolution: 0.1 % air saturation

Parameter: Dissolved Oxygen mg/L (Calculation from % air saturation, temperature and salinity)

Units: milligrams per Liter (mg/L)

Sensor Type: ROX Optical Dissolved Oxygen Sensor

Model #: 6150

Range: 0 to 50 mg/L

Accuracy: 0 to 20 mg/L, +/- 1 % of the reading or 0.1 mg/L, whichever is greater; 20 to 50 mg/L, +/- 15% of the reading

Resolution: 0.01 mg/L

Parameter: Non-Vented Level – Shallow (Depth)

Units: feet or meters (ft or m)

Sensor Type: Stainless steel strain gauge

Range: 0 to 30 ft (9.1 m)

Accuracy: +/- 0.06 ft (0.018m)

Resolution: 0.001 ft (0.001m)

Parameter: pH EDS flat glass

Units: pH units

Sensor Type: Glass combination electrode

Model #: 6561

Range: 0 to 14 units

Accuracy: +/- 0.2 units

Resolution: 0.01 units

Parameter: Turbidity

Units: nephelometric turbidity units (NTU)

Sensor Type: Optical, 90º scatter, with mechanical cleaning

Model #: 6136

Range: 0 to 1000 NTU

Accuracy: +/- 5% reading or 2 NTU (whichever is greater)

Resolution: 0.1 NTU

YSI EXO2 sensors:

Parameter: Temperature

Units: Celsius (C)

Sensor Type: Thermistor

Model#: 599870/01

Range: /5 to 50 C

Accuracy: /5 to 35: +// 0.01, 35 to 50: +// .005

Resolution: 0.01 C

Parameter: Conductivity

Units: milli/Siemens per cm (mS/cm)

Sensor Type: 4/electrode cell with autoranging

Model#: 599870/01

Range: 0 to 200 mS/cm

Accuracy: 0 to 100: +// 0.5% of reading or 0.001 mS/cm; 100 to 200: +// 1% of reading

Resolution: 0.001 mS/cm to 0.1 mS/cm (range dependant)

Parameter: Salinity

Units: practical salinity units (psu)/parts per thousand (ppt)

Sensor Type: Calculated from conductivity and temperature

Range: 0 to 70 psu

Accuracy: +// 1.0% of reading pr 0.1 ppt, whichever is greater

Resolution: 0.01 psu

Parameter: Dissolved Oxygen % saturation

Sensor Type: Optical probe w/ mechanical cleaning

Model#: 599100/01

Range: 0 to 500% air saturation

Accuracy: 0/200% air saturation: +// 1% of the reading or 1% air saturation, whichever is greater 200/500% air saturation: +// 5% or reading

Resolution: 0.1% air saturation

Parameter: Dissolved Oxygen mg/L (Calculated from % air saturation, temperature, and salinity)

Units: milligrams/Liter (mg/L)

Sensor Type: Optical probe w/ mechanical cleaning

Model#: 599100/01

Range: 0 to 50 mg/L

Accuracy: 0/20 mg/L: +//0.1 mg/l or 1% of the reading, whichever is greater

20 to 50 mg/L: +// 5% of the reading

Resolution: 0.01 mg/L

Parameter: Non/vented Level / Shallow (Depth)

Units: feet or meters (ft or m)

Sensor Type: Stainless steel strain gauge

Range: 0 to 33 ft (10 m)

Accuracy: +// 0.013 ft (0.004 m)

Resolution: 0.001 ft (0.001 m)

Parameter: pH

Units: pH units

Sensor Type: Glass combination electrode

Model#: 599701(guarded) or 599702(wiped)

Range: 0 to 14 units

Accuracy: +// 0.01 units within +// 10° of calibration temperature, +// 0.02 units for entire temperature range

Resolution: 0.01 units

Parameter: Turbidity

Units: formazin nephelometric units (FNU)

Sensor Type: Optical, 90 degree scatter

Model#: 599101/01

Range: 0 to 4000 FNU

Accuracy: 0 to 999 FNU: 0.3 FNU or +//2% of reading (whichever is greater); 1000 to 4000 FNU +//5% of reading

Resolution: 0 to 999 FNU: 0.01 FNU, 1000 to 4000 FNU: 0.1 FNU

Depth Qualifier:

The NERR System/Wide Monitoring Program utilizes YSI data sondes that can be equipped with either vented or non/vented depth/level sensors.  Readings for both vented and non/vented sensors are automatically compensated for water density change due to variations in temperature and salinity; but for all non/vented depth measurements, changes in atmospheric pressure between calibrations appear as changes in water depth.  The error is equal to approximately 1.02 cm for every 1 millibar change in atmospheric pressure, and is eliminated for vented sensors because they are vented to the atmosphere throughout the deployment time interval.

Beginning in 2006, NERR SWMP standard calibration protocol calls for all non-vented depth sensors to read 0 meters at a (local) barometric pressure of 1013.25 mb (760 mm/hg).  To achieve this, each site calibrates their depth sensor with a depth offset number, which is calculated using the actual atmospheric pressure at the time of calibration and the equation provided in the SWMP calibration sheet or Digital Calibration Log.  This offset procedure standardizes each depth calibration for the entire NERR System.  If accurate atmospheric pressure data are available, non-vented sensor depth measurements at any NERR site can be corrected.  The Research Coordinator at the specific NERR site should be contacted in order to obtain information regarding atmospheric pressure data availability.

In 2010, the CDMO began automatically correcting depth/level data for changes in barometric pressure as measured by the Reserve’s associated meteorological station during data ingestion.  These corrected depth/level data are reported as cDepth and cLevel, and are assigned QAQC flags and codes based on QAQC protocols.  Please see sections 11 and 12 for QAQC flag and code definitions.

NOTE:  older depth data cannot be corrected without verifying that the depth offset was in place and whether a vented or non/vented depth sensor was in use.  No SWMP data prior to 2006 can be corrected using this method.  The following equation is used for corrected depth/level data provided by the CDMO beginning in 2010:

((1013/BP)*0.0102)+Depth/Level = cDepth/cLevel.

Salinity Units Qualifier:

In 2013, EXO sondes were approved for SWMP use and began to be utilized by Reserves. While the 6600 series sondes report salinity in parts per thousand (ppt) units, the EXO sondes report practical salinity units (psu). These units are essentially the same and for SWMP purposes are understood to be equivalent, however psu is considered the more appropriate designation. Moving forward the NERR System will assign psu salinity units for all data regardless of sonde type.

Turbidity Qualifier:

In 2013, EXO sondes were approved for SWMP use and began to be utilized by Reserves. While the 6600 series sondes report turbidity in nephelometric turbidity units (NTU), the EXO sondes use formazin nephelometric units (FNU). These units are essentially the same but indicate a difference in sensor methodology, for SWMP purposes they will be considered equivalent.  Moving forward, the NERR System will use FNU/NTU as the designated units for all turbidity data regardless of sonde type. If turbidity units and sensor methodology are of concern, please see the Sensor Specifications portion of the metadata.

10. Coded variable definitions

wkb = Weeks Bay NERR

wq = Water quality data

Sampling Station:	Sampling Site Code:		Station Code:

Fish River			FR			wkbfrwq

Weeks Bay			WB			wkbwbwq

Middle Bay			MB			wkbmbwq

Magnolia River		MR			wkbmrwq

11. QAQC Flag Definitions

QAQC flags provide documentation of the data and are applied to individual data points by insertion into the parameter’s associated flag column (header preceded by an F_).   During primary automated QAQC (performed by the CDMO), -5, -4, -2, 2, and 3 flags are applied automatically to indicate data that is above or below sensor range, missing, or outside 2 or 3 standard deviations from the historical seasonal mean.  All remaining data are then flagged 0, as “good”.   During secondary and tertiary QAQC 1, -3, and 5 flags may be used to note data as suspect, rejected due to QAQC, or corrected.

-5	Outside High Sensor Range

-4	Outside Low Sensor Range

-3	Data Rejected due to QAQC

-2	Missing Data

-1	Open - reserved for later flag

0		Good Data

1		Suspect Data

2		Open - reserved for later flag

3		Open - reserved for later flag

4		Historical Data:  Pre-Auto QAQC

5		Corrected Data

12. QAQC Code Definitions

QAQC codes are used in conjunction with QAQC flags to provide further documentation of the data and are also applied by insertion into the associated flag column.  There are three (3) different code categories, general, sensor, and comment.  General errors document general problems with the deployment or YSI datasonde, sensor errors are sensor specific, and comment codes are used to further document conditions or a problem with the data.  Only one general or sensor error and one comment code can be applied to a particular data point.

General Errors

GIM	Instrument Malfunction

GMT 	Instrument Maintenance

GNF	Deployment Tube Clogged / No Flow

GOW	Out of Water Event

GPF	Power Failure / Low Battery

GQR	Data Rejected Due to QA/QC Checks

Sensor Errors

SBO	Blocked Optic

SCF	Conductivity Sensor Failure

SDO	DO Suspect

SDP	DO Membrane Puncture

SIC	Incorrect Calibration / Contaminated Standard

SNV	Negative Value

SOW	Sensor Out of Water

SPC	Post Calibration Out of Range

SSD	Sensor Drift

SSM	Sensor Malfunction

SSR	Sensor Removed / Not Deployed

STF	Catastrophic Temperature Sensor Failure

STS	Turbidity Spike

SWM	Wiper Malfunction / Loss

Comments

CAF	Acceptable Calibration/Accuracy Error of Sensor

CAP	Depth Sensor in Water, Affected by Atmospheric Pressure

CBF	Biofouling

CCU	Cause Unknown

CDA	DO Hypoxia (<28% sat)

CDB	Disturbed Bottom

CDF	Data Appear to Fit Conditions

CFK	Fish Kill

CLT	Low Tide

CND	New Deployment Begins

CRE	Significant Rain Event

CSM	See Metadata

CTS	Turbidity Spike

CWD	Data Collected at Wrong Depth

13. Post Deployment Information

End of Deployment Post-calibration Readings in Standard Solutions, un-cleaned probes:

[Standards used: Specific conductivity = 24.82; DO = 100%; pH = 7.00 and 10.00; turbidity = 0.0 NTU]

14. Other Remarks/Notes

Data are missing due to equipment or associated specific probes not being deployed, equipment failure, time of maintenance or calibration of equipment, or repair/replacement of a sampling station platform.  If additional information on missing data is needed, contact the Research Coordinator at the reserve submitting the data.

For more information and to link meteorological data with water quality data, researchers should request the accompanying SWMP MET data with the SWMP WQ data. These data are available at the CDMO home page http://cdmo.baruch.sc.edu/ or by contacting the Research Coordinator at Weeks Bay Reserve (see contact information above).

At sites WB, MB and MR, YSI EXO2 datasondes were used exclusively – YSI 6600-v2 datasondes were used at site FR exclusively.

From approximately June 19, 2017 through June 24, 2017; Tropical Storm Bret and Tropical Storm Cindy caused greater than average rainfall along the gulf coast from Louisiana to the western Florida panhandle.