Mission Aransas (MAR) NERR Meteorological Metadata

January – December 2008

Latest Update: 02/04/2013

I. Data Set & Research Descriptors

- Principal investigator & contact persons:

Address: 	Mission Aransas National Estuarine Research Reserve

750 Channelview Drive

Port Aransas, TX 78373

Phone:  (361)749-6711

Fax:	 (361)749-6777

Contact Persons:

Dr. Edward Buskey, Research Program Coordinator

E-mail: ed.buskey@mail.utexas.edu; (361)749-6794

Sally Morehead, Manager

E-mail: sally.morehead@mail.utexas.edu; (361)749-6771

Cammie Hyatt, SWMP Data Manager & Technician

E-mail: cammie.hyatt@mail.utexas.edu; (361)749-6774

Colt Cook, SWMP Technician

E-mail : coltcook@mail.utexas.edu; (361)749-6843

Brad Gemmell, SWMP Field Technician

E-mail: bgemmell@mail.utexas.edu; (361)749-6702

- Entry verification :

Prior to 04/17/07, data were uploaded from the CR1000 datalogger to a Personal Computer (IBM compatible). Files were exported from Loggernet in a comma-delimited format (.CDF) and uploaded to the CDMO where they underwent automated primary QAQC and became part of the CDMO’s online provisional database. During primary QAQC, data were flagged if they were missing or out of sensor range. The edited file was then returned to the Reserve where it was opened in Microsoft Excel and processed using the CDMO’s NERRQAQC Excel macro. The macro inserts station codes, creates metadata worksheets for flagged data, and graphs the data for review. It allows the user to apply QAQC flags and codes to the data, append files, and export the resulting data file to the CDMO for tertiary QAQC and assimilation into the CDMO’s authoritative online database. For more information on QAQC flags and QAQC codes, see Sections 11 and 12.

On 04/17/07, a Campbell CFM100 was installed on the CR1000 and code was uploaded via a laptop to direct the CR1000 to start sending data to the compact flash card within the CFM100. Once a month, the compact flash card is swapped out with a blank card and the full card is taken back to the lab where the file on it is converted using Loggernet. The resulting comma-delimited format (.CDF) file is then uploaded to the CDMO and handled in the same manner as listed above. Cammie Hyatt is responsible for data management.

3)  Research Objectives:

Five long-term monitoring stations have been established along the Texas Coast in the Mission-Aransas estuary to collect baseline data on environmental conditions within the bay systems. These System Wide Monitoring Program Stations (SWMPs) provide continual monitoring of temperature, salinity, oxygen and other parameters that characterize the water quality of the bays. In addition to the water quality stations, one meteorological station has been installed in Copano Bay (east) to collect data on air temperature, relative humidity, barometric pressure, wind speed, wind direction, precipitation, and solar radiation (PAR). The goal of the SWMP meteorological station is to develop quantitative measurements of short-term variability and long-term changes in the weather within the Mission Aransas NERR. This data will help us determine climactic and hydrological patterns that influence freshwater inflow in the Mission- Aransas Estuary and will be used to help coastal managers make informed decisions on local and regional issues, such as “no discharge” zones for boats and measuring the success of restoration projects. To the best of my knowledge, this is the only water-based weather station within the NERR.

4)  Research Methods:

Campbell Scientific data telemetry equipment (including the CR1000 datalogger) was installed at the MARCEMET station on 11/14/06 and transmits data to the GOAA GOES satellite, NESDIS ID #3B039516. (Where 3B039516 is the GOES ID for that particular station.) The transmissions are scheduled hourly and contain four (4) data sets reflecting fifteen minute data sampling intervals. Upon receipt by the CDMO, the data undergoes the same automated primary QAQC process detailed in Section 2 above. The “real-time” telemetry data become part of the provisional dataset until undergoing secondary and tertiary QAQC and assimilation in the CDMO’s authoritative online database. Provisional and authoritative data are available at http://cdmo.baruch.sc.edu.

The 15 minute Data are collected in the following formats for the CR1000:

Averages from 5-second data:

Air Temperature (°C), Relative Humidity (%), Barometric Pressure (mb),

Wind Speed (m/s), Wind Direction (degrees), Wind Direction Standard Deviation (degrees) (began collecting as a standard parameter on 11/18/2008), Battery Voltage (volts)

Maximum, Minimum, and their times from 5-second data:

Air Temperature (°C) (these data are not in the dataset but are available from the MAR NERR)

Maximum and their times from 5-second data:

Wind Speed (m/s) (began collecting as a standard parameter on 4/22/2008)

Totals:

Precipitation (mm), PAR (millimoles/m2), and Cumulative Precipitation (mm) (began collecting as a standard parameter on 5/6/2008)

Sensors on the weather station are inspected monthly and repaired and/or cleaned as needed. Sensors are removed and sent back to Campbell Scientific for calibration either annually or bi-annually:

- Temperature/Humidity- annual recalibration

- Rain Gauge- annual recalibration

- Wind Speed/Direction- bi-annual recalibration

- Barometric Pressure- bi-annual recalibration

- PAR- bi-annual recalibration

5)  Site location and character:

The MANERR meteorological station (MARCEMET) is located in open water in east Copano Bay near the Copano Causeway at 28º 7.941’N, 97º 2.067’W. The weather station is located on a platform that is 3.5 meters above the water (at the high water mark). Wind (speed and direction), temperature, relative humidity, and PAR sensors are mounted onto a 3 meter aluminum tower so that the sensors stand approximately 4 meters off the platform. The PAR sensor is located on one side of the 1.2 meter sensor crossarm, with the wind monitor on the other side. The rain gauge is located on the arm rail of the platform which is approximately 1.5 meters off the platform. The barometric pressure sensor is located inside the Campbell Scientific enclosure, which is 1.1 m off the platform, and has an intake valve that runs through a hole in the bottom of the enclosure that allows pressure equalization.  To the best of my knowledge, all the weather instrumentation and associated sensors are set up according to standard CDMO protocols. There are no obstructions around the platform; therefore, the tower does not receive shade and it is not blocked from the wind. This site also has a SWMP water quality monitoring station installed on the lower level of the platform.

Major climactic influences within the MANERR are temperature, precipitation, evaporation, wind, tropical storms, and hurricanes. The lowest yearly air temperature recorded at MARCEMET was 0.9 ºC and the highest temperature was 33.7 ºC, with an average temperature of 21.9 ºC. Minimum wind speed recorded was zero meters per second (m/s) and the maximum was 20.2 m/s, with an average of 6.0 m/s. The lowest yearly reading of relative humidity was18% and the highest was 98%, with an average of 76% for the year. Total rainfall for the year measured 47.9 cm.

6)  Data collection period:

Data collection at the MARCEMET began in November 2006. Data was initially recorded onto the CR1000 and was downloaded via laptop as indicated in Section 2.

Once the CFM100 was installed, the flash card holding the data was swapped out with a clean card once a month. The starting and ending times of those flash card files are listed as follows:

12/04/2007 10:15 – 01/08/2008 09:45

01/08/2008 10:00 – 02/19/2008 11:00

02/19/2008 11:15 – 04/02/2008 10:15

04/02/2008 10:30 – 04/22/2008 10:30

04/22/2008 10:45 – 05/06/2008 09:30

05/06/2008 10:00 – 05/20/2008 10:00

05/20/2008 10:15 – 07/01/2008 08:45

07/01/2008 09:00 – 07/29/2008 10:00

07/29/2008 14:00 – 08/26/2008 09:00*

08/26/2008 09:15 – 09/09/2008 09:30

09/09/2008 09:45 – 09/23/2008 09:00

09/23/2008 09:15 – 10/07/2008 09:30

10/07/2008 09:45 – 10/21/2008 08:45

10/21/2008 09:00 – 11/04/2008 09:30

11/04/2008 09:45 – 11/18/2008 10:00

11/18/2008 10:15 – 12/02/2008 09:45

12/02/2008 10:00 – 12/17/2008 10:45

12/17/2008 11:00 – 01/07/2009 10:00

*There was a problem with the flash card that was swapped on 07/29/2008 resulting in a loss of data from 07/29/2008 14:00 – 08/26/2008 09:00. Efforts were made to repair the file; however, they were unsuccessful. This data was later obtained from the CDMO and added into the yearly appended file for secondary QAQC.  Data for 2008 were collected from 00:00 January 1 thru 23:45 December 31.

7)  Distribution:

NOAA / ERD retains the right to analyze, synthesize and publish summaries of the NERRS System-Wide Monitoring Program data. The PI retains the right to be fully credited for having collected and processed the data. Following academic courtesy standards, the PI and NERR site where the data were collected will be contacted and fully acknowledged in any subsequent publications in which any part of the data are used. Manuscripts resulting from this NOAA / OCRM supported research that are produced for publication in open literature, including refereed scientific journals, will acknowledge that the research was conducted under an award from the Estuarine Reserves Division, Office of Ocean and Coastal  Resource Management, National Ocean Service, National Oceanic and Atmospheric Administration. The data set enclosed within this package / transmission is only as good as the quality assurance and quality control procedures outlined by the enclosed metadata reporting statement. The user bears all responsibility for its subsequent use / misuse in any further analyses or comparisons. The Federal government does not assume liability to the Recipient or third persons, nor will the Federal government reimburse or indemnify the Recipient for its liability due to any losses resulting in any way from the use of this data.

NERR weather data and metadata can be obtained from the Research Coordinator at the individual NERR site (please see Principal investigators and contact persons), from the Data Manager at the Centralized Data Management Office (please see personnel directory under the general information link on the CDMO home page) and online at the CDMO home page http://cdmo.baruch.sc.edu. Data are available in comma separated format.

8)  Associated researchers and projects:

In addition to the meteorological station, the Copano East site is also equipped with a YSI datasonde that measures Water Temperature, Specific Conductivity, Salinity, Dissolved Oxygen (% and mg/L), Depth, pH, and Turbidity. Water is also collected at this site for nutrient analysis to determine the amount of chlorophyll a, orthophosphate, ammonium, nitrogen (nitrite, nitrate, or nitrite + nitrate), and silicate that is present.

II. Physical Structure Descriptors

9)  Sensor Specifications:

Parameter: Temperature

Units: Celsius

Sensor Type: Platinum resistance temperature detector (PRT)

Model #: HMP45C Temperature and Relative Humidity Probe

Operating Temperature: -40 ºC to +60 ºC

Range: -40 ºC to +60 ºC

Accuracy at 20 ºC: ± 0.2 ºC

Date of Last calibration: 05/27/2008

Parameter: Relative Humidity

Units: Percent

Sensor Type: Vaisala HUMICAP© 180 capacitive relative humidity sensor

Model #: HMP45C Temperature and Relative Humidity Probe

Range: 0-100% non-condensing

Accuracy at 20 ºC: ± 2% RH (0-90%) and ± 3% (90-100%)

Temperature dependence of RH measurement: ± 0.05% RH/ ºC

Date of Last calibration: 05/27/2008

Parameter: Barometric Sensor

Units: millibars (mb)

Sensor type: Vaisala Barocap© silicon capacitive pressure sensor

Model #: CS-105 (1/01/2008-7/29/2008), CS-106 (7/29/2008-12/31/2008)

Operating Range: Pressure 600 to 1060 mb; Temperature: -40 ºC to +60 ºC; Humidity: non-condensing

Accuracy: ±  0.5 mb @ 20 ºC; ±  2 mb @ 0 ºC to 40 ºC; ± 4 mb @ -20 ºC to 45 ºC; ± 6 mb @ 40 ºC to 60 ºC

Stability: ± 0.1 mb per year

Date of Last calibration: 11/30/2005 (CS-105); 04/04/2008 (CS-106)

Parameter: Wind speed

Units: meter per second (m/s)

Sensor type: 18 cm diameter 4-blade helicoids propeller molded of polypropylene

Model #: R.M. Young 05103 Wind Monitor

Range: 0-60 m/s (134 mph); gust survival 100 m/s (220 mph)

Accuracy: ± 0.3 m/s

Date of Last calibration: 02/21/2008

Parameter: Wind direction

Units: degrees

Sensor type: balanced vane, 38 cm turning radius

Model #: R.M. Young 05103 Wind Monitor

Range: 0 to 360º mechanical, 355º electrical (5º open)

Accuracy: ± 3 degrees

Date of Last calibration: 02/21/2008

Parameter: LI-COR Quantum Sensor

Units: mmoles m-2 (total flux)

Sensor type: High stability silicon photovoltaic detector (blue enhanced)

Model #: LI190SB

Light spectrum waveband: 400 to 700 nm

Temperature dependence: 0.15% per ºC maximum

Stability: < ± 2% change over 1 yr

Operating Temperature: -40 ºC to +60 ºC; Humidity: 0 to 100%

Sensitivity: typically 5 μA per 1000 μmoles s-1m-2

Date of Last calibration: 11/20/2007

PAR multiplier = 1.25999

Parameter: Precipitation

Units: millimeters (mm)

Sensor type: Tipping Bucket Rain Gauge

Model #: TE525

Rainfall per tip: 0.01 inch

Operating range: Temperature: 0 to 50 ºC; Humidity: 0 to 100%

Accuracy: ± 1.0% up to 1 in/hr; +0, -3% from 1 to 2 in./hr; +0, -5% from 2 to 3 in./hr

Date of Last calibration: 02/26/2008

The CR1000 has two MB Flash EEPROM that is used to store the Operating System. Another 128 K Flash is used to store configuration settings. A minimum of 2 MB SRAM (4 MB optional upgrade) is available for program storage (16K), operating system use, and data storage. Additional storage is available by using a compact flash card (1GB) in the optional CFM100 Compact Flash Module.

10)  Coded variable definitions:

Sampling station:		Sampling site code:		Station code:

Copano East			CE				marcemet

11) QAQC flag definitions:

QAQC f lags provide documentation of the data and are applied to individual data points by insertion into the parameter’s associated flag column (header preceded by an F_). During primary automated QAQC (performed by the CDMO), -5, -4, and -2 flags are applied automatically to indicate data that are above or below sensor range, or missing. All remaining data are then flagged 0, as passing initial QAQC checks. During secondary and tertiary QAQC 1, -3, and 5 flags may be used to note data as suspect, rejected due to QAQC, or corrected.

-5	Outside High Sensor Range

-4	Outside Low Sensor Range

-3	Data Rejected due to QAQC

-2	Missing Data

-1	Optional SWMP supported parameter

0	Passed Initial QAQC Checks

1	Suspect Data

2 	Open – reserved for later flag

3	Open – reserved for later flag

4	Historical Data : Pre-Auto QAQC

5	Corrected Data

12) QAQC code definitions:

QAQC codes are used in conjunction with QAQC flags to provide further documentation of the data and are also applied by insertion into the associated flag column. There are three (3) different code categories : general, sensor, and comment.  General errors document general problems with the CR1000, sensor errors are sensor specific, and comment codes are used to further document conditions or a problem with the data. Only one general or sensor error and one comment code can be applied to a particular data point.

General Errors

GIM	Instrument Malfunction

GIT 	Instrument Recording Error, Recovered Telemetry Data

GMC	No Instrument Deployed due to Maintenance/Calibration

GMT	Instrument Maintenance

GPD	Power Down

GPF	Power Failure / Low Battery

GPR	Program Reload

GQR	Data Rejected Due to QA / QC Checks

GSM	See Metadata

Sensor Errors

SIC	Incorrect Calibration Constant, Multiplier or Offset

SNV	Negative Value

SOC	Out of Calibration

SSN	Not a Number / Unknown Value

SSM	Sensor Malfunction

SSR	Sensor Removed

Comments

CAF	Acceptable Calibration / Accuracy Error of Sensor

CDF	Data Appear to Fit Conditions

CRE	Significant Rain Event

CSM	See Metadata

CVT	Possible Vandalism / Tampering

13) Other remarks / notes:

Data are missing due to equipment or associated specific sensors not being deployed, equipment failure, time of maintenance or calibration of equipment, or repair/replacement of a sampling station platform.  Any NANs in the dataset stand for  “not a number” and are the result of low power, disconnected wires, or out of range readings.  If additional information on missing data is needed, contact the Research Coordinator at the reserve submitting the data.

Small negative PAR values are within the range of the sensor and are due to normal errors in the sensor and the CR1000 Datalogger. The Maximum signal noise error for the Licor sensor in +/-2.214 mmoles/m2 over a 15 minute interval.

Relative Humidity data greater than 100 are within the range of the sensor accuracy of +/-3%.

Cumulative precipitation data are recorded from 00:00 to 23:59 with the daily total recorded at the midnight mark (00:00). The midnight CumPrcp value is actually the total from the previous day.

All Reserves were required to align their wind sensors to True North by April 1, 2008.  MAR NERR’s wind sensor was aligned by that date, however prior to aligning the sensor, the alignment was only off by 4 degrees.

The following comments are referred to in the flag columns within the data set with the CSM (see metadata) flag.

Prior to 07/29/2008, the following parameters were collected with out of calibrations sensors while we were in the process of purchasing a second set of replacement sensors:

Wind speed/direction

Air Temperature

Relative Humidity

Barometric Pressure

Precipitation

PAR

There were many nights during the year in which the nighttime PAR readings exceeded the noise error of the LICOR (+/- 2.214) and were rejected. I am unsure as to the reason for the higher readings; however, since the daytime readings were lower than normal, I still rejected those nighttime readings. I suspect that the nighttime PAR readings may be caused by moisture in the air due to fog or drizzle that may not have been enough to tip the rain bucket.

There were many occasions when the daytime PAR reading suddenly dropped and there was no apparent reason for it. I suspect that a bird was either sitting on the sensor or very close to it to cause such a low reading since we saw an osprey sitting on the crossbar of the MET station on several occasions. Those data were rejected.

01/01/08 00:00 – 01/19/08 07:15  PAR data rejected due to abnormally low daytime readings that I believe were caused by a build-up of bird poo that was washed off during the rainstorm toward the end of that time period.

01/08/08 09:45-10:00			Cleaned rain bucket during MET inspection;

Corrected data for Prcp.

09:53		Swapped flash card.

01/19/08 07:30 – 07/14/08 19:30	PAR daytime readings marked as suspect even though in most cases readings seem normal, I feel that something is off with the sensor. The following daytime PAR readings appeared higher than normal for that time of day:

01/26/08 13:00-13:45

03/13/08 12:15

03/22/08 12:15

02/19/08 11:17			Swapped flash card.

03/21/08 10:50			Swapped flash card; however, when I tried to convert it back in the lab, there was nothing on the card.

04/22/08 10:43			Swapped flash card.

05/06/08 09:35			Swapped flash card.

09:50	Uploaded code to the CR1000 (MARCEMET_V5.5_040308.CR1)

05/20/08 10:05			Swapped flash card.

07/01/08 08:58			Swapped flash card.

09:15			Cleaned rain bucket;

Corrected data for Prcp and CumPrcp.

07/14/08 19:45 – 07/29/08 10:00	PAR data rejected since daytime readings are way below normal and nighttime readings are way above normal.

07/22/08 13:30 – 07/25/08 10:30	Hurricane Dolly (lots of wind and rain).

07/29/08 10:13			Swapped flash card.

07/29/08 10:15 – 13:45		Power down; swapping out sensors.

13:54			Uploaded code to the CR1000

(MARCEMET_V5.6_071508.CR1

14:30	Tipped rain bucket to make sure it was installed properly; Corrected data for Prcp and CumPrcp.

07/29/08 14:00 – 08/26/08 09:00	Data had to be collected from telemetry because we swapped the flash card prior to uploading the new program; therefore, no new data was collected on the card since it was formatted with the old program.

08/15/08 10:45			PAR reading marked as suspect since it is much higher than normal for that time of day.

08/23/08 11:45 - 12:00		PAR reading marked as suspect since it is much higher than normal for that time of day.

08/26/08 09:10			Swapped flash card.

09:15	Cleaned rain bucket; Corrected data for Prcp and

CumPrcp.

16:45 – 17:30		PAR readings much lower than normal for that time of day and comparisons with PAR readings on the roof of UTMSI do not indicate cloudiness; however, since the locations are very different for these two sensors, it is hard to make a direct comparison. There could be an isolated storm cell in the bay.

09/09/08 09:38			Swapped flash card.

09/12/08 07:30 – 09/16/08 12:00	Hurricane Ike

09/20/08 12:30			PAR reading marked as suspect since it is much higher than normal for that time of day.

09/23/08 09:10			Swapped flash card.

9:15 – 18:00			PAR readings seemed low for that time of day; however, I checked local weather reports and there were thunderstorms in the area (even though we didn’t receive much rain) and the cloud cover was responsible for the low readings. I think the moisture in the air was the cause of the negative nighttime values that night and for a few nights afterwards.

10/07/08 09:33			Swapped flash card.

10/13/08 06:30 – 18:00		Marked PAR readings as suspect since overnight readings were so high.

10/14/08 06:45 – 18:15		Marked PAR readings as suspect since overnight readings were so high.

10/21/08 08:56			Swapped flash card.

11/04/08 09:46			Swapped flash card.

11/06/08 11:45			PAR reading marked as suspect since it is much higher than normal for that time of day.

11/18/08 10:17			Swapped flash card.

11/29/08 13:45 – 16:45		PAR readings much lower than normal for that time of day but may be due to overcast weather since there were many overcast days in November and December.

12/02/08 10:00			Swapped flash card.

12/16/08 07:45 – 17:45		Very foggy day.

12/17/08 07:45 – 17:30		Very foggy day.

11:02			Swapped flash card.

Monthly maintenance checks are performed to clean the instrumentation, if necessary, and to determine whether the sensors have drifted. A handheld Kestrel 4000 is used to check the air temperature, barometric pressure, relative humidity, and wind speed. For November and December, PAR readings were checked using a handheld LICOR quantum sensor. Due to weather conditions, the June equipment check wasn’t performed until July 1, 2008 and the July check was done at the end of that month.